I.  	Purpose:

To define the Soft computer process when resulting Routine Coagulation tests.

II.  	Procedure:

Orderable Tests:				Resultable Tests:

PTINR – PT and INR				PT and INR

PTT – Partial Thromboplastin Time		PTTI

PTPTT – PTINR and PTT			PT, INR, PTTI

FIB – Fibrinogen				FIB

DDI – D-Dimer				DDI

NINR – Reflexes on when PTINR is ordered (non-reportable test)

MCOAG – Resulting Template for Main Lab Coagulation

Canned Messages:

@HEM - Sample hemolyzed

@ICT - Sample icteric

@LIP - Sample lipemic

@POLY - APTT ____ seconds when repeated using 1.5 6.0 polybrene.

@POSH - APTT shortened to ____ seconds when repeated using ___ ug/ml polybrene

indicating the presence of heparin.

@POHE - APTT did not correct when repeated using polybrene, indicating heparin is not

present.

@RDCH - ?Result double checked.

Cancellation Comments:

|X#01	Sample clotted

|X#02	Quantity not sufficient

|X#04	Laboratory accident

|X#05	Accessioning error

|X#06	Incorrect tube/container type

|X#08	Does not meet clinical time requirements

|X#10	Broken/spilled in transit

|X#11  Duplicate test

|X#13	IV contamination

|XH01	Unable to report due to specimen problems.

|XH09	Test not indicated.

Pending test list:

-Open Resulting Worklist:

∙ Double click on Resulting Worklist on main screen

∙ Result Worklist box will appear:

Template: type in MCOAG

Status: Pend and Nonver

From order # - Make From order #: and To: blank to see all

pending tests leave (or can adjust date and order numbers to see only today or previous days)

∙ Click OK

-Any order #’s that appear on the list have pending tests

-Tests that don’t have a V next to it are the tests that are pending for an order

**Also call pending test list for special coagulation (SPCOG) every hour to save samples

**(See page 5 for Special Coagulation Tests)

-To call extra labels from Resulting Worklist (to use for special coagulation aliquot tubes):

∙  Select the order that needs labels printed

∙  Click on the Printer icon

∙  Click on Selected Label

∙  Change the printer to the appropriate label printer

∙  Click on the box next to HBT (Hematology Blue Top) so there’s a check in the box

∙  If more than one label is neededClick on the Label Properties tab

∙  Change the # of copies to the # of labels to be printed

∙  Click OK and labels will print to the label printer specified

Samples to be run on BCS:

-Sample arrives at the bench

-Spin sample in centrifuge for 10 minutes

-Uncap samples, place in BCS rack and run on BCS

-Open Instrument Menu

∙  Double Click Instrument Menu on main screen

∙  Under Select Instrument: Choose either ID MBCS1, MBCS2 or MBCS3

∙  Under Select Date: Choose Loadlist and today’s results

-Tests that are pending are under Test List

∙  If there is a V in the Flags column and by a test result then that test was

autoverified but there is a pending test on the order

-Select the Order # you want to verify

-Check results for flags and errors.  Look at test history if needed.

∙  If a result fails delta check there will be %D in the Flags column.  Compare

results to previous results (Prev Res) in Lab Results and/or look at Test History

∙  If results need to be checked from past the previous result Click on the result to

be checked the Click on Tst Hist.

-A Test History box will appear with a list of previous results for the selected test.  Results will be listed numerically and in a graph.  Numerically listed results will have flags is abnormal while normal results on the graph are green and abnormal results on the graph are yellow.

-When done looking at Test History Click OK.

-To check if the patient is on any anticoagulants

∙  Click on Lab Results tab

∙  Look at the Results for test COAGS

∙  COAGS is Anticoagulant Statusresults will be Heparin, Coumadin, Both or None

∙  Click on Instrument to look at results from the instrument

-If no result comes across the instrument menu

∙  The sample was not run

∙  The result is in orange with a ? (ex. 55.9?) and will not transmit from the

instrument across the LIS – follow lab protocol on running sample on a different

wavelength or not reporting the result

∙The result is an ! and will not transmit from the instrument across the LIS

Samples to be verified on BCS:

**Results should be verified ONLY in the Instrument tab (with the exception of entering PTT >120.0 or entering tests that are performed manually) **

-If results are OK to release:

∙  Select Post All to verify results (Tech initials, date verified and time verified

will autofill)

-If results need to be repeated:

∙  Select (left click) the test to be repeated and press F4.  This will “cancel” the first run from being sent out.  A “C” will appear in the Status column.

∙  Rerun sample on instrument

∙  If the rerun is the result to be reported click Verify

∙  If the original result is to be reported:

-Select (left click) the rerun result and press F4.

-Select (left click) the original result and press F4.

∙  Verify the result by clicking Post All

-If the PT is >120.0 seconds and the numerical result is sent from the BCS:

(Example: PT 121.9 is the result on the BCS will go into Soft as >120.0)

∙  The PT result of >120.0 will be in the result box

∙  The INR should be reported as >13.00

-If the PT  is >120.0 seconds and no numerical result is sent from the BCS:

(Example: ! or no clot is the result on the BCS and no result is in Soft)

∙  Click on the Lab Results tab

∙  Select the result box that corresponds with PT

∙  Click Def Res and the result box will autofill with >120.0

∙  Select the result box that corresponds with INR

∙  Click DefRes and the result box will autofill with >13.00

∙  A Result Information box will appear containing:

-Test name

-Result and Date

-Value out of NORMAL range

-The Normal Range for the test

-Result fails delta check if applicable

∙  A Result Information box will appear containing:

-INR test information and that the critical INR needs to be called

-If the PTT is >120.0 seconds and the numerical result is sent from the BCS:

(Example: PTT 126.5 is the result on the BCS will go into Soft as >120.0)

∙  The PTT result of >120.0 will be in the result box

∙  Click Verify to verify all results

-If the PTT is >120.0 seconds and no numerical result is sent from the BCS:

(Example: ! or no clot is the result on the BCS and no result is in Soft)

∙  Click on the Lab Results tab

∙  Select the result box that corresponds with PTT

∙  Click Def Res and the result box will autofill with >120.0

∙  A Result Information box will appear containing:

-Test name

-Result and Date

-Value out of NORMAL range

-The Normal Range for the test

-Result fails delta check if applicable

∙  Click OK

∙  Make sure only the PTT result box is selected, Click Ver to verify only

the PTT and then click on the Save icon

-If a “Result double checked” comment needs to be entered:

∙ For PT and PTT 	-Click on Comment

-Click on CanMesg

-Click on @RDCH

-Click OK (There will be an M in the M column meaning

that there is a message attached to the result)

∙ For INR, FIB or DDI-Click on Comment

-Type in @RDCH (all capitol letters)

-Hit Enter (the comment will appear in the Results comment box)

-Click OK (There will be an M in the M column meaning

that there is a message attached to the result)

-If a Fibrinogen Low or Fibrinogen 57.Low result is below the curve

∙  Result of <70 will appear in the Instrument menu

-If a D-Dimer result is <0.17 then a result of <0.17 will be reported

-If a D-Dimer result is >4.30 the result will auto cancel in the Instrument menu and need to be run on D-DDi.H

-If a D-Dimer result is >34.00 then a result of >34.00 will be reported

-If a result needs to be manually entered:

∙  If have a test has to be resulted manually, example manual PT, enter the result in

Lab Result tab OR Result entry.

-If INR is >4.0:

∙  If the INR is >4.0 then the result will not autoverify and the critical value needs

to be called to the appropriate care provider

Samples that are >4 hours old:

-If a specimen is run that is >4 hours old then a canned message will automatically be resulted with test result with a disclaimer that the specimen is >4 hours old.

-The comment code that will be seen is @OLD

-The result will be shown in the result box followed by @OLD

-DO NOT remove this comment if it appears with results

Samples that require a comment:

-Select the result box of the test that requires a comment

-Click on Comment

-Click on CanMesg to open a drop down of result comments that correspond to the

selected test

-Highlight the message to be added and click OK

-A box with the contents of the canned message will appear

∙  If the comment is correct click OK

∙  If the comment is not correct click OK or on the X in the upper right hand

corner and then highlight the comment and hit Backspace on the keyboard

-The comment ID (ex. @LIP) will appear after the result

Special Coagulation Orderable Tests:

SPCOG – Resulting Template for Special Coagulation Tests

A2AP – Alpha-2-Antiplasmin Activity

ATIII – Antithrombin Enzymatic Activity

APC – APC Resistance

HA – Anti-Xa Activity, LMWH

HAUFH – Anti-Xa Activity, UFH

ARIX – Anti-Xa Activity, Fondaparinux

INHIB – Inhibitor Assay

PHOSN – Phospholipid Neutralization

PLG – Plasminogen

PRC – Protein C, Functional

PRS – Protein S, Functional

PRSTF – Protein S, Total and Free Antigen

PTM – Prothrombin Time 1:1 Mix

PTTM – Partial Thromboplastin Time 1:1 Mix

RCF – von Willebrand Factor Activity

DRV – Dilute Russell Viper Venom Time

TT – Thrombin Time

VWFAG – von Willebrand Factor Antigen

FII – Factor II Activity

FV – Factor V Acitivity

FVII – Factor VII Activity

FVIII – Factor VIII Activity

FIX – Factor IX Activity

FX – Factor X Activity

FXI – Factor XI Activity

FXII – Factor XII Activity

FXIII – Factor XIII Activity

LUPUS - Lupus Anticoagulant Panel (DRV, PTT and PHOSN)

HYPER - Hypercoagulable Panel (APC, ATIII, PRC and PRS)

VWD - von Willebrand Panel (VWFAG, RCF and FVIII)

Special Coagulation Orderable Sendout Tests:

SCOGS – Resulting Template for Special Coagulation Sendouts

ATAG - ATIII Antigen

PAIAC - Plasminogen Activator Inhibitor-1, Activity

PRCA - Protein C Antigen

TPA - Tissue Plasminogen Activator

C4BPR - C4 Binding Protein

PAIAG - Plasminogen Activator Inhibitor-1, Antigen

PAIM - PAI-1, Mutation (4G/5G)

PTFR - Prothrombin fragment 1&2

TATC - Thrombin-Antithrombin complex

VWMA - von Willebrand Multimeric Analysis

VWFPA - vWF Protease Activity w/Rflx Inhibitor

REPTM – Reptilase Time w/Rflx Rep Time 1:1 Mix

CRFB - Cryofibrinogen

III.	HISTORY:

H-1	This procedure was written by Jessica Kaul on 3/7/2010.

H-2	This procedure was revised by Jessica Kaul on 6/16/2011.

H-3	This procedure was reviewed by Nancy Kriz on 9/20/2013