- PURPOSE:

The quality control (QC) program consists of both internal and external Quality Control. 	The internal QC consists of commercial and in-lab prepared materials, which are used to 	verify the accuracy and precision of our test procedures and instrument moving average 	programs. External Quality control consists of Proficiency Testing materials obtained 	from the CAP, Streck, and/specimens submitted to another laboratory.

- SCOPE:

All technical staff who perform instrument set-up, analyze specimens, and validate/report results must participate in the quality control program. Quality control specimens must be analyzed and acceptable before patient specimens’ area analyzed.

# INTERNAL QUALITY CONTROL

### I.	CELL COUNTERS

## A. 	Performing Quality Control

1.	Commercial control:

Once each shift three levels of commercial cell control (cbc/plt/diff/retic) 			are analyzed, any control which is out of range should be repeated one 			time. If at least two of the three levels are within the acceptable range for 			all parameters the run is acceptable. If more than one sample is out of 			range then testing will cease on that instrument and a supervisor must be 			notified (all problems must be logged in instrument problem logs) and

instrument not be used for patient specimen analysis.

2.	Patient References:

Each shift two patient samples are selected; one normal, one abnormal. 				The samples are analyzed 2 times, the results reviewed to assure 					reproducibility, and then the mean values set as the target. These patient 				references should be run every 2 hours, after reagent changes or if 					instrument problem suspected. Acceptable ranges are posted in the 					computer. If at any time a sample has a result outside of acceptable 				range it should be repeated once. If the repeat is still out, testing should 				stop and a supervisor should be notified and 	problem / resolve entered in 				instrument problem log. (refer to Sapphire operators manual for definition

of normal and abnormal references, establishing patient reference

libraries).  If an instrument is not in use or in service references do not

need to be run but status of instrument must be noted in instrument log.

3.	Whole blood comparisons: Monthly 6 patient bloods are analyzed on each 				cell counter. The results means calculated and all instrument results 				compared to assure consistency of results on the cell counters.

Comparison bloods and mean of the 6 bloods must meet the established

tolerance ranges.

4.	Moving averages: Sapphires have moving average programs enabled to

continually monitor x-b, wbc, plt, retic results on an ongoing basis. The

operator is alerted of any deviation from established ranges. The moving

average results are also routinely reviewed daily by a supervisor for trends

or problems.

5.        Commercial calibrator: Performed at least every 6months, after major

instrument service, when indicated by QC results, to verify instrument

unexpected shifts/trends.

#### B. 	Review of Quality Control:

- At least weekly (usually daily) a supervisor will review the daily controls to observe patterns or consistent problems with the commercial controls or patient controls or instrument operation. Monthly the QC library  results will be printed, purged  and  signed off by a supervisor.

2.	New lot of commercial control is “overlap” assayed prior to activating the

new control for use. Prior to performing the overlap assay a supervisor

will verify that the instruments are performing satisfactorily and current

lot of  commercial control recovering with posted tolerance ranges.

Overlap assay of new controls with the existing controls is utilized to verify SD and control target mean of manufacturer or if  “in-lab” adjustment of target mean is necessary.

- Peer group statistics: Once a control lot has expired the,

control libraries for that lot are  printed and the 2 months control data is

submitted to Streckstats

### II. 	COAGULATION

##### A.	Performing Quality Control

1.	PT, PTT, d-dimer and Fibrinogen:

Once each shift of use or whenever a reagent is changed, normal and 				abnormal controls are run on the BCS. Results must be within the 					acceptable ranges posted on the instrument. If a value is outside of 					acceptable range it should be repeated once. If still out a supervisor should 			be notified, testing must not continue.

2.	Special Coagulation:

All special coagulation tests have at least two levels (abnormal and 				normal) which must be run with each batch. Results must be within the 				stated ranges for patient results to be reported.

##### B. 	Review of Quality Control:

1.         PT, PTT, d-dimer and Fibrinogen controls are reviewed from the QC libraries at 	least once per week by the Coagulation Specialist. Monthly, the computers	control libraries are printed out and reviewed by the Coagulation Specialist. Statistics are sent to Siemens QAP for computation.

2.          Special coagulation controls are reviewed by the Coagulation Specialists from the QC libraries on BCS or STA compact.

### III. 	URINALYSIS

###### A.	Performing Quality Control:

1.	Two levels of urine control are run each shift of use. The controls must be 			run on Atlas and/or Clinitek 500 if the instrument is to be used for patient 			testing. If an instrument is not used, no controls need to be run but this 			should be noted on the control log. Commercial calibrators and controls 			are run after reagent change. 	2 previously tested urines (normal and 			abnormal) are also run and results reviewed for satisfactory performance.

2.	Clinitest. Two levels of controls are run only at the time of 				                       running the test(s) with a maximum of once/shift.

3.	At least 2 times per year patient urine specimens are run on both Atlas and 		Clinitek 500 and results compared for consistency.

###### B.	Review of Quality Control:

- Monthly the QC file will be signed off by a supervisor.

IV.	 BODY FLUIDS:

A. 	Once each day saline and albumin background are evaluated under the 				microscope for bacteria. If bacteria are discovered, the affected reagent is 				discarded.

B.	Once a shift Cell chex manual control level 1 and 2 are counted. Results must be

within 	established target range. Once a shift Cell chex auto control level 1,2 and

3  are run on the fluid Sapphire. Results must be within established target range.

In  addition a peripheral blood reference is run (a new one each shift) to assure

satisfactory operation of the Sapphire

Both manual and automated commercial controls must be overlapped with current

lot prior to usage.

V. 	MISCELLANEOUS TESTING:

All quality control is discussed in the individual procedures.

All quality control runs are logged whether acceptable or unacceptable. All unacceptable 	quality control must have problem resolved and quality control material(s) retested. 	Patient analysis cannot be performed until quality control analysis is satisfactory. A 	supervisor must be informed when second analysis of a control material is unsatisfactory.

EXTERNAL QUALITY CONTROL AND PROFICIENCY TESTING PROTOCOL

I.		PROCEDURE:

Proficiency testing samples are performed in exactly the same manner as patient

tests. They are integrated into routine round the clock patient testing and are

performed by whichever technologist is assigned to a given workstation at the given 	time. Proficiency test specimens and/or results are not shared or referred to other 	Hematology laboratories.

A limited number of CAP surveys (ex stem cell, hgb identification) have multiple

laboratory  components to the survey and these are shared as per CAP survey instructions.

A.The CAP proficiency kits are delivered to lab and person initially handling completes

the first portion on the CAP Survey Board. The supervisor of the area for the survey

initiates the coordination of  the processing of the specimens, results and last portion

of the survey board.

CAP proficiency specimens are analyzed by all staff on a rotational basis.

B.      Upon receipt the survey specimen is stored in the cold room (or at temperature dictated by survey directions)  in the appropriate section of laboratory: main lab, special hematology, and special coagulation. The accompanying sheet is dated.

C.	Within ten working days a supervisor of that area will initiate testing by 				giving survey to a technologist

D.	The technologist will perform the proper warming and mixing technique 				and process the analysis.

E.	After testing, the supervisor or team member will complete the answer 				sheet, get the testing technologist and the Laboratory Director 					signatures, and then submit the results to the coordinator who submits to

CAP. A second designee will review the recording of the

results on the CAP answer sheet before the results are submitted to CAP.

F.	The coordinator will email the results to CAP or the proficiency

agency and retain the paper reports in the office file.

G.	The results of the proficiency report are reviewed by the Lab Medicine

Chair, Compliance Coordinator, Lab Director, the Lab Manager upon

receipt. The supervisor of the specific area will also review results.

H.	All data (raw and completed), reports, and action reports are retained for at 			least two years.

Clinical Hematology performs the following CAP Proficiency Testing Surveys:

| BP | Blood parasite | CGL | Coagulation, limited |
| --- | --- | --- | --- |
| FH3P | Hematology automated differential | RT | Retic |
| HG | Hemoglobinopathy | OCB | Occult blood |
| CMMP | Clinical microscopy | CMP | Clinical microscopy |
| DSC | Dipstick confirmatory | PF | Platelet function |
| PF1 | Platelet function | SCM1 | Special microscopy |
| DCM2 | Special Microscopy | G6PDS | G6PD |
| LN18 | Retic linearity | CGEX | Coagulation extended |
| CGS1 | Coag special testing | CGS2 | Coag special testing |
| CGS4 | Coag special testing | ESR | Sedrate |
| BFC | Crystals | URC | Crystals |
| HFC | Hemocytometer fld | LN9ER | Hematology linearity |
| FCFS | Fecal fat | FlAC | Fecal lactoferrin |
| SCP | Stem cell processing | RMAL | Rapid malaria |
| FLD | Body fluid | PF | Platelet aggregation |
| PF1 | Platelet function | BFC | Crystal |
| VLBS | Virtual smear |  |  |

In addition:

Reticulocyte and CBC/plt Linearity From Streck 2x/year

ZPP- Blood/lead and trace elements paired specimens submitted to ARRUP 2x/year.

Tests where no external proficiency testing is available- tests are performed in a double blind analysis by 2 technologists and final results recorded and compared for consistency.

II.       INTERPRETATION OF CAP Survey RESULTS:

The CAP results are reviewed by the Chairman of Laboratory Medicine, Laboratory 	Director and Laboratory Manager. The Laboratory Manager will coordinate the 	investigation and correction of problems if the testing is unacceptable. Problems that may 	occur are: transcription errors recorded onto the answer sheet, instrument malfunctions, 	reagent problems, and debatable CAP interpretations. Instrument controls, calibrations, 	reagents, and precision are examined when results show a bias or trend. If the results are

ungraded for any reason, they are compared to the group results. If 	discrepancies are

discovered, an investigation is performed, discoveries are recorded, and an action plan

developed. The action plan is filed with the proficiency test response.. In non-CAP

surveys, results are evaluated by the Laboratory Director or Designee. All problems are

investigated, and investigations are recorded and action plans are filed with the results.

III.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 10-10-97.

H-2	This procedure was revised by N. Ortoli-Drew on 7-2005.

H-3	This procedure was revised by N. Ortoli-Drew on 4-2007.

H-4	This procedure was revised by N. Ortoli-Drew on 7-2007.

H-5	This procedure was revised by N. Ortoli-Drew on 10-2008.

H-6       This procedure was revised by N. Ortoli-Drew 4-2011

H-7       This procedure was revised by N. Ortoli-Drew 12-2012

H-8       This procedure was revised by N. Ortoli-Drew 6-2013

H-9       This procedure was revised by N. Ortoli-Drew 9-2013