I.	PURPOSE:

Whole blood comparisons are routinely performed on  all hematology Sapphire instruments to validate comparable operations and resulting. If an instrument(s) problem is suspected a whole blood evaluation will also be performed. Since specimens are more routinely run in the closed mode that is the mode routinely evaluated. We may, at times, also need to have comparisons run in open mode.

II.	SAPPHIRES:

A.	Cover  lab barcode with blank label

B.	Purge library 24 on Sapphire QC library menu

C.	On worklist:

1.	select specimen type as QC

2.	select library 24

3.	select rack and tube

4.	enter rack number and tube position number (of rack you are using)

5.	select save

6.	do this for each blood in whole blood comparison run

D.	Click on run autoloader and let rack run on instrument (process is the same on 			each Sapphire).

E.	Print out the QC 24 library (if checking open mode enter qc as specimen type and 			library 24 for each entry).

F.	Enter results on “whole blood comparison” table.

G.	Review all results and those from YPB lab, specimen to specimen and mean

average of each instrument.

H.	Evaluate if results acceptable from all instruments, any problems or need to 			calibrate. (Tolerance limits for each parameter can be found in Sapphire

Operators Manual). Resolve any instrument operation problems and redo whole

blood comparison. If operation is acceptable, recalibrate problem parameters.

I.	Sign off on comparison sheet and file in “whole blood comparison” notebook.

III.		HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on September, 2003.

H-2      This procedure was revised by N. Ortoli-Drew 2-2011.

H-3	This procedure was revised by N. Ortoli-Drew 10-2012.