- Principle:

Platelets participate in hemostasis by forming aggregates on the vessel

wall at the site of injury.  Platelet ADP and Thromboxane A2 contribute to this

aggregation.  The process can be duplicated in vitro by adding reagents such as

collagen,  ADP,  epinephrine, ristocetin,  and arachidonic acid to the platelet rich

plasma (PRP).  These agents then cause release of the endogenous ADP and

thromboxane. The resulting in vitro aggregations can then be measured optically.

II.   	Specimen:

The collection, handling, and preparation of patient’s platelets for testing is critical for

optimum results. Venous blood is collected into a siliconized butterfly needle

and plastic syringe immediately prior to testing and making a 9 to 1, blood to

anticoagulant (3.2 g% Na Citrate) mixture. All specimen collection, handling, and

processing must be done in plastic with avoidance of agitation of the specimen. Platelet

rich plasma (PRP) is prepared by gentle centrifugation  (1000 RPM) of the citrate   anticoagulated specimen and adjustment of the PRP to 300,000. The higher the PRP count, the more response the platelets have to activators. Although the PRP count is usually higher than the platelet count in the blood from which it was prepared, aggregation becomes unreliable when the PRP is less than 200,000, meaning the patient’s whole blood platelet count must be over 100,000. Patients with platelet counts >100,000 are rejected for platelet aggregation. Test must be performed within 2 hours of collection.

### III.   	Patient Preparation:

Platelet aggregation is a scheduled test (tues and thurs).  Complete the patient’s platelet  aggregation  scheduling form at the time of scheduling.  The  patient should be instructed to refrain from taking any prescribed or over the counter drugs including herbals and topical pain relief medications which affect platelet function for ten to fourteen  days prior to testing and foods (garlic, spicy, onions) 2 days prior to testing.  If patient is taking medication that affects platelet function, the patient will need to be scheduled at least 14 days after medication stopped. Ideally patient should be free of all medication (prescribed, OTC, herbals and folk remedies). If the physician feels that the medication cannot be stopped, and it is a know platelet antagonist inform physician it will affect test results.

Registration:

Adult outpatients:

Patients registered and drawn in the Smilow phlebotomy lab 4th floor Smilow (Debbie 200-2390) Call or fax schedule reminder to Debbie the day before patient scheduled to alert her aggregation scheduled.

Pediatric outpatients:

Register in the Children’s Hospital, then drawn by Pediatric Specialty Clinic on the second Floor WP.  Fax appointment information to Pediatric Clinic no later than day before appointment.

### IV.   	Required Materials:

A.  Blood collection tubes:

Sodium citrate (2 blue tubes, 1 for PFA and 1 for coags unless assays ordered)

EDTA (1 lavender tube)

Platelet aggregation tube (polycarbonate tube with 2 cc 3.2 sodium citrate)

B.  Reagents

Ristocetin (American Biochem Cal & Pharmaceutical Ltd.)

Epinephrine (Hospital stock, 1:1000)

ADP (Sigma Co.)

Collagen (Bio-Data)

Arachidonic Acid (Bio-Data)

C.   Equipment

4 Channel aggregometer  (Chrono-Log)

50 ml. polycarbonate tubes

Computer software (Chrono-Log)

cuvettes and magnetic stir bars

12 x 75 plastic tubes

Gilson pipettes and tips

chart paper

distilled water

parafilm

### V.  	Collection of Specimens:

2 blue tubes Sodium citrate (for PT,  PTT, PFA100) freeze remainder of plasma

Sodium citrate in polycarbonate tube for aggregation;

2 cc. of 3.24% Na cit,  18 cc. of blood or multiple of 19 dilution

EDTA tube (for CBC, platelet count, smear for platelet morphology), 3 cc. minimum

Any other tubes for additional test requested to be drawn with aggregation.

Verifynow  (aspirin) tube filled to fill mark

- Perform venipuncture using 21g butterfly and plastic syringe.

- Remove butterfly and dispense blood:

18 cc. into aggregation tube

3 to 4 cc. into 4.5 ml. Na citrate tube

3 cc. into EDTA tube

Gently invert tubes to mix anticoagulant and blood. Platelets must be tested less than 2

hours from collection time.

#### VI.     Preparation of Platelet Aggregation Specimen:

A.  Centrifuge specimen ( use the platelet aggregation collection tube) 10 minutes at 1000

rpm in order to obtain platelet- rich plasma (PRP).  PRP has a ground glass appearance.

B.  Transfer PRP to 2-3 plastic 12x75 tubes and cover with parafilm.

C.  Transfer remainder of specimen into 2 plastic 12x75 tubes in order to obtain platelet-

poor plasma (PPP).

- Centrifuge for 5 minutes at 3000 rpm (patients with elevated platelet counts must

be  spun for 15 minutes to assure platelet- free PPP).

E.  Transfer PPP to plastic 12x75 tubes.

F.   Perform platelet count on PRP (use the PLTi count) and adjust PRP count to 300,000

with PPP.  Test can be performed with PRP counts less than 300,000 but results may be

uninterpretable if abnormal.

How to adjust PRP to 300,000 (if PRP count is over 300,000):

actual PRP count x 3 ml. = 3 ml. + #ml. PPP to be added

300,000

example:   600,000 x 3 = 3 + PPP

300,000

6 = 3 + PPP = 3 ml. PPP

Note:  PRP (diluted or undiluted) must equilibrate 30 minutes at room temperature  before testing.

##### VII.    Preparation of Stock and Working Reagents:

A. Ristocetin stored RT stable within expiration date

Reconstitute vial  (100 mg)  with 4 ml. distilled water yielding  working concentration

of 25 mg/ml,  stable 7days at 2-8 °C.

Make a 1:4 dilution of the working concentration to test for Von  Willebrand (VWD)

type IIB.

B. Collagen (Bio-Data Co.)

Reconstitute vial with 0.5 ml. distilled water yielding a working concentration of 1.9 mg/ml.  Reagent is stable for 7 days in refrigerator.

C. Arachidonic Acid (Bio-Data Co.)

Reconstitute vial with 0.5 ml. distilled water yielding a working concentration of 5

mg./ml.  Reagent is stable for 7 days at 2-8º C refrigeration.

D. Epinephrine (Stock bottle from Hospital Pharmacy, 1:1000 dilution)

Store at room temperature.  Verify stock bottle within expiration date. Open Stock reagent stability

within 6 months of open date.

For the working reagent, make the following dilutions:

0.05 ml. stock (1:1000) epinephrine + 1.35 ml. saline = 200uM  dilution

(final test  concentration 10uM)

0.2 ml. of 200 uM + 0.2 ml. saline = 100 uM working solution  of epinephrine

(final test concentration 5uM)

E. ADP (Sigma Co,):

Directions for making the stock dilution:

Stock 4000 uM ADP =

Molecular weight on bottle of ADP x 4 =   ______  mg ADP/ 10 ml. distilled water

100

Dispense stock in 0.1 ml. into plastic tubes, cap, and store at   -70C.  Stable for 6

months. Date the bag in which the tubes are stored. Test this ADP and currently used

ADP in parallel.  Remake ADP if >10% discrepancy of amplitude of aggregation.

For the working solution,   make following dilutions:

0.1 ml. 4000 uM ADP + 1.9 ml. saline = 200 uM ADP dilution

1.0 ml. of 200 uM ADP dilution + 1.0 ml. saline = 100 uM  ADP dilution

(final test concentration 5uM)

0.3m.l of 100 uM ADP dilution + 0.2 ml. saline =   60uM working dilution of ADP

(final test concentration 3uM)

0.2 ml. of 100 uM ADP dilution + 0.3 ml. saline = 40 uM working dilution of ADP.

(final test concentration 2uM)

F. 3.24% Sodium citrate (Na cit): 3.24 grams Na citrate + 100 ml.

distilled water.  Store at room temperature.  Expiration 2 months from date made.

##### VIII.  Preparation of Aggregometer:

A.  Turn on power switches of both units of 4-channel aggregometer.  Allow them to

warm to 37C 2C, approximately 20 minutes (record temperature on log:do not used

channel if temperature not as specified and record problem). If the

computer hard drive is off, press power button. Stirrer switch should be “ON”. Baseline

switches should  point to “1” in order to utilize one PPP for 100% baseline.

B.  On computer screen:

1. access set-up page select

a.  Chrono-Log

b.  Aggro-Link

c.  Aggregometer

d.  Test Procedure

2. At “Procedure Name” access RIS,ARACH,ADP2,EPI5 selection. When

testing the other agonists, choose the “Coll,Ristx4,Collx2,Collx3” selection (you can

create a new panel of agonists, if desired)

3. Next verify the following settings:

a.  Input ports match channels 1-4.

b.  Length of time is 10 minutes.

c.  All channels are selected for optical.

d.  Invert display.

e.  Agonists’ sequence corresponds to sequence in procedure name.

f.  Gain should be “-”.

4.  Select OK.

5.  Next select “Aggregometer”, then “Run Test”.

6.  Select patient’s “last name” and type it. Press TAB and, similarly complete the

demographics. When completed select “OK”.

7.  Set baselines.

C.  Setting Baselines of Aggregometer

1.  Pipette 0.4 ml. PPP into one cuvette. Add a magnetic stir bar into at least six

cuvettes and pipette 0.4 ml. PRP into them.

2.  Place the PPP cuvette into channel 1’s PPP well. (Verify all toggle switches

are set to “1”).

3.  Baseline 100% for each channel.  Press the red “set baseline” switch on first

channel.  Note blue line appears.  Repeat for subsequent channels:  2 -- black,

3-- red,  4 -- green.

4.  Next place a PRP cuvette into each PRP well,  channels 1 - 4.

5.  Baseline 0% the same as 100%.  Press the red baseline switch.  Tracing

should move to 0%.  Note:  Read the scale right side of the graph.

6.  Select “Aggregometer”, then “Stop test”.

7.  Select “Aggregometer”, then “Restart test” (to erase baselining).  Click “OK”

if demographics remains correct.

#### IX.  Aggregation Procedure of Specimen and Control:

After baselines are set, begin testing.

- To existing cuvettes of PRP pipet 20 ul of aggregating  reagents into bottom center of

respective cuvettes..

.          2.   Note timer is visible, so allow all agonists to run until complete aggregation occurs

or for 10 minutes (the last one may take up to 12 or 13 minutes total running time).

When the last cuvette is completed, select “Aggregometer” and “Stop test”.

| Reagent | Working Concentration | Final Concentration (1:20) |
| --- | --- | --- |
| Ristocetin | 25 mg./ml. | 1.25 mg./ml. |
| Ristocetin | 6.26 mg./ml | 0.312 mg./ml. |
| +Collagen | 1.9 mg.ml. | 0.095 mg./ml. |
| #Arachidonic Acid | 5 mg./ml. | 0.25 mg./ml. |
| *ADP | 40 uM | 2 uM |
| ADP | 60 uM | 3uM |
| ADP | 100 | 5uM |
| *Epinephrine | 100 uM | 5 uM |
| Epinephrine | 200uM | 10 uM |

3. Aggregation of PRP is complete when graph tracing plateaus.

### X.  Procedural Notes

- Normal Collagen aggregations are repeated on 1:2, 1:3, 1:5  dilutions of collagen until no aggregation occurs.

- If platelets do not respond or aggregation is followed by disaggregation,  a release reaction has not occurred.  Proceed to test with higher concentrations of ADP or Epinephrine to induce release reaction and permanent aggregation.

- If aggregation is abnormal, the special coag tech should run the Verifynow (aspirin) tube

### XI.   Quality Control

A.  Normal range is >50% aggregation for all agonists.

- A normal control must be run when a patient has abnormal response to any

agonist. Normal control is prepared in the same manner as patient specimen and tested with all (the same) aggregation agonists.

If “Normal” donor is abnormal, another “Normal” must be run.

### XII.  Reporting:

A.  Normal Response of Curves:

Epinephrine may exhibit a biphasic curve, report only as single curve

ADP may exhibit a biphasic curve, report only as single curve

Collagen exhibits a monophasic curve after approximately a one

minute lag response.

Ristocetin exhibits usually a single broad curve, but occasionally

a biphasic curve is seen, report only single curve.

Ristocetin, 1:4 dilution exhibits no aggregation.

Arachidonic acid exhibits a rapid irreversible monophasic

response.

See tracings in Appendix C.

B.  Calculations:

1.Select “Edit”, then “Set start and stop times”.

2.To set start time click arrow to left on graph’s “y-axis” and drag the line(each

channel will show a different color of line) to where the agonist was introduced

in the cuvette and the baseline is reestablished.  Release the mouse when this is

done. If primary colors change on screen, access “Deselect All” option.

3.To set the endpoint, click-and-drag the corresponding endpoint line to the apex

of the aggregation(up to 10 minutes). For biphasic curves, calculate the total

aggregation and then the primary wave only. Subtract this from the total

percent to obtain the secondary wave calculation. Report both results

of any biphasic curves on the report sent to the physicians. Comment any

disaggregation  on the calculation line on the actual report sent to the

physicians.

4.Select next channel and repeat process. When all channels are completed,

select “Done”.

4      5.Select “Edit” then “Compute ...”. If this data is correct select “OK”.

6.Computations should appear on the computer screen. % Aggregation is the

Amplitude.

7.To report comments (for disaggregation):

Select:

a.Edit

b.Test information

c.“Click” in the comment space and enter comments. When completed, select

“OK”.

C. Printing

To print select “File”, “Print”, then “OK”. When report has printed, select

“File”, “Save”(save as the patient’s last name1 and name2 up to 8 characters),

“Close”, and “OK”. Jot the file names on the respective graph pages.

XIII.  Resulting/interpretation:

A.Pltag resulting worklist template, enter % aggregation for each agonist (and normal

control if done)

B. The report and tracings are left for the resident/attending review (do not enter

anything into the LIS)

C. The laboratory director reports the clinical interpretation via Soft Flow  program

The results will be sent via fax to the ordering clinician and the Soft Flow program will

enter the results into the Soft patient report

CONDITION                                                     RESULT/PATTERN

1. Bernard Souliers’s syndrome                      Lack of response to Ristocetin

2. Severe von Willebrand’s disease                 Lack of response to Ristocetin

3. Aspirin “like” defect (can be caused            Lack of response to arachidonic

by aspirin, NSAIDS, congenital                   acid, no secondary wave to

deficiency of cyclooxygenase or                  epinephrine

thromboxane synthase)

4. Platelet storage pool disease                       Lack of or reduced response to

(dense body deficiency)                              collagen - ADP -epinephrine

(though some of these patients are

normal on aggregation and have only

serotonin deficiency)

5. Glanzman’s thrombosthenia                       Lack of response to collagen - ADP -

epinephrine

6. Myeloproliferative platelet dysfunction      Lack of response to epinephrine alone

(This can also be a normal variant)

7. Von Willebrand’s disease                            Aggregation response to low

Type IIb                                                 concentrations of Ristocetin   (0.3 mg/ml)

### XIV.  References:

1.  Triplett,  Douglas:  Platelet Function: Laboratory Evaluation and Clinical

Application.  1978.

2.  Caprino,  L.:  Platelet Aggregation and Drugs.  1974.

3.  Weiss,  Harvey:  Platelets,  Pathophysiology,  and Antiplatelet Drug Therapy.

1982.

4.  Firkin,  Barry:  The Platelet and Its Disorders.  1984.

5.  Lasslo,  Andrew:  Blood Platelet Function and Medicinal Chemistry.  1984.

6.  McPhedran,  Peter, et al:  Yale New Haven Hospital Research

7.  Hoffman,Benz,Shattil,Furie,Cohen,Silberstein:

Hematology Basic Principles and Practice. p.1928.1995. (Appendix A)

8.  Chrono-Log Corporation:  Instruction Manual, “Aggro/Link Control Software”.

9.  Beutler, Lichtman, Coller, Kipps:  Williams Hematology.  1995.

XV.	History:

H-1	This procedure was written by Natalie Ortoli-Drew on 6-2000.

H-2	This procedure was revised by N. Ortoli- Drew on 5-2010.

H-3	This procedure was revised by N. Ortoli- Drew on 2-2011.