I.              PRINCIPLE:

This is a rapid screening test for lactose deficiency using Ames Clinitest reagent

tablets. Reduction of copper ions may be brought about by any reducing substance

and detect sugars other that glucose. The reduction of cupric oxide results in the

solution changing color from blue to green, yellow then brown.

- SPECIMEN:

A.	The stool should:

1.	 Be collected free of preservatives

2.	 Be analyzed within 1 hour of collection

- REAGENTS / MATERIALS:

A.	16 x 100 mm glass test tubes

B.	distilled water

C.	clinitest reagent tablets

D.	clinitest color chart

E.	disposable 1 ml pipets

F.	positive and negative urine control

- QUALITY CONTROL:

Verify clinitest tablet performance by using the positive and negative urine control.

Controls are ordered and resulted in Soft Total Q.C. If control results are not

acceptable, repeat control testing with a new batch or lot of tablets. Patient specimen

testing cannot be performed if control results unacceptable. Notify a supervisor if this

problem condition exists.

- PROCEDURE:

A.	Pipet 1.0 ml of distilled water into a clean 16 x 100 mm glass test tube

B.	Add 0.5 ml of stool specimen and mix well

C.	Transfer 15 drops of stool suspension to another 16 x 100 mm glass test tube

D.	Add 1 clinitest tablet

E.	Read color reaction of the mixture on the 5 drop clinitest chart at 30 seconds

- REPORTING:

Report Results as “reducing substance”:

No color change                < 250 mg/dl

Trace                                 < 250 mg/dl

1+                                      < 250 mg/dl

2+                                         250-500 mg/dl

3+                                      > 500 mg/dl

4+                                      > 500 mg/dl

- SIGNIFICANCE:

Normal individuals exhibit < 250 mg/dl of reducing substances in the stool.

Patients suspicious for lactose deficiency exhibit 250-500 mg/dl.

Patients with lactose deficiency exhibit > 500 mg/dl.

- PROCEDURAL NOTE:

If this test cannot be performed within 1 hour, specimen should be frozen.

- REFERENCES:

  - Bayer Corporation, Modern Urine Chemistry. 1979, revised 1996.

  - Davidsohn and Henry, Clinical Diagnosis and Management by Laboratory

Methods. 18th edition,1991.

- HISTORY:

H-1      This procedure was written by N. Ortoli-Drew on 9/30/2003.

H-2      This procedure was revised by D. Fico on 12/2010.