- PRINCIPLE:

Clinitek Atlas 1000 is a highly automated system for performing macroscopic urinalysis. 	Fresh urine is examined for color, clarity, specific gravity, PH, glucose, protein, blood, 	bilirubin, nitrites, leukocyte esterase, and ketones

- SPECIMEN REQUIREMENTS:

Specimen requirements defined in macroscopic (general) procedure. Urines with small volumes (<2.0ml), mucoid, markedly turbid or grossly bloody must be tested manually or on the Clinitek 500.

- REQUIRED MATERIALS:

A.	Clinitek Atlas 1000 Urine Chemistry Analyzer, Siemens Healthcare Diagnostics

B.        Clinitek Atlas Reagent Pak, Siemens Healthcare Diagnostics

C.	Count -10 Trol I, Myers-Stevens Group, Montebello, California

D.	Count -10 Trol III, Myers-Stevens Group, Montebello, California

E.	Clinitek Atlas Calibrator Set, Siemens Healthcare Diagnostics

F.	Distilled Water

G.	15 ml conical plastic tubes

H.	Bleach

I.	Clinitek Atlas Rinse Additive, Siemens Healthcare Diagnostics

- REAGENTS:

A.	Rinse solution: Make up daily and discard old rinse solution.

2ml Clinitek Atlas Rinse Additive into 1000 ml of distilled water.

Mix and place on instrument.

Rinse the old container with distilled water and let air dry.

- OPERATION:

A.	Daily  Maintenance:

1.	Clean the sample wheel and handle on the carousel with a mild detergent 				and a damp cloth.

2.	Prepare fresh diluted bleach

| Na hypochlorite concentration | ml of bleach | ml of water | Final concentration |
| --- | --- | --- | --- |
| 6.0% | 11 | 1.1 | 5.22% |
| 6.15% | 11 | 1.3 | 5.21% |
| 7.35% | 11 | 1.5 | 5.25% |
| 8.25% | 11 | 1.7 | 5.00% |

3.	Clean the SG well (3 minute process): Performed beginning of each

shift.

a.	Main Menu, press "6"

b.	Place tube of fresh diluted bleach in position 43

c.	Press "Begin".

4.	Replace the rinse bottle and empty waste:

a.	open left door of instrument

b.	check volume of rinse container, discard any remaining reagent if

less than ½ full.

c.	remove waste container and dump contents

d.	rinse waste container with water and discard

e.	add a small amount of bleach to empty waste container

f.	reinstall waste container

g.	install full rinse container

h.	close instrument doors

B.		Daily Quality Control:

1.	Verify lot numbers, expiration date and open stability valid on control 				bottles	(Count-10 Trol I and III, stable through expiration date, open 				stability 5 days)

2.         Order Q.C. for Atlas being used that day in Soft Total Q.C per shift.

Q.C. for confirmatory tests-order   MUA  per shift in Soft Total Q.C.

3. 	Verify at least 2 ml of control in bar-coded conical urine tubes

4.	Bring urine control to room temperature

5.	Place tubes in any position (not 43) on instrument wheel

6.	Press Analyze, then 40 tray / routine

7.	Post Q.C. from Instrument menu. If out of range, repeat control, new order

will be generated.  If still unacceptable, do not run instrument and notify

supervisor. Patient samples are not processed when Q.C. is not acceptable.

Note: any instrument unresolved qc problems must be recorded in problem log in 	addition to informing a supervisor.

- PATIENT SPECIMEN ANALYSIS:

Note: specimens with <2ml, mucoid, or grossly bloody are run on Clinitek 500 or 		          done manually.

  - Thoroughly mix each specimen (plastic conical or preservative tube).

B.     Allow any refrigerated specimens to warm to room temperature

C.     Remove foam from any foamy specimens; analyze any extremely mucoid or 			blood urines specimens on Clinitek 500.  Grossly bloody samples are spun and   			done manually or run on Clinitek 500.

D.     Place specimens in white sample positions (1-40) in sample tray with barcodes 			facing outward:

E.    Stat specimens should be positioned in red sample positions or placed before 			routine specimens

F.     Press ANALYZE, then 40 / TRAY ROUTINE

G.    At end of run, resolve any specimens that did not sample, unread barcode or other

problems.

VII.     RESULTS VERIFICATION:

A.       Refer to specific procedures for confirmatory testing directions.

1. All colored urines auto release.

2. Specific Gravity :  >1.035 – need to do Clinitest and enter results.

3. Protein:   alkaline urine and positive protein confirm with SSA.

4. Urines with positive protein, blood, nitrite, leukocyte esterase, atypical color

and cloudy opacity reflex to a microscopic.

VIII.	TRANSLATION OF ATLAS RESULTS TO LABORATORY 				REPORTING:

A. 	Reported results are not always the same as the values indicated on the color 		chart.

Use the table below to determine the values to report.

| Test |  | Color Chart Value |  | YNHH Report |
| --- | --- | --- | --- | --- |
| pH |  | 5.0 - 8.5 |  | 5.0 - 8.5 (same) |
| Protein |  | Negative |  | Negative |
|  |  | Trace |  | Negative |
|  |  | 30 mg/dl (+) |  | 1+ |
|  |  | 100 mg/dl (++) |  | 2+ |
|  |  | 300 mg/dl (+++) |  | 3+ |
|  |  | 2000 or more mg/dl (++++) |  | 3+ |
| Glucose |  | Negative |  | Negative |
|  |  | 100 mg/dl |  | Trace |
|  |  | 250 mg/dl |  | 1+ |
|  |  | 500 mg/dl |  | 2+ |
|  |  | 1000 mg/dl |  | 3+ |
|  |  | 2000 or more mg/dl |  | 3+ |
| Ketones |  | Negative |  | Negative |
|  |  | 5 mg/dl (trace) |  | Negative |
|  |  | 15 mg/dl (small) |  | Small |
|  |  | 40 mg/dl (moderate) |  | Moderate |
|  |  | 80 mg/dl (large) . |  | Large |
|  |  | 160mg/dl (large) |  | Large |
| Blood |  | Negative |  | Negative |
|  |  | Trace (Non-hemolyzed) |  | Small |
|  |  | Moderate (Non-hemolyzed) |  | Moderate |
|  |  | Trace (Hemolyzed) |  | Small |
|  |  | Small (+) |  | Small |
|  |  | Moderate (++) |  | Moderate |
|  |  | Large (+++) |  | Large |
| Leuk Esterase |  | Negative |  | Negative |
|  |  | Trace |  | Positive |
|  |  | Small |  | Positive |
|  |  | Moderate |  | Positive |
|  |  | Large |  | Positive |
| Nitrite |  | Negative |  | Negative |
|  |  | Positive |  | Positive |
| Bilirubin | Negative |  | Negative |  |
|  | Small (+) |  | Small |  |
|  | Moderate (++) |  | Moderate |  |
|  | Large (+++) |  | Large |  |
| Urobilinogen | 0.2 E.U./dl |  | 0.2 E.U./dl |  |
|  | 1.0 E.U./dl |  | 1.0 E.U./dl |  |

IX.  	CHANGING REAGENT:

- Refer to Clinitek Atlas operating manual (Section 5 5.1-5.6) and Macroscopic Urinalysis procedure; IV. Quality Control, Section C.

- Each time calibration is performed on the Atlas, new tubes must be poured with

Calibrator #1, #2, #3 ( minimum volume of 3 ml) and Calibrator #4 ( minimum

volume of 5 ml). Calibrators of different lot#s are never to be used for instrument calibration. Following the calibration process, the tubes are then discarded.

X.        	CALIBRATION:

- Refer to Clinitek Atlas operating manual (Section 6 6.1-6.5)

- Calibration may need to be performed also after QC fails and troubleshooting complete. Calibration is generally done after PM and/or service. Recalibration may be required due to dry or wet calibration failure. All problems with calibration are logged in Problem Section of maintenance logs for Atlases.

XI.    	TROUBLESHOOTING:

A.   Refer to Clinitek Atlas operating manual (Section 10 10.2-8, Section 11 11.1-                   	11.6

B.     Section 12  12.1-12.11)

C.     Refer to troubleshooting section in Atlas “A” logbook.

D.     ALL instrument problems must be logged in problem log and a supervisor                 		notified if problem cannot be resolved. The instrument can not be run until 			problem resolved and calibration and controls validated.

E.     Atlases are scheduled for PM annually.

XII.            PROCEDURAL NOTES:

- Atlas “A” and “B” are used on alternate days. Atlas “A” (SN #110005) is

generally used on even days. Atlas “B” (SN #110830) is generally used on odd

days .Daily set-up, maintenance and quality control are the same for both instruments.

Note: The carousel sample holder is not interchanged between Atlases.

XII.     REFERENCES:

1. 	Schersten, B. and Fritz, H.: Subnormal Levels of Glucose in Urine. JAMA 	 	201:129-132; 1967.

2. 	   Kark, RM. et al.: A Primer of Urinalysis, 2nd ed. New York: Harper and Row; 	   	   1963.

3. 	   McGarry, J.D.: Lilly Lecture, 1978: New Perspectives in the Regulation of

Ketogenesis Diabetes 28:517-523; May, 1978.

4. 	Henry, J.B. et a1.: Clinical Diagnosis and Management by Laboratory  			Methods, 18th ed. Philadelphia: Saunders; 1991; pp. 396-397,415.

5. 	Siemens Healthcare Diagnostics, Elkhart, IN. Multistix 8 SG. 1994

6. 	Graff, Sister Laurine, A Handbook of Routine Urinalysis, J. B. Lippinctt,

Philadelphia 1983.

7. 	  Instructions for Use and Care PAL Digital Refractometer

8. 	  Count-trol system package insert, Myers-Stevens Group, Montebello,  Calif

9. 	  Atlas Clinitek operating manual 5/1998

XIII.  	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 10-05-03.

H-2      This procedure was revised by D. Fico on 12-20-10.

H-3      This procedure was revised by D. Fico on 3/2011.

H-4      This procedure was revised by D. Fico on 12/2011.

H-5       This procedure was revised by D.Fico on 11/2012.

H-6       This procedure was revised by D.Fico on 11/2013.