I.		PRINCIPLE:

Fibrinogen is a plasma protein that is converted from a soluble protein to an insoluble polymer by the action of thrombin, resulting in the formation of a fibrin clot. The fibrinogen determination is based on a modification of the von Clauss fibrinogen method. Excess thrombin is added to citrated plasma and the time required for clot formation is recorded. The clotting time obtained is compared with that of a standardized fibrinogen preparation.

II.		SPECIMEN:

A.	Type:

Mix nine parts of freshly collected blood with one part of 3.2% sodium citrate

anticoagulant. Invert the tube gently three or four times immediately after

venipuncture to ensure proper mixing of blood and anticoagulant. A syringe or

evacuated tubes (blue top) may be used for collection. If multiple specimens are

collected, the coagulation sample should be the second or third tube collected. If

blood is drawn from an indwelling catheter, the line should be flushed with 5.0

mL saline and the first 5 mL of blood or six dead space volumes of the catheter

discarded or used for other laboratory tests. The citrate concentration must be

adjusted in patients who have hematocrit values above 55%. Specimens that are

clotted, collected in the wrong tube, or have less than the expected  90% fill

(QNS) should be rejected. Hemolyzed specimens may be run and results reported

with a comment.

B.	Handling Conditions:

The whole blood specimen is checked for clot formation by gentle inversion and

observation. Centrifuge the capped blood specimen for 10 minutes at 3000 g.

Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells.

III.		EQUIPMENT AND MATERIALS:

A.	Equipment:

1.	BCSXP  Coagulation Analyzer

2.	Rotors

3.	Waste, rinse and disinfectant containers

4.	BCS/XP Sample Cups

B.	Materials:

1.	Multifibren U Reagent-( OWZG15)

2.	Fibrinogen Standards-(OQVK11)

3.	Control material: Control Plasma P-(OUPZ19)

4.	Imidazole-buffered saline (IBS)

5.	Wash Solution

6.	Preservative-free distilled or deionized (NERL) water

7.	70% isopropyl alcohol disinfectant

C.	Preparation:

1.	Multifibren U Reagent- 50 IU/mL bovine thrombin, fibrin-aggregation

retarding peptide, calcium chloride, stabilizers, buffer and preservative.

- Reconstitute the fibrinogen reagent with 2.0 mL NERL  water

(according to label on vial). Stability after reconstitution: 48 hours on

board (Rack 1-4 with an adaptor)

3.	Fibrinogen Standards (Calibration only) - Six (6) lyophilized standards

(approx. 70, 100, 250, 350, 600 and 800 mg/dL fibrinogen) prepared from

pooled plasma by means of dilution or addition of human fibrinogen and

containing HEPES buffer

4.	Citrol Level 1 Control -Lyophilized preparation of pooled normal plasma

stabilized with HEPES buffer solution (12g/L). Used for Quality Control-

Normal

5.         Control Plasma P-Lyophilized preparation of pooled normal plasma

adjusted to defined factor concentration , stabilized with HEPES buffer

solution (12g/L). Used for Quality Control-Abnormal

a.	Reconstitute lyophilized Fibrinogen Standards with 0.5 mL

distilled water

b.	Reconstitute lyophilized Citrol Level 1 and Control Plasma P with 1.0 mL distilled or deionized water.

c.   	Restopper vial and mix carefully to dissolve (without foam

formation).

d.	Let stand at 15-25C for at least 15 minutes before use.

Stability after reconstitution:   Stable for 4 hours stored at15-25C

Stable for 4 weeks frozen at -20C.

Fibrinogen Standards, Standard Human Plasma, Citrol Level 1, and Control Plasma

P may be quick frozen and thawed before expiration date at 37C. Thawed plasma must

be used within 2 hours when stored at 15-25C. Can only be frozen and thawed once.

D.	Performance Parameters:

Indication of deterioration: no evidence of vacuum in vial upon opening,

difficulty in reconstituting reagent, control values outside of determined range.

IV.		QUALITY CONTROL:

A.	Citrol Level 1–Normal Control

B.	Control Plasma P- Abnormal Control

1.	Controls should be tested once per shift, and with any change or addition

of reagent.

2.	Controls should be run in the same manner as the test samples.

3.	Controls should be within the tolerance limits established for the specific

lot number of control.

4.	Corrective action when tolerance limits are exceeded.

a.	Check reagents, controls and instrument performance

b.	Document actions taken to identify and correct the problem before

reporting any patient data.

V.	PROCEDURE – STEPWISE:

A.	Instrument Set-up and Calibration:

(NOTE: Calibration is done as every 6 months, or as needed, and is not necessary to perform on a routine basis. If you have not performed fibrinogen calibration before, it is suggested that you contact BCS/XP technical support for advice .

Four calibrations must be done for each BCS instrument using the new lot # of 					Multifibren U reagent:

Fib

Fib.570

Fib.low

Fib.low.570

Press Calibration Icon and choose New key.  Select test (eg: “Fib” or Fib.570”, 					“Fib.low”, or Fib.low.570”) which will be highlighted.  Select the lot number of 					reagents and standards on board for the calibration curve

Press Measure curve key to begin curve calibration.

Place the reagent and standards required for the calibration on the analyzer. Click 					OK to 	confirm the message. When the reagent and standards have been 						positioned, the analyzer starts calibration.

Acceptable and valid calibration curves will display a BCS symbol next to the 					curve and date. Unacceptable curves will display an “exclamation point”.

Fib.Multi U curves will be compared to an “Evaluation Tube Curve”. View the 					curve for validity.

Click the Calibration button. Click the Multifibren reference curve in the Curve 					overview column.

Click the Show curve key to view the reference curve. The curve will be 						displayed.

To re-measure one or more points displayed in the graphic, click the point to be 					corrected. The point will be highlighted. Click the Repeat key and close the 					dialog. The point will be re-measured and entered in the graphics.

Rack lanes 1- 4 with an adaptor are used for on board positions of Multifibren U 					reagent, and Rack lanes 4-14 for controls and standards.

# B.	Loading Patient Samples on the BCSXP:

Patient samples can be programmed and loaded on the BCSXP using one of three 					different options:

1.	Auto Host:

a.	Load bar-coded samples into any sample rack with the bar-code 							label visible through the window of the rack position.

b. 	Insert the rack into any open lane.(4-14)

c. 	Push the rack until you feel it stop.

d. 	Open the JOB LIST icon.

e.	The JOB LIST screen displays the samples loaded with a small 							green analyzer next to the sample ID.

The small green analyzer icon indicates the sample is currently 							loaded in the processing area

f. 	The test requests will automatically be picked up from the host.

2.	Assigning a Stat Test Request:

a. 	Prior to assigning test requests, double click on the sample ID 							number

b.	Click on the STAT box.  The box will be marked with a “X”.

c. 	Confirm the entry by clicking OK

d. 	Locate the test requests for each test required for each STAT 							sample listed.

e. 	Double click on the box that intersects between the sample 								identification number and the test column.  A red “X” appears in 							the intersection.  The “X” denotes that a test request for this 							sample is in process.

f. 	Sample processing begins automatically.

3.	Non – Bar – Code Samples without Host (ie: for repeating):

a.	Load the samples into the sample racks into consecutive positions.

b. 	Click on the LOADING icon.

c. 	Click on the number that corresponds to the sample rack (the rack 							number is highlighted)

d. 	Click on the Select rack button (a picture of the sample rack is 							displayed)

e. 	Click on the position number of the sample rack in which the 							sample has been placed.

f. 	Click on the Sample ID  box

g.	Enter the sample identification number.

h. 	At the question “Is the sample STAT?” answer either yes or no.

“Yes” answer will result in a red “X” appearing in the box.

“No”---proceed to the next step.

i. 	To request Individual Assays:

(1.)	Locate the assay box

(2.) 	Click on the assay

(3.) 	Repeat the above steps until all assays are requested.

.

j. 	Click on the INSERT button and the sample ID appears in sample 							position of the rack and the next position of the sample rack is 							highlighted.

k. 	To request another:

(1.) 	Sample in the same rack:   Repeat steps 6-11

(2.) 	Sample Rack:  Click on OTHER RACK button and repeat 							steps 3-11.

l. 	Insert the sample racks into any open lane (4-14) (An open lane is 							one without the LED illuminated or blinking.

m. 	Push the rack until you feel it stop.

n.	Repeat steps 12 and 13 until all sample racks have been loaded.

o. 	The JOB LIST screen displays the samples loaded with a small 							green analyzer next to the sample ID.

The small green analyzer icon indicates the sample is currently 							loaded in the processing area

A red “X” appears in the intersection between the sample ID 						number and the test column and denotes that a test request for the 						sample is in progress.

p. 	Sample processing begins automatically.

C.	Loading and Programming Controls on the BCSXP:

1. 	Place control bottles into the 5 ml reagent rack with the bar-code label on 						the bottle visible through the window of the rack position.

2.	Insert the rack into any open lane from 4-14 (an open lane is one that does 						not have a blinking or illuminated LED.)

3. 	Push the rack until you feel it stop

4.	Click on the Control Journal icon.

The small green analyzer icon indicates that the Control lot number is 					currently loaded in the processing area.

5. 	Double click on the box that intersects between the control identification 						and the test column. A red “X” appears in the intersection and denotes that 					a test request for the control is in progress.

6. 	Control processing begins immediately.

7. 	Repeat step 5 until all assays for each level of control have been requested.

.

D.	Reporting Results:

Patient values autoverify in mg/dL (or can be posted by the tech) in the computer system

Hemolyzed, lipemic, or icteric samples must be commented with the result.

E.	Reference Ranges:

Reference interval determined for fibrinogen (mean value +/- 2.0SD) varies with the lot #, but is approximately 150-450 mg/dL.

VI.		PROCEDURE NOTES:

The overall performance of Fibrinogen testing is dependent on the instrument performance as well as the reagent used. Acceptable variability (imprecision) should be such that the total day-to-day coefficient of variation (CV) of the analytic system is less than 10% on the same lot of control plasma.

A new reference curve should be established every 6 months, or with each change of reagent lot or instrument, or with any consistent deviation from control or proficiency testing limits.

VII.	LIMITATIONS OF THE PROCEDURE:

Results may be affected by fibrin (ogen) degradation products in the patient’s plasma.

Significant levels of FDP may give prolonged coagulation times and cause the test to indicate a falsely low fibrinogen level.

Heparin (up to 2 IU/mL does not interfere with the test.

With some manufacturers’ artificially prepared (e. g. diluted) control plasma for the lower

range, recovery may be above the declared range if the sample matrix is not sufficiently

plasma-like. In such cases, it is recommended that the Control Plasma P be used.

Hemolyzed and lipemic samples should reflex to the 570 tests. These tests may also be

manually requested at the discretion of the operator.

Lipemic samples may need to be ultracentrifuged in Chemistry lab before analysis.

VIII.		REFERENCES:

1.	Siemens Multifibren  U package insert, Dade Behring, Marburg,

Germany, 1998

2.	Fibrinogen Standards package insert, Dade Behring, Marburg, Germany, 1996

3.	Standard Human Plasma package insert, Dade Behring, Marburg, Germany,

January 1999

4.	Control Plasma N package insert, Dade Behring, Marburg, Germany, May 1998

5.	Control Plasma P package insert, Dade Behring, Marburg, Germany, January

1999.

6.	National Committee for Clinical Laboratory Standards. Collection, transport and

processing of blood specimens for coagulation testing and general performance of 					coagulation assays. Approved Guideline, 3rd Edition. NCCLS Publication H21-

A3. Wayne, PA, 1998.

7.	National Committee for Clinical Laboratory Standards. Procedure for the

Determination of Fibrinogen in

Plasma. Approved Guideline. NCCLS Publication H30-A, Vol. 14, No.2. Wayne,

PA, 1995.

8.	Dade-Behring BCS Operator’s Manual.

IX.		HISTORY:

H-1	This procedure was written by N. Kriz on 7-5-2006.

H-2	This procedure was revised by N Kriz on 10/21/2010.

H-3	This procedure was revised by N Kriz on 8/23/2011.

H-4	This procedure was reviewed by N Kriz on 4/24/2012.

H-5	This procedure was reviewed by N Kriz on 6/11/2013.

H-6	This procedure was revised for the BCSXP by N Kriz 11/5/13.