- PURPOSE:

The quality control (QC) program consists of both internal and external Quality Control. 	The internal QC consists of commercial and in-lab prepared materials, which are used to 	verify the accuracy and precision of our test procedures and instrument moving average 	programs. External Quality control consists of Proficiency Testing materials obtained 	from the CAP, Streck, and/specimens submitted to another laboratory.

- SCOPE:

All technical staff who perform instrument set-up, analyze specimens, and validate/report results must participate in the quality control program. Quality control specimens must be analyzed and acceptable before patient specimens’ area analyzed.

# INTERNAL QUALITY CONTROL

### I.	CELL COUNTERS

## A. 	Performing Quality Control

1.	Commercial control:

Once each shift three levels of commercial cell control (cbc/plt/diff/retic) 			are analyzed, any control which is out of range should be repeated one 			time. If at least two of the three levels are within the acceptable range for 			all parameters the run is acceptable. If more than one sample is out of 			range then testing will cease on that instrument and a supervisor must be 			notified (all problems must be logged in instrument problem logs) and

instrument not be used for patient specimen analysis.

2.	Patient References:

Each shift two patient samples are selected; one normal, one abnormal. 				The samples are analyzed 2 times, the results reviewed to assure 					reproducibility, and then the mean values set as the target. These patient 				references should be run every 2 hours, after reagent changes or if 					instrument problem suspected. Acceptable ranges are posted in the 					computer. If at any time a sample has a result outside of acceptable 				range it should be repeated once. If the repeat is still out, testing should 				stop and a supervisor should be notified and 	problem / resolve entered in 				instrument problem log. (refer to Sapphire operators manual for definition

of normal and abnormal references, establishing patient reference

libraries).  If an instrument is not in use or in service references do not

need to be run but status of instrument must be noted in instrument log.

3.	Whole blood comparisons: Monthly 6 patient bloods are analyzed on each 				cell counter. The results means calculated and all instrument results 				compared to assure consistency of results on the cell counters.

Comparison bloods and mean of the 6 bloods must meet the established

tolerance ranges.

4.	Moving averages: Sapphires have moving average programs enabled to

continually monitor x-b, wbc, plt, retic results on an ongoing basis. The

operator is alerted of any deviation from established ranges. The moving

average results are also routinely reviewed daily by a supervisor for trends

or problems.

5.        Commercial calibrator: Performed at least every 6months, after major

instrument service, when indicated by QC results, to verify instrument

unexpected shifts/trends.

#### B. 	Review of Quality Control:

- At least weekly (usually daily) a supervisor will review the daily controls to observe patterns or consistent problems with the commercial controls or patient controls or instrument operation. Monthly the QC library  results will be printed, purged  and  signed off by a supervisor.

2.	New lot of commercial control is “overlap” assayed prior to activating the

new control for use. Prior to performing the overlap assay a supervisor

will verify that the instruments are performing satisfactorily and current

lot of  commercial control recovering with posted tolerance ranges.

Overlap assay of new controls with the existing controls is utilized to verify SD and control target mean of manufacturer or if  “in-lab” adjustment of target mean is necessary.

- Peer group statistics: Once a control lot has expired the,

control libraries for that lot are  printed and the 2 months control data is

submitted to Streckstats

### II. 	COAGULATION

##### A.	Performing Quality Control

1.	PT, PTT, d-dimer and Fibrinogen:

Once each shift of use or whenever a reagent is changed, normal and 				abnormal controls are run on the BCS. Results must be within the 					acceptable ranges posted on the instrument. If a value is outside of 					acceptable range it should be repeated once. If still out a supervisor should 			be notified, testing must not continue.

2.	Special Coagulation:

All special coagulation tests have at least two levels (abnormal and 				normal) which must be run with each batch. Results must be within the 				stated ranges for patient results to be reported.

##### B. 	Review of Quality Control:

1.         PT, PTT, d-dimer and Fibrinogen controls are reviewed from the QC libraries at 	least once per week by the Coagulation Specialist. Monthly, the computers	control libraries are printed out and reviewed by the Coagulation Specialist. Statistics are sent to Siemens QAP for computation.

2.          Special coagulation controls are reviewed by the Coagulation Specialists from the QC libraries on BCS or STA compact.

### III. 	URINALYSIS

###### A.	Performing Quality Control:

1.	Two levels of urine control are run each shift of use. The controls must be 			run on Atlas and/or Clinitek 500 if the instrument is to be used for patient 			testing. If an instrument is not used, no controls need to be run but this 			should be noted on the control log. Commercial calibrators and controls 			are run after reagent change. 	2 previously tested urines (normal and 			abnormal) are also run if reagent is a different lot and results reviewed for

satisfactory performance.

- Clinitest. Two levels of controls are run only at the time of 				                       running the test(s) with a maximum of once/shift.

- U-F1000i- Two level of controls are run if used for patient testing.

4.	At least 2 times per year patient urine specimens are run on both Atlas and 		Clinitek 500 and results compared for consistency.

###### B.	Review of Quality Control:

- Monthly the QC file will be signed off by a supervisor.

IV.	 BODY FLUIDS:

A. 	Once each day saline and albumin background are evaluated under the 				microscope for bacteria. If bacteria are discovered, the affected reagent is 				discarded.

B.	Once a shift Cell chex manual control level 1 and 2 are counted. Results must be

within 	established target range. Once a shift Cell chex auto control level 1,2 and

3  are run on the fluid Sapphire. Results must be within established target range.

In  addition a peripheral blood reference is run (a new one each shift) to assure

satisfactory operation of the Sapphire

Both manual and automated commercial controls must be overlapped with current

lot prior to usage.

V. 	MISCELLANEOUS TESTING:

All quality control is discussed in the individual procedures.

All quality control runs are logged whether acceptable or unacceptable. All unacceptable 	quality control must have problem resolved and quality control material(s) retested. 	Patient analysis cannot be performed until quality control analysis is satisfactory. A 	supervisor must be informed when second analysis of a control material is unsatisfactory.

EXTERNAL QUALITY CONTROL AND PROFICIENCY TESTING PROTOCOL

I.		PROCEDURE:

Proficiency testing samples are performed in exactly the same manner as patient

tests. They are integrated into routine round the clock patient testing and are performed by whichever technologist is assigned to a given workstation at the given time. CAP proficiency specimens are analyzed by all staff on a rotational basis.   Proficiency test specimen and/or  results  are not shared or referred to other Hematology laboratories.

A limited number of CAP surveys (ex stem cell, hgb identification) have multiple laboratory  components  to the survey and these are shared as per CAP survey instructions.

- The CAP proficiency kits are delivered to the lab and logged on the CAP survey board by the receiver with all pertinent information. The kit will be checked for all satisfactory materials. The survey kit is then stored according to CAP prescribed conditions, either cold room or room temperature. The supervisor of the area for the survey initiates the coordination of the handling, processing of the specimens, timely testing and obtaining the results before due dates. All testing will be performed according to laboratory policies and procedures and in compliance with CAP kit directions.

- The assigned technologist will initiate testing by proper warming and mixing techniques

and  will perform the analysis.

- After completing the testing, the technologist will provide the supervisor with    worksheets and completed CAP result forms. Result forms and worksheets will be reviewed by the supervisor of the area. Only tests that require a supervisory review normally will be checked for accuracy. If the tech results are incorrect, the supervisor will exercise an educational and/or retraining with the tech and report the corrected results. The same process is followed if tests need a Lab Resident/Attending review. If the results from the testing technologist are deemed incorrect, a review with the tech and supervisor will take place with all proper documentation filed with the survey. In both these circumstances, documentation will be filed in the testing technologist’s competency file.

- Signatures from the testing technologist and the Laboratory Director are required on completed forms. A second designee will review recording of the results before being submitted to CAP and the coordinator will email results to CAP and retain all paper reports for the lab’s  proficiency files.

- The results of the proficiency report are reviewed by the Lab Medicine Chair, Compliance Coordinator, Lab Director, Lab Manager and the supervisor of the specific area upon receipt.

F.	All data (raw and completed), reports, and action reports are retained for at 			least two years.

Clinical Hematology performs the following CAP Proficiency Testing Surveys:

| BP | Blood parasite | CGL | Coagulation, limited |
| --- | --- | --- | --- |
| FH3P | Hematology automated differential | RT | Retic |
| HG | Hemoglobinopathy | OCB | Occult blood |
| CMMP | Clinical microscopy | CMP | Clinical microscopy |
| DSC | Dipstick confirmatory | PF | Platelet function |
| PF1 | Platelet function | SCM1 | Special microscopy |
| DCM2 | Special Microscopy | G6PDS | G6PD |
|  |  | CGEX | Coagulation extended |
| CGS1 | Coag special testing | CGS2 | Coag special testing |
| CGS4 | Coag special testing | ESR | Sedrate |
| BFC | Crystals | URC | Crystals |
| HFC | Hemocytometer fld | LN9ER | Hematology linearity |
| FCFS | Fecal fat | FlAC | Fecal lactoferrin |
| SCP | Stem cell processing | RMAL | Rapid malaria |
| FLD | Body fluid | PF | Platelet aggregation |
| PF1 | Platelet function | BFC | Crystal |
| VLBS | Virtual smear |  |  |

In addition:

CBC/plt Linearity From Streck 2x/year

Tests where no external proficiency testing is available- tests are performed in a double blinded analysis by 2 technologists and final results recorded and compared for consistency.

II.       INTERPRETATION OF CAP Survey RESULTS:

The CAP results are reviewed by the Chairman of Laboratory Medicine, Laboratory 	Director and Laboratory Manager. The Laboratory Manager will coordinate the 	investigation and correction of problems if the testing is unacceptable. Problems that may 	occur are: transcription errors recorded onto the answer sheet, instrument malfunctions, 	reagent problems, and debatable CAP interpretations. Instrument controls, calibrations, 	reagents, and precision are examined when results show a bias or trend. Repeat testing will be done with the original CAP specimens by the original testing technologist. If

discrepancies  are  discovered, an investigation is performed, discoveries are recorded, and an action plan is developed. If the original samples are not available for repeat testing, the lab will contact CAP to order repeat specimens. In the event, CAP is unable to provide new specimens. Positive and negative patient samples for the specific test in question will be tested by the original tech in a blinded fashion. The specific supervisor of the area will observe that the laboratory procedure was followed correctly by the technologist and that the technologist was competent to perform the testing and report the results. This action plan is then filed with all proper documentation regarding the remediation of the contributing problem and re-education of the tech to the proficiency test response. This process occurs for CAP surveys(graded and ungraded) and non-CAP assessment testing is evaluated by the Laboratory Director or Designee.

.

III.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 10-10-97.

H-2	This procedure was revised by N. Ortoli-Drew on 7-2005.

H-3	This procedure was revised by N. Ortoli-Drew on 4-2007.

H-4	This procedure was revised by N. Ortoli-Drew on 7-2007.

H-5	This procedure was revised by N. Ortoli-Drew on 10-2008.

H-6       This procedure was revised by N. Ortoli-Drew 4-2011

H-7       This procedure was revised by N. Ortoli-Drew 12-2012

H-8       This procedure was revised by N. Ortoli-Drew 6-2013

H-9       This procedure was revised by N. Ortoli-Drew 9-2013

H-10     This procedure was revised by D.Fico  6/2014.