- PRINCIPLE:

The sulfosalicylic acid method for urine protein is used to confirm the dipstick or 	Clinitek Atlas method when the protein is positive and the PH is alkaline. Highly 	buffered alkaline urines may give false positive results if the buffer system of the 	reagent pad is overcome. The sulfosalicylic acid method is a precipitation technique for 	semiquantitative determination of protein which is unaffected by urine PH.

- SPECIMEN REQUIREMENTS:

Fresh random urine tested within 2 hours of collection. If testing will be delayed 	specimen should be refrigerated. Specimen must be at room temperature for testing.

- QUALITY CONTROL:

On each shift a sulfosalicyclic method is performed on negative and positive urine controls. See Urine Soft Q.C. Procedure. If control results are not acceptable, patient results must not be reported and a supervisor notified.

- REQUIRED MATERIALS:

3% Sulfosalicylic acid ( 3 gr. sulfosalicylic acid dissolved and diluted to 100 ml with distilled water, stable indefinitely at room temperature)

Abnormal Urine Control, Count-IO-trol III, Myers-Stevens Group

Normal Urine Control, Count-IO-trol I, Myers-Stevens Group

Transfer pipets

Test tubes

# PROCEDURE:

- Mix equal parts of 3% sulfosalicylic acid and urine sample or control well.   Let stand at room temperature for 10 minutes.

B. 	Read the resultant turbidity against the native urine turbidity/clarity of the urine 			being tested.

C.	 Interpret as follows:

| Result | Appearance | Approx. Protein Conc. |
| --- | --- | --- |
| Negative | No turbidity / no increase in turbidity. | <5 mg/dl |
| 1+ | Perceptible to distinct turbidity but no granulation. | 10 - 50 mg/dl |
| 2+ | Turbidity with granulation but no floculation. | 200 mg/dl |
| 3+ | Turbidity with granulation and floculation. | 500 mg/dl |
| 4+ | Clumps or solid precipitate | 1000 mg/dl |

# EXPECTED RESULTS:

In most cases, this method will correlate with the dipstick or Clinitek Atlas results. In some cases, when a urine is alkaline the dipstick method will give a false positive and this method will yield a true negative. Report the SSA method in the computer. The associated comment will read as follows:

Sulfosalicyclic Acid

Positive result confirms protein is present.

Negative result confirms protein is absent.

# LIMITATIONS OF PROCEDURE:

If the urine is cloudy the sample should be centrifuged before performing the 	  		sulfoasalicylic acid method.

# REFERENCES:

## 1.	Modern Urine Chemistry, Bayer Corporation, Tarrytown, NY, 1996

| Review | Revised | Reviewer |
| --- | --- | --- |
|  | 7/2005 | N Drew |

VIIII.	  HISTORY:

H-1	This procedure was written by Natalie Ortoli-Drew on 09-2003.

H-2	This procedure was revised by Natalie Ortoli-Drew on 7-2005.

H-3 This procedure was revised by D.Fico on 9/2011.

H-4 This procedure was revised by D.Fico on 7/2014.

Appendix C

3%   Sulfasalicylic  Acid For Urine Proteins:

Reagent is used to test urine proteins. This method is used when protein can’t be read on

Multistix  or  to confirm presence of protein in alkaline urines.

Preparation:       3 gm. Sulfasalicylic acid

Dissolve in 100 ml distilled water

Stability:            Indefinitely, at Room Temperature

Reference:         Kingsbury and Clark method

Brays Clinical Laboratory methods

Reagent needs to be properly labeled.