- PURPOSE:

A.	To define what test will be ordered on HPC specimens

B.	To describe how specimens are to be accessioned

C.	To define where specimens should be forwarded

II.	POLICY:

## Specimens on HPC will be delivered to Department of Laboratory Medicine

Specimen - Receiving / processing area. Specimens will consist of 1 tube for flow 	cytometry, 2 tubes (one concentrate and one 1:5 dilution) for main hematology. Each 	specimen (s) must be accessioned, appropriate comments entered and forwarded to 	testing areas.

III.	PROCEDURE:

A.  	Each specimen will be labeled with patient information and accompanied by a

requisition slip (requisition slip is to be scanned in by specimen processing).

B.	Dependent on the sample type the following tests are ordered

- HPCA: Hematopoietic Progenitor Cells, Apheresis

- HPCCB: Hematopoietic Progenitor Cells, Cord Blood

- HPCM: Hematopoietic Progenitor Cells, Marrow

C.        1 tube of concentrate and 1 dilution x5 is delivered to Sapphire area

D.   	Specimen analysis:

- Scan the order number of the patient sample in Open tube mode

- Run 1:5 dilution specimen through Sapphire

- Print result sheet

- Multiply WBC and platelet counts by 5 with a calculator

- Press F4 and click save on instrument page to cancel all results

- Enter results in Lab results tab: enter WBCI and PLTI and verify all.

(man diff is automatically added)

- Forward undiluted specimen and result slip to smear area

- Make smear from undiluted specimen

- Perform 200 cell differential from smear of undiluted specimen

- Verify results into computer unless abnormal cells seen and the differential needs to be reviewed by supervisor.

IV.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 6-4-2003.

H-2      This procedure was revised by N. Ortoli-Drew on 3-27-2010.

H-3      This procedure was revised by N. Ortoli-Drew on 5-2011.

H-4      This procedure was revised by N. Ortoli-Drew on 4/2012.

H-5	This procedure was revised by Elizabeth Degano on 12-17-12

H-6	This procedure was revised by Elizabeth Degano on 6-20-13