- Purpose:

To provide review and documentation of comparison of automated vs. manual

differentials.  To formally document each shift daily that automated differential data id indeed correct and correlated with manual differential.

- Policy:

Each shift a technologist performing differentials will record auto and manual differential on 1 normal and 1 abnormal (flagged) specimen. These differentials must agree within the noted tolerance limits.

Tolerance limits seg 8, lymph 8 mono 6, eos 1.5

- Procedure:

  - Perform manual differential on 1 normal and 1 abnormal (flagged) specimen

  - Record auto differential and manual differential for these specimens in the comparison log.

  - Review that results agree within the specified tolerance limits noted on the bottom of the log sheet

  - If results do not agree have another tech perform a second manual differential and review

  - If discrepancy still exists, repeat specimen on Sapphire and make a new smear. If still discrepant inform a supervisor and do not use the Sapphire that produced the automated differential

IV.	HISTORY:

H-1     This procedure was written by N. Ortoli-Drew on 9-24-2003.

H-2     This procedure was revised by N. Ortoli-Drew on 1-2011.

Appendix A