I.	PRINCIPLE:

Routine coagulation screening tests are performed through photometric detection of changes in sample optical density. Results are generated for prothrombin time (PT), International Normalized Ratio (INR), activated partial thromboplastin time (APTT), fibrinogen, and d-dimer.

The INR is calculated from the patient PT divided by the geometric mean of the PT reference range, using as an exponent the International Standardization Index provided by the reagent manufacturer. (Refer to specific procedure for Set up of INR on New Lot Number of Reagent for additional information.)

In addition, d-dimer testing is performed by analysis of the change in turbidity of a sample 	upon addition of polystyrene particles coated with monoclonal antibody to human d-Dimer 	fragments. (Refer to specific d-Dimer procedure for additional information.)

IMPORTANT NOTES:

- The BCSXP must ALWAYS be run with the BCSXP, Autoassistant, and Easy Talk icons minimized at the bottom of the screen.  If the Autoassistant is not running, tests will not reflex (eg, to long tests, or to 570 wavelength); if Easy Talk is not running, the instrument will not communicate with Soft for autorequest of tests nor will it autopost results.

- When adding a new bottle of reagent (as, for instance, when the old bottle is empty or expired) the old bottle must be taken off and the new bottle must be placed in either a different rack or in a different position in the same rack. When a rack is ejected, the BCSXP retains the memory for one hour of the vials which were on the instrument; if a new bottle is placed in the same position as the old bottle, the instrument will NOT recognize it as a new bottle, and will retain the same expiration and volume information that it had for the used bottle.

- There are passwords in place for the BCSXP and these must be used when the instrument is shut down and brought back up.

  - After a complete shut down (software and instrument) bring the instrument back up by turning on the switch first, then clicking on the BCSXP icon. A password is needed, and the User Name will be filled in already BCSXPADMINISTRATOR. The Password for this log on is

BCSXP Admin and is case-sensitive.

  - After a restart of the software only, or after logging on as above, another username Labmanager and password BCS will be used. As above, the user name and password are case sensitive.

- REAGENTS:

NOTE: After opening any vial to place onto the instrument, label that vial with the expiration date, referring to the stability information provided here.  Discard reagent when expired.

A. 	Siemens InnovinR—lyophilized recombinant human tissue factor, and calcium in buffer-- for performance of the Prothrombin time. For use, add 10 ml NERL water, allow 15 minutes to dissolve. After reconsitution, stabile for up to 5 days in refrigerator, 48 hours on board the BCSXP (track 1 -4).

B.	Siemens ActinR FSL--purified soy and rabbit brain phosphatides in ellagic acid with added buffer, stabilizers, and preservative-- for performance of APTT. Ready for use; after opening, stabile 5 days in refrigerator; 48 hours on board the BCSXP (track 1 -4).

C.	Siemens CaCl2 0.025 M--for performance of APTT. Ready for use; after opening, stabile 8 weeks in refrigerator.  Stabile 48 hours on board BCSXP (track 1 -4).

D.	Siemens Multifibren U-for fibrinogen determination. Reconstitute each vial 				with 2 mL NERL water and allow to sit for 15 minutes at room temperature;  swirl 			to mix 	and place on BCS; run fibrinogen testing as described in separate procedure. 			Stabile 48 hours on board BCSXP (track 1 -4 with an adaptor).

E.	Imidazole-buffered saline (IBS)—diluent for Multifibren U fibrinogen testing. Place 			about 5 ML (approximately half a bottle) of prepared IBS in a large brown reagent 			bottle which carries a “Saline” barcode label. Replace each shift, or if necessary  			prior to that. Stabile 8 hrs on board BCS (track 1-4.

F.	Siemens Citrol Level 1 (normal) and Citrol Level 3 (abnormal) controls. Reconstitute each with 1.0 ml NERL water and let stand 15 minutes; swirl to mix before loading; after opening, stable 8 hours refrigerated after running. Stabile 8 hours on board  the BCSXP (racks 4-14 ).

G.	Siemens Control Plasma P (abnormal) for Multifibren U low fibrinogen and derived fibrinogen determinations. Reconstitute with 1.0 mL NERL water and let sit 15 minutes; swirl before loading. Stabile after opening 8 hours refrigerated after running; 8 hours on board the BCSXP (racks 4-14) .

H.	Innovance d-Dimer reagents. Prepare as needed and place on BCS instrument according to d-Dimer procedure described elsewhere.  Stable 72 hours(3 days) on board BCSXP .  D-dimer reagent, supplement and buffer should be placed in track 1-4 with adaptors in the wells; d-dimer diluent should be placed in track 5-14.

I.	Innovance d-Dimer Control 1 and Control 2. Prepare and place on BCSXP instrument in track 5-14 according to procedure. After running controls, and assuring that results are within acceptable range, re-cap vials and store in refrigerator. Handled in this way, controls are stable for 72 hours (3 days) in refrigerator at 2-8C.

J.  	Siemens BC Thrombin reagent—reconstitute 1 vial with 5 mL of buffer provided in kit.  Place on BCSXP in track 5-14 and run test as described in a separate procedure. Stabile 5 days when capped and stored in refrigerator at 2-8C after use.

K.	Siemens Citrol 1 (normal) for thrombin time control—reconstitute 1 vial 				with 1.0 mL of NERL water. Allow to dissolve completely and run according to 			procedure described elsewhere. Stabile 8 hours on board BCSXP (racks 4-14).  			Also run abnormal control, referring to Thrombin Time procedure for directions.

L.	Siemens washing solution—ready for use. Place fresh bottles in left and 				center positions as needed. Stabile on board BCSXP for 5 days, unless problems 		occur with patient or control results.

M.	Disinfectant solution—70% Isopropyl alcohol ready for use.  Stable 3 months at room temperature in canister provided.

N.	Distilled water (from distilled water tap in lab)—for diluent/ rinsing. Change 			daily during morning set up.

O.	Siemens BCSXP cuvette rotors—load into RIGHT rotor store only.  Load only 			5 rotors at a time.

P.	10% bleach solution for cleaning counters/surfaces—Prepare in with distilled 			water from the DW tap. Make fresh daily for counter cleaning.

III.	PROCEDURES

A. 	MORNING SET-UP:

- Perform Daily Maintenance as described in separate procedure; include a                              Shift Change on each instrument before running controls. After the shift change

  - delete the previous day’s patient results

  - shut down the BCSXP software by going to File, then Quit; Select Delete for patient results, and de-select Disinfect (if you don’t deselect, the instrument will do another shift change); select Continue.

2. 	Prepare and load routine reagents as needed:

a. 	Siemens CaCl2 (0.025 M)

b.	Siemens Innovin

c.	Siemens FSL

d.         Multifibren U

e. 	IBS

f. 	Innovance d-Dimer reagents

g. 	Innovance d-Dimer control 1 and control 2

3. 	Reconstitute reagent(s) before controls  and allow to sit at room temperature 			for 15 minutes; swirl to mix just before loading onto the BCSXP.

4. 	Reconstitute 1 vial each of Citrol Level 1 (normal control), Citrol Level 3 				(abnormal control), and Control P (abnormal fibrinogen control) with 1.0 ml 				NERL water. Allow the vials to sit for 15 minutes before placing on BCSXP 				(in racks 4-14); swirl to mix just before loading onto the BCSXP.

5.   	Run controls; if control values are within acceptable limits, post the results into Soft referring to Soft procedures as necessary.  If QC results fall

outside acceptable limits (blue exclamation point appears next to result), proceed as follows:

a.	Perform an intensive Wash of both pipettors (System, Pipettors, 				change wash intensity to Intensive, check both left and right, then 				Start). Recheck pipettor dispensing to assure that the stream of 				liquid dispenses straight down into washing well.

b. 	Rerun controls; if still out, prepare new controls with fresh 					NERL water.

c. 	Rerun controls. If controls still out of range, change appropriate 					reagent(s).

d.	Rerun controls. If controls still out of range, change washing 					solution(s) if not changed during set up.

e.	Rerun controls, If controls still out, notify supervisor and/or contact Siemens instrument support hotline for further help. (Toll free number is posted on each instrument. Also have ready the serial number of the instrument with the problem. )

Normal and abnormal controls must be run once per shift, and whenever reagents are added or changed.

B.	 RUN PATIENT SAMPLES:

1.	Insert tubes into racks with barcodes visible in the opening. Place racks 				into lanes to be brought into the instrument.

2.	Tests will be requested automatically by LIS computer. Certain samples (eg: 				those with hemolysis, lipemia, PTT “no clot” or PT >60.0 sec) will 					reflex to appropriate programmed repeat).

3.  	Completed results will autorelease according to programmed parameters, or will appear for operator verification on Instrument Menu. Compare results with previous ones, repeating, checking for clotted/QNS where appropriate. Also refer to separate Autoverification Guidelines for complete information.

4. 	If a sample is lipemic or icteric, it is possible that the light intensity will not be sufficient to accurately sense a clot formation; these samples will automatically repeat using a different wavelength (PT 570 or PTT 570). In the case of a lipemic sample, an invalid result may be obtained even on the 570 repeat. In that case, the sample should be ultracentrifuged in the chemistry lab and rerun. If a valid result cannot obtained, cancel the test and notify a caregiver to send a repeat sample.

- Samples which are hemolyzed may be run (as long as not QNS or clotted), and may also reflex to a 570 repeat; those results should be reported accompanied by a comment noting the hemolysis and stating that results may be affected.

- EVENING SHIFT

  - Eject racks from one instrument and check volumes and expirations; replenish reagents as needed. Continue running patient samples on the other BCSXP.

  - Replace IBS

  - Prepare controls as for day shift set up.

  - When controls are ready to run, gently mix before placing onto instrument.

  - Request controls to be run, and monitor results when done. Assure that control values are within acceptable ranges before continuing with patient samples.

  - Check and initial evening maintenance logs

  - Run patient samples.

  - Set up the second BCSXP in the same manner.

IV.	HISTORY:

H-1	This procedure was written by N. Kriz on 8-14-2006.

H-2	This procedure was revised by N. Kriz on 7/26/2010.

H-3	This procedure was reviewed by N Kriz on 2/25/2011

H-4	This procedure was reviewed by N Kriz on 1/31/2012

H-5	This procedure was reviewed by N Kriz on 3/11/2013

H-6	This procedure was revised by N Kriz on 11/5/2013 for 2014.

H-7	This procedure was reviewed by P Bahel on 08/19/2014.