I.	PRINCIPLE:

Polybrene is a specific heparin antagonist that will shorten the APTT when the prolongation 	is due to the presence of heparin. Polybrene will not correct an APTT prolonged by liver 	disease, factor deficiencies, or 	coumadin.

II.	PROCEDURE:

A. 	Place 0.5 ml of patient plasma into sample cup.

B. 	Add 0.05 ml (50 ul) of 1.5 ug/ml final concentration polybrene (in refrigerator).

C. 	Run a PTT on the mixture on the BCSXP in the usual manner.

D.	If there is partial or no response of the PTT to 1.5 ug/ml of polybrene, add an 			additional 0.05 ml (50 ul) of 	6.0 ug/ml polybrene to the same 0.5 ml of patient 			sample and repeat the PTT.

(If original PTT is > 50.0 seconds, the 6.0 ug/ml concentration may be used directly.)

E. 	Make an appropriate comment against the PTT using canned comments for correction or non-correction of the PTT .

Report result of whichever concentration shortened the PTT the most.

F.   To prepare  POLYBRENE  (for heparin neutralization  of PTT):

a. 	In a 100-ml volumetric flask, dissolve 66 mg of polybrene 						(hexadimethrine bromide-hydroscopic, kept tightly sealed in 						refrigerator) in 100 ml of distilled water. Stir to mix well. 						Concentration of this solution is:

66 mg/ 100 ml ~ 66,000 ug/ 100 ml ~ 660 ug/ ml

b. 	Dilute 660 ug/ml 1:10 in DW:

10 ml of 660 ug/ml + 90 ml DW = 66 ug/ml (stock)

c.	Dilute 66 ug/ml 1:4 in DW:

25 ml of 66 ug/ml + 75 ml DW = 16.5 ug/ml (stock)

d.	For use:

0.5 ml patient + 50 ul 66 ug/ml=   6.0 ug/ml

plasma stock            final concentration

0.5 ml patient + 50 ul 16.5 ug/ml= 1.5 ug/ml

plasma stock             final concentration

NOTE: Order hexamethrine bromide from Sigma (vendor #9304), 				  	Catalogue # H9288,   5 gm vial, line account 513340, price 15.50/ 5 gm vial.

III.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-16-2004.

H-2	This procedure was revised by N Kriz on 2/2/2010

H-3	This procedure was revised by N Kriz on 2/3/2011

H-4	This procedure was reviewed by N Kriz on 12/1/2011

H-5	This procedure was reviewed by N Kriz on 12/19/2012

H-6	This procedure was reviewed by N Kriz on 12/27/2013