- PURPOSE:

A.	To define what test will be ordered on HPC specimens

B.	To describe how specimens are to be accessioned

C.	To define where specimens should be forwarded

II.	POLICY:

## Specimens on HPC will be delivered to Department of Laboratory Medicine

Specimen - Receiving / processing area. Specimens will consist of 1 tube for flow 	cytometry, 2 tubes (one concentrate and one 1:5 dilution) for main hematology. Each 	specimen (s) must be accessioned, appropriate comments entered and forwarded to 	testing areas.

III.	PROCEDURE:

A.  	Each specimen will be labeled with patient information and accompanied by a

requisition slip (requisition slip is to be scanned in by specimen processing).

B.	Dependent on the sample type the following tests are ordered

- HPCA: Hematopoietic Progenitor Cells, Apheresis

- HPCCB: Hematopoietic Progenitor Cells, Cord Blood

- HPCM: Hematopoietic Progenitor Cells, Marrow

C.        1 tube of concentrate and 1 dilution x5 is delivered to Sapphire area

D.   	Specimen analysis:

- Scan the order number of the patient sample in Open tube mode

- Run 1:5 dilution specimen through Sapphire

- Print result sheet

- Multiply WBC and platelet counts by 5 with a calculator

- Press F4 and click save on instrument page to cancel all results

- Enter results in Lab results tab: enter WBCI and PLTI and verify all.

(man diff is automatically added)

- Forward undiluted specimen and result slip to smear area

- Make smear from undiluted specimen

- Perform 200 cell differential from smear of undiluted specimen

- Verify results into computer unless abnormal cells seen and the differential needs to be reviewed by supervisor.

F. 		For specimen analysis for CAP Surveys only

1. Perform manual differential and report all WBCs cells seen including total nRBCs /100 WBCs from the smear (do not use the Sapphire result)

2. To obtain the White Blood Cell Count: use the formula for the WBC from the Abbott Sapphire as:

WBC = (Sapphire absolute WBC + Sapphire Absolute nRBC) – (Sapphire absolute WBC x (total number of nRBCs seen/100 WBCs on smear))

3. To obtain the total nucleated cell count, use the following calculation: Total nucleated cells = Sapphire Absolute WBC + Sapphire absolute WBC x (Total nRBC/ 100 WBC)

IV.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 6-4-2003.

H-2      This procedure was revised by N. Ortoli-Drew on 3-27-2010.

H-3      This procedure was revised by N. Ortoli-Drew on 5-2011.

H-4      This procedure was revised by N. Ortoli-Drew on 4/2012.

H-5	This procedure was revised by Elizabeth Degano on 12-17-12

H-6	This procedure was revised by Elizabeth Degano on 6-20-13

H-7	This procedure was revised by Elizabeth Degano on 10-24-14.