I.			PRINCIPLE:

A.		The Clinitek500 is a semi-automated, reflectance spectrophotometer designed to 			"read" Bayer Reagent Strips for Urinalysis. The Clinitek500 instrument analyzes 			the color and intensity of the light reflected from the reagent area and prints the 			results in clinically meaningful units.

II.		REQUIRED MATERIALS:

A.		Clinitek500 Siemens Healthcare Diagnostics Siemens     Diag.Multistix

B.      	Siemens Healthcare Diagnostics Multistix 8 SG

C.      	Distilled Water

D.      	Abnormal Urine Control, Count-l0-trol III, Myers-Stevens Group

E.      	Normal Urine Control, Count-l0-trol I, Myers-Stevens Group

III.		INSTRUMENT SET - UP:

A.	Daily cleaning: (Clinitek500 manual section 5.1-5.3)

Each morning the Clinitek500 must be cleaned.

1.	With the instrument at the Ready / Run screen (see manual p.4.1)

2.	Turn instrument OFF

Note: Must be in this screen when turning the instrument off for daily 			cleaning. Any other screen will inhibit the replacement of the platform.

3.	Remove the push bar by tilting it slightly upwards and pulling straight out.

4.	Remove the plastic waste liner and discard used reagent strips.

5.	Remove the fixed platform by pulling entire assembly forward.

6.	Remove the moving table (black piece) under platform in same manner.

7.	Remove the holdown plate from fixed platform by pressing up against the 			tab at the back of the plate. Then pull the other end from its retaining hole.

8.	Wash all removed parts with mild soap and warm water.

No bleach is used.

Note: when cleaning platform, avoid wiping across white calibration 	bars. Clean bars with cotton swabs and plain water. Allow to air dry.

9.	Reinstall moving table by aligning the grooves and pushing as far as it 			goes.

10.	Reinstall the holdown, arrow side facing up and arrow pointing in the 			back.  You should hear a snap when you press back in place.

11.	Reinstall fix platform, push all the way on.

If there is difficulty replacing platform after cleaning, turn instrument 	on and that will move the black ledge (the piece which moving table 	rests on) to lowest position. Then turn power off again and reinstall 	moving table then platform.

12.	Hold push bar by flattened end and slightly upward to insert in position. 			Extra care should be made to correctly seat push bar back into 			position; otherwise error message 24 or 25 will occur.

13.	Then turn the instrument back on.

IV.	QUALITY CONTROL:

A. 	Daily Instrument Calibration is not necessary on Clintek 500. Calibration is 	performed at each readhead immediately before reagent strip is read.

B.	Daily: directions for control preparation in macroscopic general procedure.

1.	Verify normal and abnormal controls are valid to run (5 day open vial 	stability)

2.	Warm normal and abnormal controls to room temperature.

3.  	Run controls on Clinitek 500 following “running specimens” procedure.

Follow Soft procedures for Urinalysis Q.C.

4.            Any unacceptable control runs will generate new orders. Controls should

be  repeated and noted in problem log. Instrument should not be run when

controls  are not acceptable and a supervisor should be alerted.

Note:   Controls should be run after opening new bottle of Multistix.

C.	Each shift of operation:

- Normal and Abnormal controls are ordered in Soft Total Q.C. and run on Clinitek 500 each shift.  Controls are then posted in “ Clinitek Instrument Menu.”

- Access Result Entry in Total Q.C. to add color and opacity.

V.	RUNNING SPECIMENS:

  - Touch display screen. This prompts the Ready / Run screen. Press ID key and with Barcode Reader scan barcode of urine sample.

Same process if adding a list of specimens.

B. 	After entering all accession numbers, then press             to return to Ready/Run

Screen.

C.	With the push bar on the left side, dip strip into specimen, tap off excess on

side of the tube and place strip with reagent pad up parallel to the push bar and against rear wall of the platform. The push bar will move to the right and then back to the left, ready for the next urine specimen.

Note:  “ Place strip” will be seen on screen when it is time to dip the next

specimen and  ID( accession #)  of the next specimen in the list is displayed

on the screen.

D.	To stop a run press                 .

E. 	The instrument will continue in “Run” mode until all reagent strips have moved to

waste container. On screen, “Completing Run   Please wait.” will be displayed

and results will be printed out.

F. 	Perform confirmatory test(s) on urines with positive findings as specified in

macroscopic general procedure.

VI.	RESULT REPORTING:

A.	The Clinitek 500 directly reports results in a slightly different format. Below are 			the results the instrument prints and the corresponding value we report. Normally 			this translation is down on the interface computer.

| TEST | Instrument Report | YNHH REPORT |
| --- | --- | --- |
| PH | 5.0 - 8.5 | 5.0 - 8.5 (same) |
|  | >=9.0 | >8.5 |
| Protein | Negative | Negative |
|  | Trace | Negative |
|  | 30 mg/dl | 1+ |
|  | 100mg/dl | 2+ |
|  | >=300 mg/dl | 3+ |
| Glucose | Negative. | Negative |
|  | 100 mg/dl | Trace |
|  | 250 mg/dl | 1+ |
|  | 500 mg/dl | 2+ |
|  | >=1000 mg/dl | 3+ |
| Ketones | Negative | Negative |
|  | # Trace | Negative |
|  | 15 mg/dl | Small |
|  | 40 mg/dl | Moderate |
|  | 80 mg/dl | Large |
| Blood | Negative | # Negative |
|  | Trace-Lysed | Trace |
|  | Trace-Intact | Trace |
| Blood | Small | Small |
|  | Moderate | Moderate |
|  | Large | Large |
| Leuk. Esterase | # Negative | Negative |
|  | Trace | Positive |
|  | Small | Positive |
|  | Moderate | Positive |
|  | Large | # Positive |
| Nitrite | # Negative | Negative |
|  | # Positive | Positive |

B.	Specific gravity is done on refractometer and those results reported. (there is a 			bias in the upward trend on Clinitek 500.)

VII.	TROUBLESHOOTING:

A.	Problems or error messages consult Section 8 “Troubleshooting” of Clinitek 500

manual.

B.	Error code 24 or 25 – can result from improper installment of platform or push 			bar.

1.	If push bar has erratic movement along metal rod- wipe rod with alcohol 				wipe and lightly grease with Cell Dyn tube grease.

VIII.	REFERENCES:

1.	Clinitek 500 Operating Manual, Siemens Healthcare Diagnostics

2.	Siemens Multistix8sg package insert, Siemens Healthcare Diagnostics

3.	Graff, Sister Laurine, A Handbook of Routine Urinalysis and Body Fluids,  Lippincott, Philadelphia, 2011.

IX.  	HISTORY:

H-1	This procedure was written by Donna Fico on 9-2006.

H-2      This procedure was revised by D.Fico on 12/2010.

H-3      This procedure was revised by D.Fico on 1/2012.

H-4      This procedure was revised by D.Fico on 11/2012.