I.	PRINCIPLE:

The BCSXP program uses the Control Journal to run daily controls. When a control value falls within the 	programmed acceptable range, a green checkmark appears next to it. Corresponding patient samples may 	then be run.  If the control value is out of acceptable range, a blue exclamation point appears next to the 	result; patient samples may not be run for that test until the control issue is addressed and repeat values are 	within acceptable range. (If patient samples are run when control values are out of range, all of those results 	will also be flagged with a blue exclamation point.)

All Control Journal results are stored in the QC Software via buffer file called ASCII file. This buffer file does not delete when Shift Change Delete is done, although the BCSXP Control Journal will be cleared of its results.

During start-up the QC Software all controls in the ASCII file are imported into their respective QC libraries. The means, CVs, and SDs are calculated and results plotted on Levy-Jennings graphs.

Note: The communication between the BCSXP and the QC Software program is one-way; updates on the QC Software will not update the BCSXP side. Therefore, any adjustments made to control ranges, etc, must be done on the BCSXP side.

In addition, the control for which values are being updated must not be on board the instrument when 	adjustment is attempted—the BCSXP will not allow an update to a control in use at the time. Be aware that 	changing a range will automatically close the library of that control, and open a new one, and there is no 	way to merge the data back into one library.

II.	PROCEDURE:

After controls have been run:

A.	If there is a green checkmark by the result, nothing further needs to be done.

B.	If a blue exclamation point appears next to the result:

1.	repeat control once; if result is still out, then

2	do an intensive wash of the pipettors (refer to maintenance procedures)

3.	repeat control once more; if still out, make new control, and repeat. If still out,

4.	change appropriate reagents/ diluents and repeat. If control still out, notify supervisor/ 				contact Dade Behring technical support.

C.	It is necessary to go into the QC Software and delete the out-of-range controls from the QC 			calculations. To do this:

- Click on QC button in the control Journal of BCSXP System. The QC software starts. During opening, the current data of the BCSXP system are automatically read in, analyzed and displayed in the QC journal dialog. Click on Close to start evaluation of the results. At the end of the QC results are transferred to the database and the periods are displayed in the main window of the QC Software.

- Select the desired period (s) in the period list.

- Click on the Result List button or Double click on the desired period to open the Period dialog.

- Click in the Period dialog on the Result List tab.

- Controls within the acceptable range will have a green checkmark; if a control was low  or high  the corresponding symbol will appear in red.

- Double-click on the unacceptable line. Check box  “Block this result. It will not be used for calculations and graphics” AND type in the reason for the deletion/ the action taken (eg. “Control rerun; result OK”; “Maintenance performed”, “New reagent added”, etc). Type your initials after the comment. Click OK.

*Note: A comment MUST be made if result is deleted from the QC calculations.*

7.	Close the dialog and check other control values in the same way.

D.	To exit QC Software

1.	Click on File in upper left corner

2.	Click Exit

III.  	HISTORY:

H-1 	This procedure was written by N. Kriz on 7-5-2006.

H-2	This procedure was reviewed by N Kriz on 10/12/2010

H-3	This procedure was reviewed by N Kriz on 4/1/2011

H-4	This procedure was reviewed by N Kriz on 2/13/2012

H-5	This procedure was reviewed by N Kriz on 3/12/2013

H-6	This procedure was revised for BCSXP by P Bahel on 11/20/14