- I.	PRINCIPLE:

The STA-Staclot Protein S kit, (Cat # 0746), is intended for the quantitative

measurement of the functional protein S level based on the principle of factor Va

inhibition.  The assay is based upon the cofactor activity of activated protein C.  This

enhancement is reflected by the prolongation of the clotting time of a system enriched

with factor Va which is  a physiological substrate for activated protein C.  Protein S is a

plasma vitamin K-dependent protein.  Protein S is synthesized in the liver as an

inactive precursor.  The active form is obtained after carboxylation of glutamic residues by a    vitamin K-dependent carboxylase.  Protein S has an essential anticoagulant

function.  It acts as the cofactor of activated protein C, with which it appears to form a

stoichiometric complex.  This complex binds strongly to phospholipid surfaces and thus

regulates the coagulation process, inhibiting thrombin-stimulated factors V and VIII.

Protein S greatly enhances the anticoagulant function of activated protein C, probably

by increasing protein C affinity for phospholipid membranes.

- II.	SPECIMEN:

A.  	Citrated blood 9:1(blood to anticoagulant) 3.2% sodium citrate.

B.   	Centrifugation:	20 minutes for 2,500 g

C.   	Plasma storage	6 months at -70 C.

Frozen plasma should be thawed  only once at 37 C for 5 			lminutes.

D.   	Unacceptable specimens: Samples that are short draws or clotted.

III.	REAGENTS AND EQUIPMENT

A.	Equipment:

- Centrifuge

Pipettes

Pipette tips

Plastic test tubes

STA/STA Compact

Cuvette roll (contains 1000 cuvettes)

- B.	Reagents:

1.  	STA Unicalibrator: Freeze-dried citrated human plasma.  This plasma

is intended for use as a calibration plasma for functional assays.  					Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes 				at room temperature. Swirl gently. Reconstituted stability on the 					STA/STA Compact is 4 hours.

2.  	STA System Control N + P: Control N is citrated normal human plasma 				and Control P is citrated abnormal human plasma, both freeze-dried.  				Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes 				at room temperature. Swirl gently.  Reconstituted stability on the 					STA/STA Compact is 8 hours.

3.	Owren-Koller Buffer: Ready to use.  Request the sample drawer to open 			through the MAIN MENU under LOADING and bar code the buffer.  			Place the buffer into the sample drawer.  Stability on the STA/STA 			Compact is 120 hours.

4.	Calcium Chloride 0.025 M solution: Ready to use (Use about 2.0 ml in a 			microcup).  Request   the reagent drawer to open through the MAIN 			MENU under LOADING. Barcode and load. Stability on STA compact is 			120 hours.

5.	Reagent 1: Protein S Deficient Plasma: Lyophilized human plasma, free of  		Protein S. Reconstitute each vial with 1.0 ml reagent grade water (make 2 			vials—put into a   microcup).

Let sit  60 minutes at room temperature.  Swirl gently. Request the reagent 		drawer to open through the MAIN MENU under LOADING and bar code 			the reagent. Place the reagent into the reagent drawer.  Reconstituted 			stability on the STA/STA Compact is 4 hours.

6.	Reagent 2: PCa: Human activated protein C, freeze-dried.  Reconstitute 			each vial with 1.0 ml reagent grade water ( make 2 vials—put in a 				microcup).  Let sit 60 minutes at room temperature.  Swirl gently.  				Request the reagent drawer to open through the MAIN MENU under 			LOADING and bar code the reagent.  Place the reagent into the reagent 			drawer.  Reconstituted stability on the STA/STA Compact is 4 hours.

7.	Reagent 3: F.Va: Preparation containing bovine Va, freeze-dried.  				Reconstitute each vial with 1.0 ml reagent grade water (make 2 vials—put 			into a microcup).  Let sit 60 minutes at room temperature.  Swirl gently.  			Request the reagent drawer to open through the MAIN MENU under       			LOADING and bar code the reagent.  Place the reagent into the reagent 			drawer.  Reconstituted stability on the STA/STA Compact is 4 hours.

- IV.	CALIBRATION (if needed):

A.	STA Unicalibrator: After the reconstitution period, request the sample drawer 	(STA) or the reagent drawer (STA Compact) to open through the MAIN MENU 	under LOADING and bar code the calibrator.  Place the calibrator into the 	appropriate drawer.  The assay value for Protein S is transferred to the STA/STA 	Compact when the calibrator is loaded and prompted for the bar code information.

B.	The standards are automatically prepared by the STA/STA Compact by diluting 	the STA-Unicalibrator with Owren-Koller Buffer according to the parameters 		entered in the Test

C.	STA-System Control N + P are used by the STA/STA Compact to validate the 	 	calibration (refer to the QUALITY CONTROL section).

D.	To order the calibration: Through the MAIN MENU select CALIB/CONTROL.

Select CALIBRATION and enter.  Select CLOT S by pressing F1 and

then F10 to run the calibration.

E.	To examine the calibration curve: Through the MAIN MENU under 	 	       	CAL/CONTROL select CALIBRATION and enter.  Place the cursor on CLOT S 	and enter.  The calibration curve will appear on the STA screen.

F.	To verify the calibration curve is valid: Check that the correlation coefficient 	      	(r value) of the curve is 0.90 or greater.  Compare the actual vs the interpolated 	values of the curve.  If a data point difference is greater than 10% for the actual 	value vs the interpolated value, the point can be deleted from the curve data.  	When this is done, the STA/STA Compact will automatically re-draw the curve, 	re-calculate the r value and re-run the controls.  If any controls are outside the 	expected range, the control must be re-run.  If the value is still outside the 	expected range, a new calibration curve must be run along with fresh controls.

G.	To print the calibration curve: While examining the calibration curve on the 	 	STA/STA Compact screen, press ESC for options.  Select PRINT.  The STA/STA 	Compact cannot print a calibration curve while the STA/STA Compact is running.

- V.	QUALITY CONTROL:

A.	STA System Control N + P: After the reconstitution period, request the reagent 	drawer to open  through the MAIN MENU under LOADING and bar code the 	controls.  Place the controls into the appropriate drawer.

B.   	QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC if the 	 	patients’ samples have been loaded.  The QC can also be ordered manually before 	loading patients’ samples.

C.  	All control ranges are monitored automatically by the STA/STA Compact.  If any 	 	      	controls are outside the ± 2 SD range, the STA/STA Compact will audibly and 			visually alarm the operator.  Otherwise, the results can be found in the AUTO    			CONTROL file (STA only) and the individual QC files. Control results are 			automatically filed in the STA/STA Compact QC file.  All results for a 24 hour  			period will be converted to a mean value at midnight.  This mean is set in the 			statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- VI.	PROCEDURE:

A. 	Refer to the START-UP procedure for the STA/STA Compact before running 		     	patient samples on the STA/STA Compact at the start of each shift.

B.  	Run the calibration if needed.

C. 	Request quality control if using a stored curve: Through the MAIN MENU under

CAL/CONTROL select QUALITY CONTROL and enter.  Select CLOT S

by pressing F1 and then F10 to run the QC.

D.	Load patients’ samples: Access the sample drawer(s) through the MAIN MENU 	 	     	under LOADING.  After the drawer opens, identify the type of specimen, such as 	  	     	micro-sample, with F8, or stat, with F12.  Identify the sample by bar coding or

manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.	In MANUAL MODE, the operator must order the test from the menu (order PS 	CLOT), and then press F10 to save.

F.	In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	         	selected in the AUTO MODE profile.

G.	In TELELOADING MODE, the STA/STA Compact will query the host computer 	and download the test as well as assign the status, i.e. stat.

H. 	As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

I.  	All dilutions of the calibrator, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the reportable assay range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

dilution provided that the option has been entered in the Test Set-up.

J.   	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

K.	For results in question that need operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left hand corner of the screen,

i.e. re-run test.

- VII.	CALCULATIONS:

A.   	The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to % activity.

B.   	The assay used a dilution of 1:10, sample to buffer.  The STA/STA Compact will 	 	      	automatically dilute the sample lower if the result falls below the lowest data 			point on the calibration curve.  The STA/STA Compact will also automatically 			dilute the sample at a higher dilution if the result falls above the highest data point 		on the calibration curve.  These parameters must be pre-set in the Test Set-up.

VIII.	REFERENCE  RANGE:

A.	50 --120 %  Each laboratory should establish its own reference range.

Assay is linear from 10-150% (results >150% will be run at 1:40 dilution;

results  < 150%   will be run at 1:5. These dilutions are performed automatically.)

IX.	REPORTING RESULTS:

A.    	The results for Protein S are reported out to the nearest whole number % and can be posted from Soft Instrument Menu (Stago 2) (Example:  130%).

For a result below the linearity threshold, manually enter “<10%”.

B.	The upper limit of reporting is 344%; for higher results, report “>344%”.

- X.	NOTES:

A.	When the STA/STA Compact re-tests a patient sample at a more appropriate 	dilution, (as pre-set in the Test Set-up) the results in the TEST PANEL screen will  	appear in blue and have already been corrected by the STA/STA Compact for the 	dilutional difference.

B.  	If a patient’s functional Protein S result is <50%, antigenic testing (Protein S 			Total and Free) is performed reflexively.

## XI.	LIMITATIONS OF THE PROCEDURE:

A.  		Always use the correct anticoagulant/blood volume ratio.

B. 	Heparin levels up to 1 IU/ml do not affect test results.

C.    	Factor VIII levels up to 250% do not affect test results.

D.   	In the described assay conditions, there is no in vitro dissociation of the

C4b-BP-protein S complex.

E.    	Lupus anticoagulant, if present, might affect results by elevation of the PTT.

F.	Hirudan (lepirudan) or other direct thrombin inhibitors, if present, will affect results by elevation of the PTT (makes Protein S level falsely elevated). If DTIs are present (call floor/ caregiver to determine this) Prot S functional assay should be canceled and Protein S total and free antigen testing should be done instead.

G.	Hemolyzed samples may yield incorrect results (comment on hemolysis in 	computer).

XII.	REFERENCE:

1.    STA-Staclot Protein S, Clotting Assay for Protein S by STA.

Package insert 24541-Revised February 1997.

2.    STA Unicalibrator, Calibration Plasma for Functional Assays of Coagulation

Parameters on the STA.  Package insert 26344 - Revised May 1996.

3.    STA System Control N + P, Control Plasmas for Assays of Coagulation

Parameters on STA.  Package insert 26384 - Revised May 1997.

- STA/STA Compact Operators Manual.

XIII.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-17-08.

H-2	This procedure was reviewed by N Kriz on 10/15/2010

H-3	This procedure was reviewed by N Kriz on 7/7/2011

H-4	This procedure was revised by N Kriz on 4/24/2012

H-5	This procedure was reviewed by N Kriz on 6/11/2013

H-6	This procedure was revised by N Kriz on 12/30/2013