I.	PRINCIPLE:

The STA-Stachrom Protein C kit is intended for use with the STA analyzers for the quantitative 	measurement of the functional protein C level by the synthetic chromogenic substrate method.  In this 	method, Protein C is activated by the specific activator (Reagent 1) derived from the venom of Agkistrodon 	c. contortrix. The quantity of enzyme thus formed is measured by its amidasic activity on the synthetic 	chromogenic substrate CBS 42.46 (Reagent 2), paranitroaniline release at 405 nm. The intensity of the 	color produced is directly proportional to the level of protein C initially present in the test plasma.

II.	SPECIMEN:

Citrated blood 9:1 (blood to anticoagulant) in 3.2% sodium citrate. Centrifuge at 2500 rpm 20 minutes. 	Freeze plasma at -70 C until tested. Hemolysis may affect results.

III.	REAGENTS:

A.	STA Unicalibrator: Freeze-dried citrated human plasma intended for use as a calibration plasma 			for functional assasys. Reconstitute eaaach vial with 1.0 ml reagent grade water.  Let set 30 			minutes at room temperature; swirl gently; allow to stabilize on instrument 10 minutes; stability 			on instrument 4 hours.

B.	STA System Control N & P: freeze-dried normal and abnormal human plasma, respectively. 			Reconstitute each vial with 1.0 ml reagent grade water; let set 30 minutes at room temperature; 			swirl gently; allow to stabilize 10 minutes on instrument; stability on instrument 4 hours.

C.	Owren-Koller buffer: ready to use. Stability on instrument is 120 hours.

D.	Reagent 1: freeze-dried, purified extract of Agkistrodon c. contortrix venom. Reconstitute each 			vial with 3.0 ml d.w.; let set at room temperature 60 minutes; swirl gnetly; add reducer to vial and 			allow to stabilize on instrument 10 minutes. Stability on instrument is 21 days.

E. 	Reagent 2: freeze-dried chromogenic substrate CBS 42.46. Reconstitute each vial with 6.0 ml 			d.w.; let set at room temperature 60 minutes; swirl gently; add reducer to vial; allow to stabilize 10 		minutes on instrument; stability on instrument 21 days.

IV.	PROCEDURE:

A.	If needed, request standard curve (calibration), being sure that appropriate calibrator, controls and 			reagents are in place.

B	If calibration is already in place, request only quality control to be run.

C. 	Load patient samples, requesting protein C assay (or dependent PC 6 test) as needed; close sample 		drawer; when the drawer closes, the TEST STATUS screen will appear. If there is not enough 			reagents to run the test, the suspect reagent(s) will appear in red with the amount of deficiency. 			This deficiency will BLOCK the sample pipettor. When this occurs, add the deficient reagent(s) to 		clear the block message.

D.	All dilutions of calibrator, controls, and patient samples are automatically prepared by the STA 			according to parameters in the test protocol. If so designated in the protocol, all patient results 			falling outside of assay range will be automatically retested at an appropriate dilution. All results 			are displayed on the TEST PANEL screen and automatically print out.

E. 	For results in question that need operator intervention, cusor to the ID number in the TEST 			PANEL screen and enter; this will bring up the FILE PROCESSING screen, where operator can 			follow the options in the left hand cornor of the screen (eg: rerun test).

V.	NOTES:

A.	Results are reported out to the nearest whole number %.

Normal range= 70-140%.

B. 	Assay is linear 0-150% (results >150% are automatically rediluted). The upper limit of reporting is 300%; if result is higher, report “>300%”.

C.	Heparin levels up to 1.0 u/ml should not affect test results. Higher levels (ie: non-therapeutic 			levels, as may be found in samples drawn from a central line) may lead to an over-estimation of 			the protein C level.

D. 	It has been reported that highly elevated levels of factor VIII (> 250%) may lead to and under-			estimation of the protein C level.

VI.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-22-2004.

H-2	This procedure was reviewed by N Kriz on 10/15/2010

H-3	 This procedure was reviewed by N Kriz on 5/2/2011

H-4	This procedure was reviewed  by N Kriz on 4/19/2012

H-5	This procedure was reviewed by N Kriz on 6/11/2013

H-6	This procedure was revised by N Kriz on 12/30/2013