I.	PRINCIPLE:

Patient’s plasma is diluted 1:5 with Factor V-deficient plasma  (V-DEF) which also 	contains polybrene, a heparin antagonist. An APTT is performed on the mixture, using 	the silica-based APTT reagent and CaCl2 that has been colyophilyzed with activated 	protein C (APC-CaCl2). Normal subjects will exhibit a marked prolongation of the APTT 	in the presence of APC, and their post-APC/ pre-APC ratio will be >2.20.

II.	REAGENTS:

A.	From kit:

1.	1 vial (8 ml) CaCl2 0.025 mol/L (Stable 1 month 2-8C)

2.	1 vial (16.0 ml) silica-based APTT reagent. (Stable 1 month 2-8C)

3.	4 vials APC colyophilized with CaCl2 (reconstitute each vial needed with 1.5  mL NERL water. Leftover product may be frozen in original vial for one reuse. (Stable 3 months at -70C.)

4.	4 vials V-DEF plasma with polybrene (reconstitute each vial needed with 				4 ml distilled water. Leftover product may be frozen in original vial for 				one reuse. (Stable 3 months at -70C.)

B.	Also included in kit (to be run as “samples”):

1.	1 vial normal control (reconstitute with 1 ml distilled water. Aliquot 				leftover control into 0.2 ml amounts in microtainers and freeze for later 				use. (Stable 3 months at -70C.)

2.	1 vial abnormal control (reconstitute, aliquot and freeze as for normal 				control. (Stable 3 months at -70C)

C.	To be run as “control”:

1.	George King pooled normal plasma, as QC for assay protocol.

III.	PROCEDURE:

A.	Reconstitute all reagents as described above and allow tosit for 20 minutes before 			using.

B.	Vortex the silica-based APTT reagent before using.

C.	Reagents should be placed in reagent drawer on the STA-Compact using the 				following designations:

1.	In original vial:

- 9000	A..PTT		10.0(ml)	24(hrs)		1111(“lot”#)

2.	In microvolume:

- 9001	CACL2 0.025M	2.0(ml)		24(hrs)		1111

3.	In microvolume:

- 9002	APC/CACL2		2.0(ml)		8(hrs)		1111

4.	G.K. Normal “control”, in microvolume:

NORMAL	NORMAL	0.4(ml)		8(hrs)		2222

D. 	 In Sample drawer, as diluent, place:

1.	 F5DP	  F5 DEF PLASMA	          2.0(ml)		8(hrs)		1111

2.	 In Sample drawer, as “patient”, place:

a.	0.3 ml Normal control from kit

b.	0.3 ml Abnormal control from kit

E.	Request “control”  (G.K. pooled normal plasma) for tests APC-1 (“first” PTT, 				with CaCl2 only) and APC-2 (“second” PTT, with APC/CaCl2).

F.	Thaw patient samples and place on the STA-Compact in the usual manner.

G.	For normal and abnormal control (placed on as patient samples), request the tests 				as follows:

- APC-1

- APC-2

- APC1APC2  (this calculates the ratio). Ensure that this ratio falls within the 			acceptable range indicated on the kit assay sheet.

H.	When the controls have finished running, request the tests as above for each 				patient specimen.

I.	After the run has completed, check patient ratios for abnormals. The cutoff 				value is 2.20; patients with an APC resistance ratio of 2.20 or  greater should 				be considered normal and those with a ratio of <2.20 should be considered 				abnormal.

IV.	NOTE:

A.		The APC ratio of the normal control should be about 2.20-3.10 (see product 			insert, as this range varies with lot #).

B.		The APC ratio of the abnormal control should be about 1.40-1.90 (see product 			insert, as this range varies with lot #).

C.		Patients with an APC ratio of <2.20 are classified as abnormal (i.e. consistent 			with presence of Factor V Leiden mutation).

V. 	 HISTORY:

H-1 	 	This procedure was written by N. Kriz on 12-23-04.

H-2		This procedure was reviewed by N Kriz on 7/26/2010

H-3		This procedure was reviewed by N Kriz on 2/25/2011

H-4		This procedure was reviewed by N Kriz on 1/31/2012

H-5		This procedure was revised by N Kriz on 3/11/2013