- PRINCIPLE

STA- Stachrom Antiplasmin, (Cat # 0659), is used for the quantitative determination of

the antiplasmin level by the synthetic chromogenic substrate method.  In the presence of

added plasmin, the test plasma exhibits a rapid and powerful potential for inhibiting

plasmin, and this inhibition is exerted by the antiplasmin present in the test plasma.  The

assay is carried out in two steps:

-  incubation of the test plasma with a known excess of plasmin

- Antiplasmin + plasmin > Plasmin-antiplasmin + plasmin

(excess)                                                 (residual)

-determination of the residual plasmin by its amidolytic activity on the synthetic

chromogenic substrate CBS 10.85 (pNA release at 405 nm)

- Plasmin (residual) + CBS 10.85 >  pNA

- The residual quantity of plasmin (as measured by the pNA release) is inversely

- proportional to the antiplasmin level of the test sample.

- SPECIMEN

1.   Citrated blood 9:1 (blood to anticoagulant) 3.2% or 3.8% sodium citrate.

Follow NCCLS guidelines H3-A3 and H21-A2.  No other anticoagulant

is acceptable.

2.   Centrifugation:	20 minutes at 2,500 g

3.   Plasma storage:

3 month at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

4.  Unacceptable specimens: Samples that are QNS or clotted ; hemolyzed may

yield incorrect results and that should be commented.

- REAGENTS AND EQUIPMENT

- Centrifuge

Pipettes

Pipette tips

STA/STA Compact

Cuvette roll (contains 1000 cuvettes)

- REAGENTS

1.   STA Unicalibrator: Freeze-dried citrated human plasma.  This plasma is

intended for use as calibration plasma for functional assays.  Reconstitute

each vial with 1.0 ml reagent grade water.  Let set 30 minutes at room temperature.

Swirl gently.  Reconstituted stability on the STA/STA Compact is 4 hours.

2.   STA System Control N and P: Control N is citrated normal human plasma and

Control P is citrated abnormal human plasma, both freeze-dried.  Reconstitute

each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at room

temperature.  Swirl gently.  Reconstituted stability on the STA/STA Compact is 8 hours.

3.   Owren-Koller Buffer: Ready to use.  Request the sample drawer to open through

the MAIN MENU under LOADING and bar code the buffer.  Place the buffer

into the sample drawer.  Stability on the STA/STA Compact is 120 hours.

4.   Reagent 1: Plasmin: Human plasmin, freeze-dried.  Plasmin activity per vial

is approximately 26 nKat.  Take one vial of Reagent 3 and shake it well.  Then,

pour its entire contents into a vial of Reagent 1 from the same kit.  Replace the

perforated plastic cap on the vial.  Let set 60 minutes at room temperature.  Swirl 	 	      gently.  Request the reagent drawer to open through the MAIN MENU under  LOADING and bar code the reagent.  Place the reagent into the reagent drawer,

RD1.  Reconstituted stability on the STA/STA Compact is 8 days.

5.   Reagent 2: Chromogenic substrate CBS 10.85, approximately 12 moles per vial.

Reconstitute each vial with 6.0 ml reagent grade water.  Replace the perforated

plastic cap on the vial.  Let set 60 minutes at room temperature.  Swirl gently.  		      Request the reagent drawer to open through the MAIN MENU under LOADING

and bar code the reagent.  Place the reagent into the reagent drawer, RD2.

Reconstituted stability on the STA/STA Compact is 8 days.

6.   Reagent 3: Plasmin Solvent: Ready to use.  Shake vial well before use to

reconstitute Reagent 1.

- CALIBRATION (if needed—minimum of every six months)

1.   STA Unicalibrator: After the 30 minute reconstitution period, request the sample

drawer to open (STA) or the reagent drawer to open (STA Compact) through the 		      MAIN MENU under LOADING and bar code the calibrator.  Place the calibrator

into the appropriate drawer.  The assay value for Antiplasmin is transferred to the

STA/STA Compact when the calibrator is loaded and prompted for the bar code

information.

2.   All standard dilutions are automatically prepared by the STA/STA Compact by 	 	      diluting the STA Unicalibrator with Owren-Koller Buffer according to the     parameters entered in the Test Set-up.

3.   To order the calibration: Through the MAIN MENU select CALIB/CONTROL.

Select CALIBRATION and enter.  Place the cursor on AP and enter.  Select AP

by pressing F1 and then F10 to run the calibration.

4.   STA System Control N + P are used by the STA/STA Compact to automatically 	 	      validate the calibration.  (Refer to the QUALITY CONTROL section).

5.   To examine the calibration curve: Through the MAIN MENU under

CALIB/CONTROL select CALIBRATION and enter.  Place the cursor on AP

and enter.  The calibration curve will appear on the STA/STA Compact screen.

6.   To verify the calibration curve is valid: Check that the correlation coefficient

(r value) of the curve is 0.90 or greater.  Compare the actual vs the interpolated

values of the curve.  If a data point difference is greater that 10% for the actual

vs the interpolated value, the point can be deleted from the curve data.  When this

is done, the STA/STA Compact will automatically re-draw the curve, re-calculate

the r value and re-run the controls.  If any controls are outside the expected range,

the control must be re-run.  If the value is still outside the expected range, a new 	 	      calibration must be run along with fresh controls.

- To print the calibration curve: While examining the calibration curve on the 	 	       STA/STA Compact screen, press ESC for options.  Select PRINT.

- The STA/STA Compact cannot print a calibration curve while the STA/STA Compact is running.

- QUALITY CONTROL

1.   STA System Control N + P: After the reconstitution period, request the sample

drawer to open (STA) or the reagent drawer to open (STA Compact) through the 	 	      MAIN MENU under LOADING and bar code the controls.  Place the controls

into the appropriate drawer.

2.   QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC when the

patients’ samples have been loaded.  The QC can also be ordered manually before 	 	      loading patients’ samples.

- All control ranges are monitored automatically by the STA/STA Compact.  If any 	   controls are outside the  2 SD range, the STA/STA Compact will audibly and 	       visually alarm the operator.  Otherwise, the results can be found in the AUTO 	       CONTROL file (STA only) and the individual QC files.

- Control results are automatically filed in the STA/STA Compact QC file.  All results for a 24 hour period are converted to a “mean” value at midnight.  This mean is used in the statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- PROCEDURE

1.   Refer to the START-UP procedure for the STA/STA COMPACT before running 	 	      patient samples on the STA/STA Compact at the start of each shift.

2.   Run the calibration if needed.

3.   Request quality control if using a stored curve: Through the MAIN MENU

under CAL/CONTROL select QUALITY CONTROL and enter.  Place the

cursor on AP.  Select AP by pressing F1 and then F10 to run the QC.

4.   Load patients’ samples: Access the sample drawer(s) through the MAIN

MENU under LOADING.  After the drawer opens, identify the type of specimen,

such as microsample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

5.   In MANUAL MODE, the operator must order the test from the menu and then

press F10 to save.

6.   In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	 	      selected in the AUTO MODE profile.

- In TELELOADING MODE, the STA/STA Compact will query the host computer and download the test as well as assign the status, i.e. stat.

8.   As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

9.   All dilutions of the calibrator, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

Dilution provided that the option has been entered in the Test Set-up.

10.  All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

11.  For results in question that need operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options the left hand corner of the screen,

i.e. re-run test.

- CALCULATIONS

- The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to % activity.

- The STA/STA Compact uses a dilution of 1:10, sample to buffer.  Following pre-

set parameters, the STA/STA Compact will automatically dilute the sample lower

if the result falls below the lowest data point on the calibration curve.  The

STA/STA Compact will also automatically dilute the sample at a higher dilution if

the result falls above the highest data point on the calibration curve.

3. Linearity of the assay is 0-130%.

REFERENCE RANGE

__70___ - __140___ % Each laboratory should determine its own reference range.

- REPORTING RESULTS

- The results for Antiplasmin are reported out to the nearest whole number %.  		      (Example:  103%). Results are linear to 130%; automatic dilutions allow for reporting of higher values. If result is higher than 130%, report “>130%”.

- Detection limit of assay is 7% ; if lower, report “<7%”.

- NOTES

1.   When the STA/STA Compact re-tests a patient sample at a more appropriate

dilution, (as pre-set in the Test Set-up) the results in the TEST PANEL screen will 		      appear in blue and have already been corrected by the STA/STA Compact

for the dilutional difference.

2.   Since most patients’ results are higher than 100% and will probably be

re-tested at a more appropriate dilution, it is more efficient to run the

patient samples initially at a higher dilution, i.e. 1:20.  In order to do this,

a Dependent Test with the selected dilution must be created.  (See the

Operators Manual - Dependent Test).

- LIMITATIONS OF THE PROCEDURE

1.   Do not use plasmas collected with sample collection tubes containing either

plasmin inhibitors or heparin, since the concentration of these substances

can interfere strongly in the test.  However, the same substances at

therapeutic levels have little or no effect on the test.

2.   In certain therapeutic cases a slight interference by alpha2-macroglobulin

cannot be avoided, for instance when the alpha2-antiplasmin level is greatly

depressed in the presence of a high level of alpha2-macroglobulin.

- REFERENCES

- STA-Stachrom Antiplasmin, Cat # 0659, Colorimetric Assay of Antiplasmin by

STA Analyzer.  Package insert 23460-Revised May 1996.

- STA Unicalibrator, Cat # 0675, Calibration Plasma for Functional Assays of Coagulation Parameters on the STA.  Package insert 26344 - Revised May 1996.

- STA System Control N + P, Cat # 0678, Control Plasmas for Assays of Coagulation Parameters on STA.  Package insert 26384 - Revised May 1997.

- STA/STA Compact Operators Manual.

5.  HISTORY:

H-1	This procedure was written by N. Kriz on 01-06-2005.

H-2	This procedure was revised by N Kriz on 2/1/2010

H-3	This procedure was reviewed by N Kriz on 2/3/2011

H-4	This procedure was reviewed by N Kriz on 1/2/2012

H-5	This procedure was reviewed by N Kriz on 1/28/2013

H-6	This procedure was revised by N Kriz on 12/27/2013