# Hematology Lab Weekly Meeting

# Meeting Minutes

May 25, 2017

- The Arkray seems to be working well. A few reminders:

- Only add one bottle of sticks at a time on the analyzer. ( Click Consumables > Test strip> wand bar code> Execute)

- CHM QC is always run  after a new bottle is put on analyzer.

- After running QC, need to document. ( Click Controller > Consumable Log> QC ok and your initials)

- Do not take strips from one analyzer to put in another. Let each bottle run out. No topping off.

- Every Monday – both Arkrays need S.G. well cleaned with bleach- See procedure.

- Also we are still checking every Monday, 5% Nacl on digital refractometer

- Another reminder: The blood testing for Know Your Numbers began on May 15. Even if this is completed, there is a $350.00 credit to the cost of health insurance.

- Details to follow, but a more formal process for documenting results across to the EMR after computer downtime or maintenance will b put in place.

- Any ideas for Lab Week 2018 please let Natalie know, same for 2017 Christmas Party. If anyone is interested in joining committee also let Natalie know.

- PFA weekly normal control day will be Tuesday. If there is an aggregation (or a normal control for aggregation) that should be analyzed as the normal control. If there is no aggregation, a normal patient came be used as the normal but needs to be run on both instruments.

- Good job entering instrument problems onto problem log. If you see problem recently entered in log let a supervisor know.

- Starting, Thursday, June 1- we will be starting to use the new logs for Urines and Fluids.