|  | TITLE: ACL TOP 750 Operational Procedure |  | DEPT OF LAB MEDICINE CLINICAL HEMATOLOGY Policy and Procedure Manual |
| --- | --- | --- | --- |
|  |  |  | DOCUMENT # HEM215 |
|  |  |  | PAGES 1-42 |
| WRITTEN BY: Parveen Bahel H (ASCP) | EFFECTIVE DATE: | REVISION/REVIEW: H-1 (New); 09/2019 |  |

- INTENDED USE

The ACL TOP Family 50 Series (ACL TOP 750; ACL TOP 750 CTS; ACL TOP 750 LAS; ACL TOP 550 CTS; ACL TOP 350 CTS) are benchtop, fully automated, random access analyzers designed specifically for in vitro diagnostic clinical use in the hemostasis laboratory for coagulation and/or fibrinolysis testing in the assessment of thrombosis and/or hemostasis. The systems provide results for both direct hemostasis measurements and calculated parameters.

- PRINCIPLE

The ACL TOP instrument performs the following types of tests:

Coagulometric (Turbidimetric) Measurements

The principle of coagulometric (turbidimetric) clot detection is used in the system to measure and record the amount of time required for a plasma specimen to clot. This technique assesses coagulation endpoint by measuring change in optical density.

Clot detection is based on the principle that light passing through a medium in which fibrinogen is converted to fibrin is absorbed by the fibrin strands. Light (405 or 671 nm) is transmitted through a sample onto a photodetector, which is positioned 180° to the source.

Light absorption increases as fibrin clot formation progresses. Consequently, light transmittance through the sample continuously decreases and is measured by the photodetector.

The corresponding electrical signal output from the photodetector changes according to the detected light. The signal output is processed via software through a series of algorithms to determine the clot point.

Chromogenic (Absorbance) Measurements

Chromogenic tests can be either direct or indirect.

Direct test– Test where the analyte of interest (e.g.proteinC, plasminogen) acts directly on a specified synthetic substrate.

Indirect test–Test where the analyte of interest (antithrombin, plasmin inhibitor) reacts with a fixed quantity of enzyme to form inactive complexes. Under optimized test conditions, residual enzyme activity is then measured using a specific synthetic substrate.

In most cases, the reaction is monitored at 405 nm by the continuous release of paranitroaniline (pNA) from the synthetic substrate.

The chromogenic channels use the colorimetric principle of measuring absorbance in the cuvette. An optical sensor reads light (405 nm) that passes through the cuvette. The light is absorbed by the fluid in the cuvette in direct proportion to the concentration of pNA. The amount of light reaching the photodetector is converted into an electrical signal that is proportional to the enzyme activity.

Immunological Measurements

The principle of immunological measurement is used on the system to directly measure and record the amount of an analyte. This technique assesses the physical concentration of the analyte (and not its activity) by measuring change in optical density.

Although similar to the turbidimetric method, the immunological method relies on the formation of antigen- antibody complexes to affect light transmission.

Immunological testing of the ACL TOP uses the 405 nm or the 671 nm channels depending on the test and the reagent formulation.

Both the 405 nm and the 671 nm channels use the principle of measuring absorbance in the cuvette. An optical sensor reads the light (405 nm or 671 nm) that passes through the cuvette. The light is absorbed by the fluid in the cuvette in direct proportion to the concentration of antigen-antibody complexes. The amount of light reaching the photodetector is converted into an electrical signal that is proportional or inversely proportional to the analyte concentration.

- SAFETY  PRECAUTION

All patient specimens should be considered potentially infectious and must be handled with precautions used for human blood, as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29CFR part 1910.1030.  Follow specimen handling, use of standard precautions; proper PPE wears gloves and a lab coat. Wear safety glasses if there is a risk of splashing.

- SPECIMEN REQUIREMENTS

Mix nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant.

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing of blood and anticoagulant.

A syringe or evacuated tubes (blue top) may be used for collection. If multiple specimens are collected; the coagulation sample should be the second or third tube collected. If only coagulation testing is to be performed, a red-top tube, which has no additives, should be drawn first and discarded prior to drawing the blue-top coagulation tube.

The patient cannot be on anti-coagulants when the test specimen is collected. Sufficient time after discontinuance of heparin should be allowed for heparin to be cleared from the patient’s blood, usually 6 hours.

If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values above 55%.

Specimens that are clotted, collected in the wrong tube or serum, overfilled or have less than the 90% expected fill should be rejected.

Handling Condition and Stability

The whole blood specimen is checked for clot formation by gentle inversion and observation. Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x109/L for 10 minutes at 4000 g.  Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells. For special coagulation testing, spin samples 20 minutes at 4000 g, separate plasma into plastic tubes, label and freeze all aliquots at –70C located in the Special coag area until ready to use. Always check for clots after aliquoting with applicator sticks. Always tracks aliquots in BEAKER under YH Coag Hold before freezing them. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested.

Specimen stability at ambient temperature: 4 hours; Frozen at -70° C 6 months.

Specimen Labelling Specimen should be properly labeled with at least 2 unique patient identifiers. Patient's full name and medical record number (MRN) and should have the date and time of collection. The patient's birth date may substitute for the MRN if the MRN is not available.

- SUPPLIES AND REAGENTS

  - Supplies

    - Nerl Water: pH 7.0

    - Gauze

    - Citrated blue top tubes

    - Frosted tubes for aliquots

    - Cuvettes

    - ACL Top sample cups

    - HemosIL Cleaning solution Clean A and Clean B

    - HemosIL Rinse and waste

    - HemosIL factor Diluent

  - Reagents

NOTE: After opening any vial to place onto the instrument, label that vial with the open and expiration date, referring to the stability information provided here.  Discard reagent when expired.

| Reagents Stability after reconstitution |  |  |  |
| --- | --- | --- | --- |
| Description | On Board | 2-8 oC | (- 20 o C) |
| HemosIL dRVVT Screen | 3 Days | 15 Days | x |
| HemosIL dRVVT Confirm | 3 Days | 15 Days | x |
| HemosIL Silica Clotting Time | 5 Days | 20 Days | x |
| HemosIL Factor V Leiden (APCR V) | 3 Days | 30 days for CaCl2 and aPTT reagent | 3 Months for Factor V reagent and APC/CaCl2 |
| HemosIL Liquid Antithrombin (2mL) | 48 Hrs | 5 weeks | x |
| HemosIL Protein C | 5 Days | 3 Months | x |
| HemosIL Protein S Activity | 6 Hrs | 24 Hrs | x |
| HemosIL Free Protein S | 7 Days | 1 Months | x |
| HemosIL Plasminogen | 5 Days | 3 months | 6 Months |
| HemosIL Plasmin Inhibitor | 24 hrs | 5 Days | 3 Months |
| HemosIL von Willebrand Antigen | 7 Days | 3 Months | x |
| HemosIl von Willebrand Activity | 5 Days | 1 Months | x |
| HemosIL Factor XIII Antigen | 4 Days | 1 Months | x |
| HemosIL RecombiPlasTin 2G (8mL) | 10 Days | 10 Days | x |
| HemosIL SynthASil | 10 Days | 30 Days | x |
| Calcium Chloride | 30 Days | 30 Days | x |
| Description | On Board | 2-8 oC | (- 20 o C) |
| HemosIL Factor II deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL Factor V deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL Factor VII deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL FVIII def. plasma (normal VWF) | 24 Hrs | 24 Hrs | 3 Weeks |
| HemosIL Factor IX deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL Factor X deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL Factor XI deficient plasma | 24 Hrs | 24 Hrs | x |
| HemosIL Factor XII deficient plasma | 24 Hrs | 24 Hrs | x |

| IL Special Coag Test QC |  |  |
| --- | --- | --- |
| Name | Test QC | Onboard Stability After reconstitution |
| HemosIL Normal Control assay | All Tests set up on both Instruments | 24 hours |
| HemosIL Special Control Level 2 | All tests except vWD Activity and protein S Activity | 24 hours |
| HemosIl Special Control Level 1 | vWD Activity | 24 hours |
| HemosIL Low Abnormal Control Assayed | Protein S Activity | 24 hours |
| HemosIL LA Positive Control | Silica Clotting and DRVVT tests | 24 hours |
| HemosIL LA Negative Control | Silica Clotting and DRVVT tests | 24 hours |
| APC Normal and Abnormal Control | Activated Protein C resistance | 6 hours |

Rinse and Cleaning Solutions: Expiration as indicated on the product

| Rinse and Cleaning Solutions | On-board stability after opening |
| --- | --- |
| HemosIL Rinse solution | 1 month |
| HemosIL Cleaning Agent | 1 Month |
| HemosIL Factor diluent | 1 month |
| HemosIL Cleaning solution (Clean A) | 1 month |

- ENVIRONMENTAL CONDITIONS FOR RELIABLE TEST PERFORMANCE:

The instrument functions correctly in an ambient temperature of 15° C to 32° C (59° F to 89° F) with a relative humidity of 15% to 85% (non-condensing).

In accordance with the IEC regulations, no instrument failures occur in the presence of short-term ambient temperatures as low as 5° C or as high as 40° C.

The ACL TOP Family 50 Series is compliant with IEC 60068-2-40 to 2000 meters. The instrument should not be used at an altitude greater than 2000 meters.

The instrument should be placed in an area free from dust, fumes, vibrations and excessive variations of temperature.

The heat generated by the instrument during normal operation is exhausted from the bottom, the front-right and the left side of the unit.

According to IEC 61010-1, the maximum audible noise emission should be 80 dBA. The ACL TOP Family 50 Series is compliant with IEC 61010-1 Third Edition.

The room temperature and humidity percent are monitored and documented on the Routine Coagulation checklist.

- ACL TOP 750 OVERVIEW

Modules:

Control Module (CM): The CM provides a user interface and operation control.

Analytical Module (AM) The AM consists primarily of sample and reagent handling hardware including

AM Computer
o Cuvette Handling
o Sample Area
o Diluent Area
o Reagent Area
o Bulk Fluids
o Waste Handling
o Sample Handling
o Reagent Handling
o Reaction and Detection
o Interconnect and Power Supply
o Access-restricting Cover with Safety Interlocks o Instrument-supporting Structure-chassis
o Safety Interlock for Cuvette Waste Drawer

Cuvette Loading Area The cuvette loading area is on the left-most side of the AM. The cuvette loader can be filled with up to 20 clips of 10 cuvette strips each, for a maximum of 200 cuvette strips. cuvettes can be loaded while the analyzer is running.

Sample Area The Sample loading area is on the left side of the track control panel. The sample material is placed in racks that while being inserted, are identified by the bar code reader. The sample rack holds both sample tubes and sample cups. The sample area is at ambient temperature.

Reagent and Sample Racks, Adapters

Reagent Rack(R)		Diluent Rack(D)	Sample Rack(S)

Reagent Rack Adapters 		Diluent Rack Adapters			Sample rack 
10 mL & 4 mL vials 			4 mL, 10 mL & 16 mL vials			tube adaptor

Software Overview

The Menu Bar contains the following options:

Instrument: Actions: Analysis: QC: Calibration: NPP: Setup: System: Help:

Instrument: General instrument functions
Actions: Functions applicable to the current screen/data set displayed
Analysis: Access to analytical area maps (i.e. Samples, Diluents, and Reagents)

QC: Review QC results and QC status for tests
Calibrations: Review status, stored calibrations or perform calibration
NPP: Review Normal Pooled Plasma results
Setup: Access area for software configuration
System: Perform/review maintenance and diagnostic functions
Help: Access on-line help (Operator’s Manual)

The Tool Bar main functional icons:

Software Screen Examples

Reaction Curve Analysis

A reaction curve is a visual representation of the coagulometric or absorbance reaction taking place in a cuvette on the analyzer. The curve can easily be displayed when needed for all results on the ACL TOP Family of analyzers.

A curve is a valuable tool for investigating abnormal or questionable results.

Parts of a Clotting Reaction Curve

Valid clot reaction curves have the 3 segments shown on the curve above. Normal clot reaction curves have what is referred to as an “S” shape categorized into 3 segments:

1. Baseline (sample and reagents are mixed and reaction begins)
2. Acceleration segment (for clotting assays Fibrin is now forming in this segment which changes the mAbs readings)
3. Plateau segment (for clotting assays, all of the fibrinogen has converted to fibrin and Therefore optical change in reaction levels off)

Reagent Area

Double Click on a placed rack to view rack details shown below. Rack details show materials in the rack, lot number, onboard stability, and remaining measured volume. Colored circles indicate status: Examples – Green = Acceptable for analysis, Red = Not Acceptable for analysis.

Sample Rack Screen

Double click on a sample rack to view the details (tests programmed) on the samples within the rack (shown on next page). The “ambulance” symbol indicates the sample is a STAT. Samples are color-coded for status. Common colors include Blue (Programmed), Purple (Analyzing), Green (Analysis Complete). A red exclamation mark in the circle indicates a warning on the results.

- PERFORM SHUTDOWN, START-UP, AND BACK-UP

  - Shutdown: Software –►Analyzer Module (AM) -►  module (AM)

To shut down the instrument:

- Make sure the Analytical Module in the READY state.

- Select Instrument > Exit in the menu bar or single click on the X on the upper right corner of the screen.

- Select OK at the Are you sure you want to exit? Prompt.

- Once the software has closed, turned off the Analytical Module (AM). The switch is located on the right side of the analyzer.

- Move the cursor to the left corner of the Control module (CM) to display the Windows Task Bar.

- Choose Start.

- Click Shutdown.

- Choose Shutdown from the dropdown list then OK.

- Verify sample and Reagent access covers are closed.

  - Starting the Instrument up

Startup: Control module (AM) - Analyzer Module (AM)- Software

- Make sure the access cover is closed.

- Turn on the Computer (Control Module).

- Immediately turn on the Analytical Module using the On/Off switch on the right side of the analyzer.

- Double-click the ACL TOP  icon on the desktop.

- The analyzer will turn on and begin a start-up.

- The ACL TOP software will display the login screen

    - Type “topcoag” for the Username (Case sensitive)

    - Type “topcoag” for the Password

    - Click [OK]

- The analyzer will initialize, which takes about 3 minutes.

- Reload all racks.

  - Backing up the Database

To back up the database:

- Exit the ACL TOP® Family software, by choosing Instrument ►Exit from the Menu Bar or by choosing the X in the upper right-hand corner.

- On the desktop, click ACL Top Database back-up and restore icon.

- Do Not select Enable multi-volume backup

- Choose Delete patient information from the backup. (Patient information - demographic information- is not included in backup).

- Select Backup from the Database Backup and Restore window. (If you are prompted to insert a disk into a drive, select Cancel to display the Save As dialog box). Use Backup folder on the desktop in the Save location field.

- Verify the serial number and date is in the File name: field.

- Choose Save.

- Choose OK at the “The database backup has been executed successfully” prompt.

- Choose Close.

- Open the ACL TOP by clicking ACL TOP Version icon on the desktop.

- Copy the backup file to another location, such as a CD or external drive, to prevent accidental

deletion. See Copying the Database Backup file.

- Start the ACL TOP application and and follow the steps 5 to 8, section B (Starting the Instrument up).

- INSTRUMENT MAINTENANCE

Daily Maintenance

A.	To Perform Enhanced Clean for All Probes

- Place a 10mL vial of Clean B (Cleaning Agent) in an empty Reagent Rack and load in Reagent track position R5.

- Place a second 10mL vial of Clean B in an empty Diluent Rack and load it in the D3 position.

- Place a third 10mL vial of Clean B in an empty Diluent Rack and load it in either the D1 or D2 position.

- Choose Menu Bar ► System ► Maintenance.

- Select Enhanced Clean for all Probes.

- Single-click or touch Perform Activity .

- A Comment window will display when the maintenance activity is finished. Enter your initials (optional) then choose OK.

B.	Remove the (3) Clean B vials from the analyzer after the cleaning is complete. Once removed, enter your initial. (Clean B is a corrosive material. Remove it from the instrument immediately after performing any Enhanced Clean procedure).

C. Preparation of Fresh Clean B Diluted

- Remove the Reagent Rack with the Clean B Diluted. Discard the remaining liquid in the vial. Rinse vial, if needed, with DI water.

- Make fresh Clean B Diluted by measuring 14 mL DI water then adding 2mL of Clean B. Mix and transfer into the 20mL bottle.

- Load the fresh Clean B Diluted and enter initial.

D. Temperature Check

On Analyzer screen, go to System ►Temperature, assure that temperatures for Incubators, probes, and cooler are with in the specifications. Check off the temperature boxes in the maintenance logbook. If temperatures are out of range, notify the supervisor, and call IL service.

##### E.	Complete the Maintenance Log: Record all maintenance on the maintenance log and initial. Record any problems on the problem log, initial and date.

Weekly Maintenance

1. choose Menu Bar ► System ► Maintenance.

- Select Wipe-down Sample Probe.

- Single-click or touch the Perform Activity icon .

- The following message appears

- Open the sample cover.

- Wipe the probe tip and exterior with an alcohol wipe.

- Close the sample cover and click OK.

- Enter your initials in the Comment window then select OK.

- Wait for the instrument to reinitialize.

Monthly Maintenance: Clean Air Filter

- Turn the power off on the Analytical Module.

- Remove the front filter by sliding it forward.

- Rinse the filter under running water with the dirty side down.

- Allow the filter to dry completely.

- Remove the filter on the rear (located in the center holder) of the AM by sliding it sideways to the right

- Rinse the filter under running water with the dirty side down.

- Allow the filter to dry completely.

- Replace both with an extra dry filter.

- Power up the analyzer.

- Wait for the instrument to reinitialize.

- Check off the appropriate boxes on the log sheet in the maintenance notebook; date and sign with your initials in the appropriate spaces. Also, check off due maintenance log in the instrument.

Operational Checks and Preventive Maintenance of instruments and equipment

Preventive maintenance procedures will be performed once the first year and followed by 6-month intervals each year by the field service engineer from IL.

Any problem calls for the instrument should be made to IL technical consultant service at 1-800-678-0710. Be prepared to give the instrument serial number written on the IL instruments and a description of the problem. Document problems in the problem log

- CALIBRATION

Instrument Set-Up and Calibration (only if necessary): Calibration or recalibration frequency is based on Policy HEM 179 (Calibration and Analytical measurement Policy)

Calibration and storage of a valid specific Special Coag Assays calibration curve are required to obtain test results. Calibration is performed:

- With a change of reagent lot numbers

- With a change of major instrument components

- To satisfy local regulatory requirements

- At laboratory discretion

- ALWAYS check maintenance log before calibration and make sure all maintenance is current (not overdue) and replace Factor diluent with fresh from a new bottle.

- Choose Setup, Materials List, Click Scan and Scan to 2D barcode on the top of the box of the calibrator if a new lot. This will upload all the information about lot number, expiration date, and assay values (Skip step 3 – 7below). Repeat for all reagents.  If 2D barcode is not on the box, double-click on the appropriate calibrator to open the Materials Definition screen.

- Choose the Lot Specific Information tab and enter the Calibrator Lot Number and Expiration Date.

- Enable Lot Management from the Lot Specific Information tab.

- Select the Save icon to store the lot number. Once the lot number is saved, the Assign Values icon becomes available.

- Select the Assign Values icon.

- Enter the calibration value from the package insert.  Press OK.

- Choose the Previous Screen icon to exit.

- Load the specific test Reagents, Calibration Plasma, Factor Deficient Plasma(s) if Factor assay, Factor Diluent, and Diluted Clean B onto the instrument.

Note: Always use fresh Factor Diluent on-board if calibrating assays.

Refer to test feasibility screen for loading of reagents/ calibrator and controls.

- Select Calibration, Status List.

- Double-click on the appropriate test Assay to be calibrated to open the Calibration Details screen.

- Choose the Run icon.

- Select OK at the “Do you confirm the operation?” prompt.

- Choose the Previous Screen icon to exit.

- Verify the Job Status for the Assay test code says Active.

- Once the calibration is complete, review calibration results.  The Instrument will fail the calibration if the r2 value is less than the required value. See specific procedures for each assay.

- Choose the Calibration Information tab to ensure that no errors or warnings.   If there are no errors/failures or flag and the calibration is acceptable, choose the Validate icon to validate the calibration curve.

- Always print Calibration Curve and put it in the ACL TOP Calibration binder with initial and date.

- TO ACCESS FEASIBILITY SCREEN

    - Select the Sample Area , Diluent Area or Reagent Area   icon in the toolbar to open the Sample, Diluent or Reagent Area.

    - Select Actions > Review > Test Feasibility List in the menu bar, or select the Test Feasibility

List   icon in the toolbar.

    - In the Test Feasibility List window, view the information on the Test Feasibility List and QC Feasibility List tabs.

Test Feasibility List window

Test feasibility list window contains the Test Feasibility List tab and the QC Feasibility List tab

Use the Test Feasibility List tab to determine the feasibility of tests to run on samples.

- QUALITY CONTROL

Reconstitute reagent(s) and controls and allow sitting at room temperature per recommendation; swirling to mix just before loading onto the instrument

Refer to test feasibility screen for loading of reagents, calibrators, and controls.

How to program and Run QC

- Place QC in Diluent Rack (with Factor Diluent) ► Load rack in D1 and D2. Before loading QC rack, make sure the analyzer is in Ready mode.

- Go to Menu Bar ► QC ► Result List ► double click on any QC ► QC Statistic screen

- Clear the previous selection by single-clicking box next to “Material Tests” to remove checkmark or else all tests from the last run will be repeated (system remembers last QC programming).

- Select the checkbox next to the tests to order the QC.

- Single-click or touch Program QC   to start the analysis.

- To Review QC, a single click on the Previous screen (back arrow)  will return to QC Result list.

- If the control is acceptable, click on the   data point, click on the comment icon  , and type your initials in the comment box. If control is outside the acceptable range, the Status of the QC in red ‘failed’ and QC alarm at the bottom will alert you. Take an appropriate QC corrective action below.

- Controls should be prepared and tested once each 8-hour shift and tested again whenever reagents are added or changed and after each new calibration curve. Tech has to review shift control and placed an initial in the comment box under each control.

- Controls should be run in the same manner as the test samples, and by all techs that perform special coagulation testing.

- Control tolerance limits--the range is calculated based on +/-2SD from the mean control value.

- Corrective action for Unacceptable QC

    - Rerun control after swirling QC and reagents.

    - If still out, check reagent expirations; make new if indicated. Ensure fresh Factor Diluent is on-board and change if necessary (if not changed during set up). Perform Enhanced clean for all the probes.

    - If control still out, prepare new controls or reagents depending on one level of QC is out or both levels, allow to sit for 20 minutes, mix gently, and rerun.

    - If controls still out, Recalibrate the assay and notify the supervisor.

    - Remove the results that are outside the acceptable range by clicking on the unacceptable point and then clicking the omit icon.  On the next data point, indicate the corrective action that was performed in the comment box along with your initials. The control results are recorded in the ACL TOP 750 QC files and are reviewed monthly by the supervisor.

Note: All ACL TOP assay controls for the ACL TOP 750 are not formatted in the BEAKER QC program but set up and reviewed in the instrument QC Software file.

Review of QC Statistics

    - Single click on the words for any QC level on the Material/Tests Definition Tree.

NOTE: The checkmark before the QC level is only used to select which QC levels to run. The QC level backlit either in gray or blue is the QC level that is displayed to the right.

2. Single-click or touch the Switch View icon to switch between the following views:

Toggles from the current view to Chart view (AKA Levey- Jennings).

Toggles from the current view to Grid view (AKA Table of results)

Toggles from current view to Chart & Grid view.

3. Choose another QC level to view by single-clicking on the words.

Review Chart View (AKA Levey-Jennings)

1. Use Switch View to display the Levey-Jennings.

Review QC Event Lines on Levey Jennings

Vertical colored Event lines on the QC Levey Jennings display indicate a change has been made to this particular test/control combination. The event lines are color-coded (Yellow, Blue, Green, Red), with the colors corresponding to a particular type of QC Setup change.

- Open the QC Statistics for any QC.

- Verify the Levey-Jennings is displayed. You may need to single click

or touch Chart/Grid View .

- Place your mouse on any line on the chart to display the Event

4. Look through each QC file to find examples of each type of event.

Review Chart/Grid View

  - Single-click or touch Switch View to return to the Chart/Grid View (table of results/Levey-Jennings together).

- Select a data point on the table of results.

Add Comment to a QC Data Point

1. find a QC that has a failed result in any view mentioned above

- Select the data point.

- Single-click or touch Comment .

- Enter “Repeat the same vial of control.” and your initials then choose OK.

Omit and Restore a QC Data Point

    - Verify the Levey-Jennings is displayed for any level.

    - A single click on the last data point (far right) to Select it.

    - Single-click or touch Omit to omit the result from the statistics.

    - Choose OK at the Do you confirm to omit selected result? Prompt.

    - Enter a comment as to why the result was omitted in the Comment dialog box and choose OK.

    - Single-click or touch Switch View to display the Split Screen View.

    - A in the Omit column indicates the result has been removed from the statistics.

- Select the same point again then choose Action ► Results ► Restore Point ► OK to restore it.

- Add an additional comment in the Comment dialogue box as to why and choose OK.

Review Comments

- Use the scroll bar to scroll over to the Comment column.

- Click on the Comment column header twice to move all data points with comments to the top of the list.

- Hover over any check in this column to see the comment.

Review of Quality Control:

- Technologist Responsibilities

- Control results are evaluated at the time of assay by the technologist who is performing the assay. It is the responsibility of the technologist to notify the Coagulation specialist immediately when QC failures occur.

- All QC violations must be recorded electronically on Instrument QC Software or written in the problem log sheet.

- Coagulation Specialist QC Responsibility

Special coagulation test controls are reviewed --- by the Coagulation Specialist from the QC libraries on ACL TOP instruments. Monthly, the computer control libraries are printed out and reviewed by the Coagulation Specialist. The acceptable variability should be checked for all the coagulation instruments to validate comparable precision by accessing day to day, run to run, as well as operator variance. The acceptable total coefficient of variation of the analytical system (within-device CV) on the same lot of control plasma shown in Appendix 2. If any of the tests CV is greater than the limit %, appropriate action is taken based on QC trends and shifts, Calibration verification is required, and QC ranges are adjusted as needed. The acceptable criteria for adjusting QC ranges are

1) a new lot of reagent, 2) a new lot of calibrator or controls and,  3) major Instrument maintenance. If controls show an unusual trend and cannot be corrected by maintenance or troubleshooting, contact TAC support and request for service.

- RUNNING SAMPLES

- To Run Patient Samples with barcode

  - Place sample tubes in a sample rack with barcodes facing outwards.

  - Select an available sample track and load the sample rack when the barcode reader is in position. It will upload the patient ID. Order the test manually. Double click on the box to the right.  Choose the test in the Tests and profile box. For example, factor [FII(P), FV(P), FVII(P), FX(P)] will be under the PT Factors tab, and AT LIQ, PRC will be under Thrombophilia tab in the Tests and Profiles box.

  - Verify the samples have been identified and have a test ordered. If not, program the sample ID manually, order the test manually and order the Parallelism from the test and programming window.

  - Choose the Run icon if an instrument is not currently running.

- To Run Patient Samples without barcode

  - Place sample cup in sample rack and label with sample name.

  - Click on the sample area icon.  Double click on the rack to the left.

  - Enter the sample ID.

  - Double click on the box to the right.  Choose ATIII Liq or PRC tests under thrombophilia panel or Plasminogen or Factor XIII Ag under Specials.

  - Click the insert rack icon.  Load into an available track, S1-S12.

  - If the instrument is currently running and the run icon is greyed out, the sample(s) will be added to the active list and will be run.  If the run icon is purple, click it to start the test(s).

- REPORTING RESULTS

Refer to individual Special Coag Assay procedures for reporting results.

Note: Record results in the computer system; Post results through the outstanding list / manual reporting referring to Beaker bench manual as needed.

Hemolyzed, lipemic, or icteric samples must be noted with the result.

- REFERENCE INTERVALS

Normal range data for the adult population was validated by the Hematology lab from the hospital and non-hospital patients while the age-specific ranges are from the literature available by

Age dependency for coagulation parameters in paediatric populations-Pierre Toulon, Micheline Berruyer,Marie Brionne-Francois Grand, Dominique Lasne, Caroline Telion, Julien Arcizet, Roberta Giacomello, Neila De pooter; Thromb Haemost 2016; 116; 9-16

Development of the hemostatic system in the neonate and infants. Am J pediatr Hematol Oncol 12:95.1990

See below “Appendix 1” for Reference ranges or Refer to the individual procedure.

- CRITICAL VALUES

Refer to Critical Call Policy Hematology Laboratory for calling critical results and documentation

| Test | Critical Value | Time Frame to Notification | Frequency |
| --- | --- | --- | --- |
| PT/PTT Factor assays | <5.0 % | <30 minutes | First occurrence |
| INHIB | >0.5 BU | <30 minutes | Each occurrence |

- LIMITATIONS OF THE PROCEDURE:

Refer to individual procedure for each assay for interference

Limitations and Interference substances

Samples with excessive hemolysis, icterus, or lipemia, should not be used. Also, refer to individual assay package insert/procedure for relative interference claims.

Analytical limitations:

On the ACL TOP instrument, sample carryover1 is negligible. In most situations, inaccuracy attributed to carryover is within the normal imprecision of the method, and therefore not statistically or clinically significant.

- COURSE OF ACTION WHEN TEST SYSTEM UNAVAILABLE AND OR NOT USEABLE

Leave a note on the instrument stating the instrument down.

Call to IL technical assistance center at 1-800-678-0710. Be prepared to give the instrument serial number written on the IL instruments and a description of the problem. Document problems, Reference number, and any troubleshoot suggested by an expert in the problem log (Appendix A).

- PROFICIENCY TESTING

We are currently enrolled in CAP Survey Proficiency Testing. See Appendix 4

- PERFORMANCE SPECIFICATION FOR ACCURACY, PRECISION, ANALYTICAL SENSITIVITY AND SPECIFICITY

See Approved Precision. Accuracy, Reference Ranges, linearity, and Correlation study on L-Drive.

L:\Hematology\4 - Coags\2019 IL  Installation\IL Validation Data\Method Validation for special Coag instruments\Post validation

- REFERENCES

- ACL TOP® Family On-Line Help Manual

- Clinical and Laboratory Standards Institute. Collection, Transport, and Processing of Blood Specimens for Testing Plasma-Based Coagulation and Molecular Hemostasis Assays; Approved Guideline - Fifth Edition, CLSI Document H21-A5; Vol. 28 No. 5

- Westgard JO, and Barry PL. Cost-Effective Quality Control; Managing the Quality and Productivity of Analytical Process, AACC Press, 1986

- Clinical and Laboratory Standards Institute.  One Stage Prothrombin Time (PT) Test and Activated Partial Thromboplastin Time (APTT) Test; Approved Guideline – Second Edition, CLSI Document H47-A2; Vol. 28 No. 20

- HemosIL Factor II, V, VII, VIII, IX, X, XI, XII, AND XIII ANTIGEN Deficient plasma package insert.

- HemosIL RecombiPlasTin 2G (PN 0020002950/0020003050) package insert

- HemosIL Calibration Plasma (PN 0020003700) package insert.

- Reference Clinical and Laboratory Standards Institute.Preparation and Testing of Reagent Water in the Clinical Laboratory; Approved Guideline. Fourth, Edition, CLSI Document C3-A4;Vol.26 No.22

- Age dependency for coagulation parameters in paediatric populations-Pierre Toulon, Micheline Berruyer,Marie Brionne-Francois Grand, Dominique Lasne, Caroline Telion, Julien Arcizet, Roberta Giacomello, Neila De pooter.

- HISTORY

This procedure was written by P Bahel on 9/12/2019

| Appendix 2 |  |  |  |
| --- | --- | --- | --- |
| ANANLYTICAL PERCISION PERFORMANCE REQUIREMENTS |  |  |  |
| Assay | Result Unit | Control | Total Imprecision (%CV) |
| Extrinsic Factors (II, V, VII, X) | % Activity | HemosIL Normal Control | ≤ 10.0 |
|  |  | Special Test Control Level 2 | ≤ 14.0 |
| Inrinsic Factors (II, V, VII, X) | % Activity | HemosIL Normal Control | ≤ 10.0 |
|  |  | Special Test Control Level 2 | ≤ 14.0 |
| Factor XIII Antigen | % | HemosIL Normal Control | ≤ 12.0 |
|  |  | Special Test Control Level 2 | ≤ 14.0 |
| Von Willebrand Factor Activity | % Activity | HemosIL Normal Control | ≤ 10.0 |
|  |  | Special Test Control Level 1 | ≤ 12.0 |
| Von Willebrand Factor Antigen | % VWF:Ag | HemosIL Normal Control | ≤ 7.0 |
|  |  | Special Test Control Level 2 | ≤ 15.0 |
| Antithrombin | % Activity | HemosIL Normal Control | ≤ 8.0 |
|  |  | Special Test Control Level 2 | ≤ 15.0 |
| Protein C | % Activity | HemosIL Normal Control | ≤ 6.0 |
|  |  | Special Test Control Level 2 | ≤ 12.0 |
| Protein S Activity | % Activity | HemosIL Normal Control | ≤ 8.0 |
|  |  | HemosIL Low Abnormal | ≤ 12.0 |
| Free Protein S | % Activity | HemosIL Normal Control | ≤ 10.0 |
|  |  | Special Test Control Level 2 | ≤ 12.0 |
| Plasmin Inhibitor (Alpha-2-Antiplasmin) | % Activity | HemosIL Normal Control | ≤ 8.0 |
|  |  | Special Test Control Level 2 | ≤ 12.0 |
| Plasminogen | % Activity | HemosIL Normal Control | ≤ 8.0 |
|  |  | Special Test Control Level 2 | ≤ 12.0 |
| Factor V Leiden (APCR-V) | Ratio | APC Control Plasma Level 1 | ≤ 10.0 |
|  |  | APC Control Plasma Level 2 | ≤ 10.0 |
| dRVVT Screen dRVVT Confirm | Seconds | LA Positive Control | ≤ 6.0 |
|  |  | LA Negative Control | ≤ 6.0 |
| Silica Clotting Time | Seconds | LA Positive Control | ≤ 6.0 |
|  |  | LA Negative Control | ≤ 6.0 |