|  |  |  |  |
| --- | --- | --- | --- |
|  | TITLE:  **HemosIl von Willebrand Activity-- ACL TOP 750** | | **DEPT OF LAB MEDICINE**  **CLINICAL HEMATOLOGY**  **Policy and Procedure Manual** |
| **DOCUMENT # HEM218** |
| **WRITTEN BY:**  Parveen Bahel, MT(ASCP) | **EFFECTIVE DATE:** | **REVIEW/REVISION**  H-1 (New); 10/2019 | **Pages 1-12** |

1. **Intended Use** Automated latex enhanced immunoassay for the quantitative determination of von Willebrand Factor Activity (VWF Activity) in human citrated plasma on IL Coagulation Systems.
2. **Purpose**

This procedure provides instructions for the quantitative determination of von Willebrand Factor Activity (VWF Activity) in human citrated plasma using HemosIL von Willebrand Factor Activity Kit on ACL TOP® Family.

1. **Summary and Principles**

The diagnosis of von Willebrand Disease (VWD), probably the most common congenital bleeding disorder, requires a number of special tests at the laboratory level. The measurement and comparison of von Willebrand Factor Antigen (VWF:Ag), VWF Activity and Factor VIII (FVIII) levels in plasma aid in the differentiation of quantitative defects (type 1 or type 3) or qualitative defect (type 2) of VWF and therefore to diagnose the different types of VWD.

When an extremely low or undetectable level of VWF:Ag is obtained, a type 3 VWD could be expected. If a moderate or even normal result is obtained, VWF Activity and FVIII assays must be performed and compared with the VWF:Ag result. If all three values are within the normal range, VWD and Hemophilia A may be excluded. If at least one parameter is abnormally low, it is necessary to calculate the ratios VWFActivity/VWF:Ag and FVIII/VWF:Ag. If both ratios are close to 1 (some authors suggest 0.7 as cut-off), a VWD type 1 may be diagnosed.1-2 When the VWF/Activity/VWF:Ag ratio is low (0.7 is also the suggested cut-off), types 2A, 2B or 2M may be diagnosed. These subtypes are characterized by its abnormal multimeric pattern and/or its altered platelet affinity.1-2 Additional laboratory tests as RIPA (Ristocetin Induced Platelet Aggregation), multimeric analysis and binding assays are required in order to be able to distinguish the different subtypes.

When the FVIII/VWF:Ag ratio is low (0.7 is also the suggested cut-off), a type 2N or Hemophilia A may be diagnosed and a FVIII binding assay is necessary to discriminate among them.

The VWF Activity kit is a latex particle enhanced immunoturbidimetric assay to quantify VWF Activity in plasma. The activity of VWF is determined by measuring the increase of turbidity produced by the agglutination of the latex reagent. A specific anti-VWF monoclonal antibody adsorbed onto the latex reagent, directed against the platelet binding site of VWF (Glycoprotein Ib receptor), reacts with the VWF of patient plasma. The degree of agglutination is directly proportional to the activity of VWF in the sample and is determined by measuring the decrease of transmitted light caused by the aggregates.

1. **Interpretation of Results**

The diagnosis of von Willebrand Disease (VWD) requires a number of special tests. The measurement and comparison of von Willebrand Factor Antigen (VWF:Ag), VWF Activity and Factor VIII (FVIII) levels in plasma aid in the differentiation of quantitative defects (Type 1 or 3) or qualitative defect (Type 2).

When an extremely low or undetectable VWF:Ag result is obtained, a Type 3 VWD could be expected.

If a moderate or even normal result is obtained, VWF Activity and FVIII assays must be performed and compared with the VWF:Ag result.

If all three values are within the normal range, VWD and Hemophilia A may be excluded.

If at least one parameter is abnormally low, it is necessary to calculate the ratios VWF Activity/VWF:Ag and FVIII/VWF:Ag.

If both ratios are close to 1 (some authors suggest 0.7 as a cutoff), a VWD Type 1 may be diagnosed.

When the VWF/ACT/VWF:Ag ratio is low (0.7 is also the suggested cutoff), Types 2A, 2B or 2M may be diagnosed. These subtypes are characterized by its abnormal multimeric patterns and/or altered platelet affinity. Additional specialized tests such as Ristocetin Induced Platelet Aggregation (RIPA), multimeric analysis and binding assays are required for differentiation between different subtypes.

When the FVIII/VWF:Ag ratio is low (0.7 is suggested cutoff), a Type 2N or Hemophilia A may be diagnosed and a FVIII binding assay is necessary to discriminate between them.

1. **Specimen Type**

Mix nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant.

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing of blood and anticoagulant.

A syringe or evacuated tubes (blue top) may be used for collection. If multiple specimens are collected; the coagulation sample should be the second or third tube collected. If only coagulation testing is to be performed, a red-top tube, which has no additives, should be drawn first and discarded prior to drawing the blue-top coagulation tube.

The patient cannot be on anti-coagulants when the test specimen is collected. Sufficient time after discontinuance of heparin should be allowed for heparin to be cleared from the patient’s blood, usually 6 hours.

If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values above 55%.

Specimens that are clotted, collected in the wrong tube or serum, overfilled or have less than the 90% expected fill should be rejected.

.

1. **Handling Condition and Stability**

The whole blood specimen is checked for clot formation by gentle inversion and observation. Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x109/L for **10 minutes at 4000 g.** Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells. For special coagulation testing, spin samples 20 minutes at 4000 g, separate plasma into plastic tubes, label and freeze all aliquots at –70C located in the Special coag area until ready to use. Always check samples for clots after aliquoting. Always track aliquots in BEAKER under YH Coag Hold before freezing them. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37°C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4°C until tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4°C until tested.

**Specimen stability** at ambient temperature: 4 hours; Frozen at -70° C 6 months.

**Specimen Labelling:** Specimen should be properly labelled with at least 2 unique patient identifiers

1. **Environmental Operating Conditions**

The instrument functions correctly in an ambient temperature of 15° C to 32° C (59° F to 89° F) with a relative humidity of 15% to 85% (non-condensing).

In accordance with the IEC regulations, no instrument failures occur in the presence of short-term ambient temperatures as low as 5° C or as high as 40° C.

The ACL TOP Family 50 Series is compliant with IEC 60068-2-40 to 2000 meters. The instrument should not be used at an altitude greater than 2000 meters.

The instrument should be placed in an area free from dust, fumes, vibrations and excessive variations of temperature.

The heat generated by the instrument during normal operation is exhausted from the bottom, the front-right and the left side of the unit.

According to IEC 61010-1, the maximum audible noise emission should be 80 dBA. The ACL TOP Family 50 Series is compliant with IEC 61010-1 Third Edition.

The room temperature and humidity percent are monitored and documented on the Routine Coagulation checklist.

1. **Equipment and Materials**
   1. **Supplies**
      * Nerl Water: pH 7.0
      * Gauze
      * Citrated blue top tubes
      * Frosted tubes for aliquots
      * Cuvettes
      * ACL Top sample cups
      * ACL TOP 750
   2. **Reagents**

* HemosIL Calibrator Plasma
* HemosIL Normal control Assayed
* HemosIL test Control level 1
* HemosIL von Willebrand Factor Activity kit which contain latex reagent and Buffer
* HemosIL Cleaning solution Clean A and Clean B
* HemosIL Rinse and waste
* HemosIL factor Diluent

1. **Product Information**

The **HemosIL von Willebrand Factor Activity** kit (PN 0020004700) consists of:

**Latex Reagent:** 2 vials of a lyophilized suspension of polystyrene latex particles coated with purified anti-VWF mouse monoclonal antibody directed against a functional epitope of VWF. It also contains bovine serum albumin, stabilizers and preservative.

**Buffer:** 2 vials of Tris buffer containing bovine serum albumin, stabilizers and preservative

1. **Reagent Preparation**

**Buffer:** Ready for use.

**Latex Reagent:** . Dissolve the contents of each vial by pouring the entire contents of one vial of Buffer into one vial of Latex Reagent. Ensure that all drops of Buffer are transferred. Replace the stopper and swirl gently for a minimum of 20 seconds to completely dissolve the lyophilized latex.

Ensure complete reconstitution of the product. It must appear as a homogenous and slightly milky suspension. Keep the reagent at 15-25˚C for 30 minutes and invert to mix before use. Do not shake.

**Note:** Avoid foam formation when homogenizing reconstituted reagents. Bubbles on top of the liquids may interfere with the instrument liquid sensors.

**Cleaning Agent** (Clean B Diluted): Make fresh Clean B Diluted every day, 1 Part Clean B + 7 parts of Nerl water.

**HemosIL Calibrator Plasma (Lyophilized):** Reconstitute with 1 mL of Nerl water. Used for calibration, if needed.

**HemosIL Normal control Assayed** **(Lyophilized):** Reconstitute with 1 mL of Nerl water.

**HemosIL Special control level 1 (Lyophilized):** Reconstitute with 1mL of Nerl water.

1. **Reagent Storage and Stability**

Unopened reagents are stable until the expiration date shown on the vial when stored at 2-8°C.

For optimum stability, remove reagents and calibrator from the system and store them at 2-8oC in the original vial.

**Latex reagent-** Stability after reconstitution:

1 month at 2-8 oC in the original vial after opening or 5 days at 15°C on the ACL TOP. Do not freeze

**Calibrator, Control Storage:** Unopened calibration plasma and controls are stable until the expiration date shown on the vial when stored at 2-8˚C.

Stability of HemosIL Calibrator after reconstitution is 8 hours at 2-8˚C in the original vial. Use reconstituted calibrator within 2 hours for assay calibration

Stability of Normal and abnormal controls after reconstitution are 24 hours at 15°C on the instrument.

1. **Calibration Details**

Calibration or recalibration frequency is based on Policy # HEM 179 (Calibration and Analytical measurement Policy).

Calibration and storage of a valid VWF Activity calibration curve are required to obtain Von willebrand Activity results. Calibration is performed:

* With a change of reagent lot numbers
* With a change of major instrument components
* To satisfy local regulatory requirements
* At laboratory discretion


**Refer to test feasibility screen for loading of reagents, calibrator, and controls.**

**Steps to follow calibration:**

1. **ALWAYS** check maintenance log before calibration and make sure all maintenance is current (not overdue) and replace Factor diluent with fresh from a new bottle.
2. Choose **Setup, Materials List, Click Scan** and Scan to 2D barcode on the top of the box of the calibrator if a new lot. This will upload all the information about lot number, expiration date, and assay values **(Skip step c – g below).** Repeat for all reagents.If 2D barcode is not on the box, double-click on the appropriate calibrator to open the **Materials** **Definition** screen.
3. Choose the **Lot Specific Information** tab and enter the Calibrator Lot Number and Expiration Date.
4. Enable **Lot Management** from the Lot Specific Information tab.
5. Select the **Save** icon to store the lot number. Once the lot number is saved, the **Assign Values** icon becomes available.
6. Select the **Assign Values** icon.
7. Enter the calibration value from the package insert. Press **OK**.
8. Choose the **Previous Screen** icon to exit.
9. Load the VWD: Act Latex, Calibration Plasma, Factor Deficient Plasma(s), Factor Diluent and Diluted Clean B onto the instrument.

**Note: Always use fresh Factor Diluent on-board if calibrating VWF Activity.**

**Refer to test feasibility screen for loading of reagents/ calibrator and controls.**



1. Select **Calibration, Status List**.
2. Double-click on the VWF Act test code to be calibrated to open the **Calibration Details** screen.
3. Choose the **Run** icon.
4. Select **OK** at the “Do you confirm the operation?” prompt.
5. Choose the **Previous Screen** icon to exit.
6. Verify the Job Status for the VWF Activity test code says **Active.**
7. Once the calibration is complete, review calibration results. The Instrument will fail the calibration if the r2 value is less than 0.970.
8. Choose the Calibration Information tab to ensure that no errors or warnings. If there are no errors/failures or flag and the calibration is acceptable, choose the Validate icon to validate the calibration curve.
9. Always **print Calibration Curve** and put it in the ACL TOP 1 Calibration binder with initial and date.
10. **Quality Control**
    1. Load all appropriate Reagents VWD: Act Latex, along with Diluted Clean B onto the instrument. Before loading the reagent rack, make sure the analyzer is in Ready mode.

**Refer to test feasibility screen for loading of reagents/ calibrator and controls.**

* 1. Place HemosIL normal control assay and HemosIL Special control level 1 with the barcodes facing out in a Diluent Rack and load on the instrument in a Diluent track D1 or D2.
  2. Choose **QC** from the Main Menu and select **Test Status List**.
  3. Double-click on VWF ACT test code to reveal the Test Materials Definition tree in the **QC statistics screen**.
  4. Select the box in front of the VWF ACT QC box test and choose the **Program QC** icon. This will run all QC levels for that test.
  5. To Review QC, single click on **Previous screen (back arrow)**  will return to **QC Result list.**
  6. If the control is acceptable, click on the page5image3395804144**data** point, click on the **comment icon** page5image3395808928, and type your initials in the comment box. If control is outside the acceptable range, the Status of the QC in red ‘failed’ and QC alarm at the bottom will alert you. Take an appropriate QC corrective action below.
  7. Controls should be prepared and tested once each 8-hour shift and tested again whenever reagents are added or changed and after each new calibration curve. Tech has to review shift control and placed an initial in the comment box under each control.
  8. Controls should be run in the same manner as the test samples, and by all techs that perform special coagulation testing.
  9. Control tolerance limits--the range is calculated based on +/-2SD from the mean control value.

**Corrective action when tolerance limits are exceeded**:


* + 1. Rerun control after swirling QC and reagents.
    2. If still out, check reagent expirations; make new if indicated. Ensure fresh Factor Diluent is on-board and change if necessary (if not changed during set up). Perform Enhanced clean for all the probes.
    3. If control still out, prepare new controls or reagents depending on one level of QC is out or both levels, allow to sit for 20 minutes, mix gently, and rerun.
    4. If controls still out, Recalibrate the assay and notify the supervisor.
    5. Verify reagent performance.
    6. Check instrument performance
    7. Document actions taken to identify and correct the problem before reporting any patient data.
    8. Remove the results that are outside the acceptable range by clicking on the unacceptable point and then clicking the omit icon. On the next data point**, indicate** the corrective action that was performed in the comment box along with your initials. The control results are recorded in the ACL TOP 750 QC files and are reviewed monthly by the supervisor.
    9. If the problem cannot be resolved, call for Service if necessary and properly document in troubleshooting log.Notify supervisor.

**Note: VWF Activity** **controls for the ACL TOP 750 are not formatted in the BEAKER QC program but set up and reviewed in the instrument QC Software file.**

1. **Procedure:**
2. Load reagents onto the instrument. Calibrate, if necessary (see calibration section of this procedure).
3. Place QC materials with the barcodes facing out in a Diluent Rack and load onto an instrument Diluent track.
4. Choose **QC** from the Main Menu and select **Test Status List.**
5. Double-click on the **VWF ACT** test code item to reveal the **Test Materials Definition tree.**
6. Select the box in front of the **VWF ACT QC** Control and choose the **Program QC** icon. This will run all QC levels for that test. See
7. Place sample tubes in a sample rack with barcodes facing outwards.
8. Select an available sample track and load the sample rack when the barcode reader is in position.
9. Verify the samples **have been identified and have a test ordered. If not, program the sample ID ma**nually and/or order the test manually from the test and programming window.
10. Choose the **Run** icon if the instrument is not currently running.

**To Run Patient Samples without barcode**

* + - * Place sample cup in sample rack and label with sample name.
      * Click on the sample area icon. Double click on the rack to the left.
      * Enter the sample ID.
      * Double click on the box to the right. Choose the **VWF ACT** test under the Special tab in the Tests and Profiles box.
      * Click the **insert rack** icon. Load into an available track, S1-S12.
      * If the instrument is currently running and the run icon is greyed out, the sample(s) will be added to the active list and will be run. If the run icon is purple, click it to start the test(s).



1. **Reporting Results**

VWF Activity results are reported in % normality.

Linearity for this assay is 19-130%.

If the result is greater than the reportable range (130%) the sample will be automatically repeated with a 1:3 dilution.

The upper limit of reporting is 390%; if the result is higher than 390%, it will flag as “result above linear range” report “>390%” and any result lower than 19 will be reported as <19%.

Record results in the computer system; Post results through the outstanding list / manual reporting referring to the Beaker bench manual as needed.

Hemolyzed, lipemic, or icteric samples must be noted with the result.

**Reference Interval:** Normal range data for adult population was validated by the hematology lab from hospital and non-hospital patients while the age specific ranges are from the literature available by “*Age dependency for coagulation parameters in paediatric populations-Pierre Toulon, Micheline Berruyer,Marie Brionne-Francois Grand, Dominique Lasne, Caroline Telion, Julien Arcizet, Roberta Giacomello, Neila De poote:Thromb Haemost 2016; 116; 9-16”*

**Adult Normal Range: 58 – 163%**

**Age-specific reference ranges from Literature;**



**Reflex Criteria**: All abnormal results reflex MD Interpretation.

1. **Critical Results:** No critical result for the procedure.
2. **Procedural note:**

The overall performance of von Willebrand Factor Activity testing is dependent on reagent and instrument performance. Acceptable variability (imprecision) should be such that the total coefficient of variation (CV) of the analytic system is less than <=10% on the same lot of Normal control plasma and <=12% on the same lot of abnormal control plasma.

The measuring range is defined by the concentration of the calibrators used and the extrapolation limits set

1. **Specific Performance Characteristics**

Within-run and total (run to run and day to day) precision was assessed over multiple runs using both normal and abnormal control samples with a specific lot of PT reagent.

Please refer to the appropriate package insert for precision study results.

1. **Limitations and Interference substances**

VWF Activity results on the ACL TOP® Family are not affected by:

Hemoglobin up to 70 mg/dL

Bilirubin up to 4.2 mg/dL

Triglycerides up to 1020 mg/dL

Rheumatoid Factor up to 200 IU/mL

1. **References**
2. HemosIL von Willebrand Factor Activity package insert
3. ACLTOP® Family On-Line Help Manual
4. Clinical and Laboratory Standards Institute. Collection, Transport, and Processing of Blood Specimens for Testing Plasma-Based Coagulation and Molecular Hemostasis Assays; Approved Guideline - Fifth Edition, CLSI Document H21-A5; Vol. 28 No. 5
5. HemosIL Calibration plasma package insert
6. Westgard JO, and Barry PL. Cost-Effective Quality Control; Managing the Quality and Productivity of Analytical Process, AACC Press, 1986
7. Age dependency for coagulation parameters in paediatric populations-Pierre Toulon, Micheline Berruyer,Marie Brionne-Francois Grand, Dominique Lasne, Caroline Telion, Julien Arcizet, Roberta Giacomello, Neila De pooter, Thromb Haemost 2016; 116; 9-16
8. Development of the hemostatic system in the neonate and infants. Am J pediatr Hematol Oncol 12:95.1990
9. **History**

This procedure was written by P Bahel on 10/10/2019