|  | TITLE: ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT) (CS 5100--ACTIN FSL) |  | DEPT OF LAB MEDICINE CLINICAL HEMATOLOGY Policy and Procedure Manual |
| --- | --- | --- | --- |
|  |  |  | DOCUMENT # HEM 217 - 1 - |
| WRITTEN BY: Parveen Bahel H(ASCP) | EFFECTIVE DATE: 07/14/2020 | REVIEW/REVISION: H-1 (New) | Pages 1 of 18 |

- INTENDED USE

Liquid purified soy and rabbit brain phosphatides with plasma activator. An activated partial thromboplastin reagent with increased sensitivity to lupus-like inhibitors foR use in the determination of the activated partial thromboplastin time (APTT) and related coagulation procedures.

- SUMMARY AND EXPLANATION

The activated partial thromboplastin time, a global screening procedure used primarily to evaluate coagulation abnormalities in the intrinsic pathway, will also detect severe functional deficiencies in factors II, V, X, or fibrinogen. The APTT has also been widely advocated3-6 as a means to monitor the effectiveness of unfractionated heparin therapy, where the clotting time is prolonged in proportion to the level of heparin. In patients receiving oral anticoagulants, the circulating levels of factors II, VII, IX, and X are depressed, therefore the APTT can be expected to be prolonged. The presence of non-specific inhibitors, such as the lupus-like anticoagulant, may prolong the APTT but this effect is variable and generally recognized as being related more to the nature of the APTT reagent employed. In summary, the APTT is a clinically important screening test with wide applicability for the diagnosis of coagulant disorders and therapeutic monitoring of both hemorrhagic and thrombotic disease.

- PRINCIPLE:

Factors of the intrinsic coagulation system are activated by incubating the plasma with the optimal amount of phospholipids and a surface activator. The addition of calcium ions triggers the coagulation process, and the clotting time is then measured.

- SPECIMEN REQUIREMENTS

The only acceptable sample for routine coagulation is plasma, from whole blood drawn into blue-stoppered tubes containing 3.2% sodium citrate (nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant).

- 4.5 ml-draw tube must have a MINIMUM ACCEPTABLE VOLUME of 4.0 ml total (3.5 ml blood +0.5 ml anticoagulant).

- 2.7 ml-draw tube must have a MINIMUM of 2.5 ml (2.2 ml blood + 0.3 ml anticoagulant).

- A “micro method” sample (newborn PT/PTT/FIB only) must be a full 1.0 ml specimen (0.9 ml of whole blood drawn with a syringe added to 0.1 ml 3.2% sodium citrate anticoagulant.)

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing of blood and anticoagulant.

A syringe or evacuated tubes (blue top) may be used for collection. If multiple specimens are collected; the coagulation sample should be the second or third tube collected. If only coagulation testing is to be performed, a red-top tube, which has no additives, should be drawn first and discarded prior to drawing the blue-top coagulation tube.

The patient cannot be on anti-coagulants when the test specimen is collected. Sufficient time after discontinuance of heparin should be allowed for heparin to be cleared from the patient’s blood, usually 6 hours.

If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values above 55%.

Specimens that are clotted, collected in the wrong tube or serum, have less than the 90% expected fill should be rejected, the floor must be notified, a comment entered into the computer.

Handling Condition and Stability

The whole blood specimen is checked for clot formation by gentle inversion and observation. Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x10 9/L for 10 minutes at 4000 g or 3 minutes at 8000 rpm (STAT centrifuge) for STAT samples or 7 minutes at 5200 RPM as soon as possible after collection. If immediate testing is to be done, the plasma may remain on the packed cells or separated.

Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells. For special coagulation testing, spin samples 20 minutes at 4000 g, separate plasma into plastic tubes, label and freeze all aliquots at –70C located in the Special coag area until ready to use. Always check for clots after aliquoting with applicator sticks. Always tracks aliquots in BEAKER under YH Coag Hold before freezing them. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested.

Specimen stability at ambient temperature: 4 hours; Frozen at -70° C 6 months.

Specimen Labelling Specimen should be properly labeled with at least 2 unique patient identifiers. Patient's full name and medical record number (MRN) and should have the date and time of collection. The patient's birth date may substitute for the MRN if the MRN is not available.

- ENVIRONMENTAL OPERATING CONDITIONS

The room temperature and humidity percent are monitored and documented on the Routine Coagulation checklist.

- EQUIPMENT AND MATERIALS

Equipment:

    - Nerl Water: pH 7.0

    - Gauze

    - Citrated blue top tubes

    - Frosted tubes for aliquots

    - Cuvettes

    - SLD Mini - cups

    - CS5100 sample cups – 4 mL

    - CA Clean I – 50 mL (On-board stability – 5 days)

    - CA Clean II – 5 L (On-board stability – 2 months)

    - OVB (Buffer) (On-board stability – 24 hours)

    - Trash bags

Materials:

1.	Dade Actin FSL and CaCl2

2.	Control material: Ci-Trol Coagulation Controls, Levels 1 and 3

- PREPARATION OF REAGENTS

NOTE: After opening any vial to place onto the instrument, label that vial with the open and expiration date, referring to the stability information provided here.  Discard reagent when expired.

- Siemens Actin FSL Activated PTT Reagent -purified soy and rabbit brain phosphatides in 1 x 10-4 M ellagic acid with added buffer, stabilizers and preservative.

- The reagent is liquid and ready for use.

- Mix by inversion immediately before use (6-8 times).

- Store at 2-8C when not in use. Do not freeze.

- Stable for 72 hours on board the CS 5100 (10 C)

- B.	Siemens Calcium Chloride (0.025M)

- The reagent is liquid and ready to use.

- Store at 2-8C when not in use.

- Stable for 72 hours on board the

C.	Siemens Ci-Trol Level 1 (Normal control) –Lyophilized, preparation of fresh human, normal citrated plasma. Used for Quality Control.

D.	 Siemens Ci-Trol Level 3 (Abnormal control)-Lyophilized, preparation of fresh human, abnormal citrated plasma. Used for Quality Control.

- Reconstitute Citrol 1 and Citrol 3 with 1 mL NERL water

- Re stopper vial and allow to stand for 15 minutes until dissolved.

- Invert gently to mix. Do not shake.

- Stability after reconstitution of Citrol 1 and 3—24 hours on board (10 C) the CS 5100 instrument

- PERFORMANCE PARAMETERS

Indication of deterioration: Normal plasma or controls will show deviations in results from the established laboratory range.

- LOADING REAGENTS and QC

- Reagent table loading:

Note: Always load reagent in the orange racks (A racks) and the control and calibrators in the yellow racks (C racks).

Reagents cannot be loaded while running the patients.

- Press Reagent on the toolbar.

- Verify system is in solid green ready status.

- Open light shield lid.

- Open reagent table cover.

- Remove expired or empty reagent vials by pressing Change/Add icon. It will bring the rack in front you want to remove reagents.

- Place new reagents of Actin FSL and Calcium Chloride in a rack with barcode facing out.

- Load the reagent rack.

- Press and lock reagent table cover.

- Press OK on screen to barcode read.

- Wait for barcode reading to read the barcode.

- On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

B.	Loading the Buffer

1. Press Reagent on the toolbar.
2. Verify buffer table cover LED is green.
3. Open buffer table cover.
4. Place OV Buffer in an adaptor with barcode facing out.

5. Place into the buffer table and close the cover.
6. Press OK on the barcode reading screen.
7. On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

C.	Loading QC

C-rack and SLD mini cup

- Reconstitute Citrol 1 and Citrol3 QC with 1 mL of Nerl water.

- Aspirate entire content of QC vial (GW5) using transferring pipette.

- Slowly dispense the entire volume into a new SLD mini cup, avoiding any air bubble formation.

- After reagent transfer, carefully check for bubbles and, if required, remove using smaller pipette tip.

- Set SLD mini cup into corresponding GW5 vial.

- Set GW5 with SLD into yellow-handled C-rack. Push down on vial to ensure proper seating in position.

- Place the rack back in reagent table and close lid.

- Lock lever and Press OK to read barcode.

- On Reagent screen, highlight vial just loaded and press Change to update date/time.

Note the SLD mini cups are for controls and calibrators only.

- How to order QC:

Processing from the reagent table: Control lot

1. Load QC onto a C-rack.

Refer to the Loading Reagents tab to load QC material in a C-rack using a SLD mini cup.

2. Select Order.

3. Select Switch Order.

4. Select Holder QC Order

5. Press Order Entry.

6. Select control material from the list.

7. Place cursor in the Lot No. field and select lot from list.

8. Select the appropriate assay(s) – Citrol 1 for PTT

9. Press the down arrow to order the next control – Citrol 3 for PTT

10. Press OK once controls have been ordered.

11. Press Start.

- How to review qc and add a comment or initial to a point:

1. Press QC Chart.
2. Highlight PTT QC Chart.

3. Place cursor on point of Citrol 1 and Citrol 3.
4. Press Confirm Data key.
5. Enter your initial or comment in field if there is any QC issue.
6. To exclude point: check box.

7. Press Update.

1.	Controls should be tested once each 8-hour shift, and after reagent changes / additions. Tech has to review shift control and placed an initial in QC Chart Software under each control. Also QC should be documented in QC log in the L: Drive

2.	Controls should be run in the same manner as the test samples and by the same technologists who work in the routine coagulation area.

3.	Control tolerance limits--the range is calculated based on +/- 2 SD (Standard deviations) from the mean control value.

4.	Corrective action when tolerance limits are exceeded (results will be flagged with error message):

a.	Rerun control after swirl QC and reagents.

b.	If still out, Check reagent/QC for bubbles and turbidity.

c.	Check reagents/ QC expiration. Ensure fresh OVB buffer is on-board and change if necessary (if not changed during set up). Prime the instrument.

d.	Make fresh reagent or QC if needed and rerun again

e.	Document actions taken to identify in QC Software and correct the problem before reporting any patient data.

f.	If controls still out, Recalibrate the assay if required and notify supervisor.

Remove the results that are outside the acceptable range by clicking on the unacceptable point and then clicking Confirm Data Key. On the next data point, indicate the corrective action that was performed in the comment box along with your initials. The control results are recorded in the CS 5100 QC files and QC log on L: Drive and are reviewed monthly by the supervisor.

QC Corrective Action:

NOTE: If QC is out of range after running twice, corrective action and an evaluation need to be performed. Up to 5 patient samples when QC was acceptable will need to be rerun to determine if there is a significant clinical difference in the patient results. Documentation is to print results from the Instrument of the repeated run of the patient samples, fill out corrective action log; print Beaker sample label and file in Corrective Action notebook. Acceptable agreement should be +/- 10% of the original results (See Appendix A for log). If significant clinical changes have been determined, there will be an amended report with corrected data and a caregiver will be notified promptly. In these circumstances, the analyzer will be put out of service and troubleshooting/service dispatched. Continue to run samples on the other functional instrument, if needed set up CS 5100-3 instrument for routine coag tests.

- If this occurs during off-shifts and the immediate concern is to determine if patient data has been compromised, the patient samples from the problem instrument should be rerun on the other functioning instrument and results compared. See corrective action log for acceptable criteria. Troubleshooting will then be performed timely on the problem instrument, QC run and evaluated and patient comparisons performed before instrument is put back into service.

Note: All Routine Coag Assays controls for the CS 5100 are not formatted in the BEAKER QC program but set up and reviewed in the instrument QC Software file.

- Loading patient samples on the cs 5100

- LIS order processing

- Load bar-coded samples into any sample rack (7-20) with the bar-code label visible through the window of the rack position and Place rack on the sampler.

- Check host connection (HC) status. HC status icon must be green or orange.

- Press Start.

- After barcode reading, confirm sample order status and progress on the Joblist screen.

- Manual order processing

- Place rack with sample tubes on sampler.

- Press Order.

- Enter rack number

- Select tube position to input an order.

- Press Order Entry.

- Press Ordinary Sample.

- Place cursor in Sample No. and input sample ID if the sample does not contain a barcode.

- Select assays to be analyzed.

- Press the down arrow to order the next sample.

- Press OK.

- Press Start.

- Confirm sample order status on the Joblist.

- Micro mode

- Follow Manual Order Processing steps on page 9 to ID sample and order test.

- Press in Mc column on Order screen.

- Load uncapped tube on to system.

- Select Start.

- Change to longer measurement time

- Follow Manual Order Processing steps on page 9 to ID sample and order test.

- Press Detailed Settings button.

- Click box below Measurement Time.

- Select measurement time

- Select OK.

- Select Start.

- Joblist information: Status

| Status | Meaning |
| --- | --- |
| Pending | Orders are registered but analysis is not yet started |
| Processing | Assay groups are being analyzed |
| On Hold | All assay groups have been analyzed but a validated calibration curve is missing |
| (Blank) | All assays are normal |
| Review | Analysis completed but needs a review of the results |
| Error | Hardware error has occurred. |

- Joblist information: Symbols

Result flags

Analysis Flags

- Results  management

The initial PTT result with Flag if run in normal mode – when sample run normal mode has flags like Analysis Time over, Slight Coagulation, No Coagulation or Early Reaction Error: Slow Reaction will auto-reflex to sub wavelength 600 seconds allowing that samples extra time for clot formation. It will create a second line on the Joblist. The final result will populate in the third line with all the tests results.

Sample run in micro-mode - re-analysis rules will not happen if sample run in micro mode and the longer Measurement Time (Sub wavelength) will need to be requested manually.

And also, Interference comment (sample hemolyzed/lipemic) won’t populate when run in micro mode. Enter comment before reporting result if any interference present (Hemolysis or lipemia).

Lipemic plasmas may have to be microfuge before running.

.

Refer to Sysmex 5100 Operational procedure for auto verification guidelines, normal and flagged results.

- REPORTING RESULTS:

- aPTT is not a calibrated assay. Results are reported in seconds.

These results should be related to the reference interval for the aPTT. Hemolyzed, lipemic, samples should be noted with the result. Results may be auto- or manually verified (refer to a separate Auto verification Guidelines section in the Operational procedure).

- >120 PTT result (patient not on heparin), “no clot” and a short PTT of <20.0 sec, will not auto verify and should be checked for a clot before being posted.

- REFERENCE RANGE:

Reference interval values are verified for each new lot number of Actin FSL (+/- 2.0 SD) of the mean put in use and varied with the lot#. Manufacturer recommended reference range is 23.9 – 29.9 seconds. If the verified range mean falls in the manufacturer recommended range, lab will accept the manufacturer range. If not follow the procedure to establish the reference range. The current ranges are found in the reagent rollover data binder on supervisor’s desk and are updated in BEAKER by IT personnel prior to putting the new reagents in use.

Reference range for the PTT assay = 23.9 – 29.9 seconds.

Procedure for validating new lot of PT reagent for reference range and lot to lot comparisons refer to “New Reagent Lot Number --Correlation with Old Lot Number and Validation of Reference Range”.

Critical results

APTT values greater than 120 seconds and PTT “No Clot” must be phoned to the patient’s caregiver as a critical value, and the call documented according to established protocol.

Regarding Reporting PTT “No Clot” result in BEAKER

  - There are two rows for PTT result, PTT and PTT – No Clot.

  - Double click on the patient in the outstanding list

  - Go To edit mode,

  - Type “NO CLOT” result in the PTT – No Clot row.

  - Verify the result. (You will see the comment “No clot detected- this is a critical value, suggest retest.” after verifying the result.

  - Go to Follow up Worklist for critical call, Call the result to the care giver.

- PROCEDURE NOTES

Overall performance of aPTT testing is dependent on reagent and instrument performance.

Acceptable variability (imprecision) should be such that the total day-to-day coefficient of variation (CV) of the analytic system is less than 5% on the same lot of control plasma.

- Analytical Measuring Range:  20 - 139 seconds

- Dilutions: N/A

- Sensitivity: APTT testing encompasses the entire clotting process from contact activation to fibrin formation and is therefore more susceptible to variations than specific individual tests.

ACTIN FSL may provide variable APTT results in samples containing the lupus-like anticoagulant.

Action of heparin as an anticoagulant is related to its ability in conjunction with a plasma cofactor to interfere with several aspects of the coagulation mechanism, thus retarding the rate of fibrin formation

Low levels of antithrombin III can decrease the responsiveness of patients to heparin therapy.

The APTT test will not detect qualitative or quantitative platelet disorders, isolated factor VII and factor XIII deficiencies or vascular disorders. APTT values for patient samples containing non-specific lupus-like anticoagulants are prolonged using ACTIN FSL

- Critical results:  ptt>120 SECONDS AND ptt = NO CLOT

- interferences:

- LIMITATIONS OF THE PROCEDURE:

APTT testing encompasses the entire clotting process from contact activation to fibrin formation and is therefore more susceptible to variations than specific individual tests. The control and use of APTT is therefore subject to inherent limitations. Control of plasma sample conditions is strictly emphasized. Studies have shown that sample decomposition may occur more rapidly in stored samples that are not refrigerated. Extremely small plasma volumes (prior to testing) are to be avoided since pH changes in the plasma from physiological conditions may be encountered. Such changes may lead to the decomposition of plasma components of the blood coagulation system. It should be noted that APTT testing may be affected by a number of commonly administered drugs. Decrease in time of APTT determination in conjugated estrogen therapy in males and oral contraceptive administration in females has been reported. Increase in the APTT has been seen in diphenylhydantoin, heparin, warfarin, naloxone and radiographic agent administration. Therapeutic doses of hirudin or other direct thrombin inhibitors may prolong clotting times Lipoglycopeptide antibacterial drugs (such as oritavancin or telavancin) may interfere with APTT based assays. Consult Instructions for Use of respective drugs.

In addition, the choice of anticoagulant (i.e. citrate vs. oxalate) and the condition of the

specimen (e.g. hemolyzed, lipemic, parenteral feeding, etc.) may affect results6,15,16.

The latter is particularly true of optical instrumentation measurements of the APTT. Blood

clotting factor deficiencies which should produce prolonged clotting times may be

compensated for or made to appear normal by elevated levels of one or more different clotting factors. Similarly, the presence of active intermediates which would tend to reduce the clotting time may also mask conditions that would normally lead to prolongation of the APTT. Mild or minor deficiencies in several factors may have an additive effect on increasing the APTT.

ACTIN FSL may provide variable APTT results in samples containing the lupus-like anticoagulant.

Unexpected abnormal APTT results should always be followed by additional coagulation

studies to determine the source of abnormal results.

Action of heparin as an anticoagulant is related to its ability in conjunction with a plasma

cofactor to interfere with several aspects of the coagulation mechanism, thus retarding the

rate of fibrin formation (see Section "Monitoring of Unfractionated Heparin Therapy with

APTT"). Low levels of antithrombin III can decrease the responsiveness of patients to heparin therapy. Antithrombin III slowly forms an inactive complex with the coagulant serine proteases, thrombin and factors IXa, Xa, XIa and XIIa. Heparin greatly accelerates this inactivation and this forms the basis for the therapeutic effect of these naturally occurring polysaccharide compounds. The test will not detect qualitative or quantitative platelet disorders, isolated factor VII and factor XIII deficiencies or vascular disorders. APTT values for patient samples containing non-specific lupus-like anticoagulants are prolonged using ACTIN FSL .

Siemens has validated use of these reagents on various analyzers to optimize product

performance and meet product specifications. User defined modifications are not supported by Siemens as they may affect performance of the system and assay results. It is the responsibility of the user to validate modifications to these instructions or use of the reagents on analyzers other than those included in Siemens Application Sheets or these Instructions for Use.

Results of this test should always be interpreted in conjunction with the patient’s medical

history, clinical presentation and other findings

Specific Performance Characteristics

ACTIN FSL has been carefully prepared to perform according to the results and within the limits described when used in the determination of activated partial thromboplastin time and in other coagulation procedures requiring an activated partial thromboplastin reagent.

Precision studies using the methodologies listed in this insert show that properly performed

APTT tests should result in a standard deviation (SD) which corresponds to a coefficient of

variation (CV) of less than 3 % in the normal range. In additional clinical and in-house studies, duplicate determinations on abnormal, normal plasma and control plasma were performed.

Results indicate that the APTT should agree within 3 % when performed properly.

- REFERENCES:

    - Sysmex CS 5100 Operator manual

    - Sysmex CS 5100 Instruction for Use.

    - Sysmex CS 5100 Reference Guide

    - Sysmex CS 5100 System Evaluation and Algorithm

4.	National Committee for Clinical Laboratory Standards. Collection, transport and processing of blood specimens for coagulation testing and performance of coagulation assays. Approved Guideline, 2nd Edition. NCCLS Publication H21-A2. Wayne, PA, 1991.

XV.	HISTORY:

H-1 	This procedure was written by P Bahel on 6/30/2020

Appendix-A

| QC Corrective Action Log Test |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clinical Hematology--PS5 Instrument: |  |  |  |  |  |
| Date | Specimen ID (Label) | Test result with previous QC | test Result after accepting QC | Acceptability Criteria (See Note below) | Tech initial/ Supervisor initial |
| Note: Acceptability Criteria |  |  |  |  |  |
| Coags /Fluid/ ESR: %age difference (old result-new result)/Old result*100; Acceptable within +/- 10% |  |  |  |  |  |