|  |  |  | TITLE: CS 5100 – QUANTITATIVE FIBRINOGEN (DADE THROMBIN) |  |  |  | DEPT OF LAB MEDICINE CLINICAL HEMATOLOGY Policy and Procedure Manual |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | DOCUMENT # HEM 218 - 1 - |
|  | WRITTEN BY: Parveen Bahel H(ASCP) |  | EFFECTIVE DATE: 07/14/2020 | REVIEW/REVISION: H-1 (New) |  |  | Page 1 of 12 |

- INTENDED USE:

For use in the quantitative determination of fibrinogen in plasma and to accelerate coagulation of anticoagulated samples for immunohematology studies.

- PRINCIPLE:

The enzyme thrombin converts the soluble plasma protein fibrinogen into its insoluble polymer, fibrin. The clotting time for diluted plasma is inversely proportional to the fibrinogen concentration of the plasma. By using this principle, Clauss developed a simple procedure for determining fibrinogen based on measuring the clotting time of diluted plasma after the addition of thrombin. The clotting time obtained in this manner is then compared with that of a standardized fibrinogen preparation.

- SUMMARY AND EXPLANATION

Depressed levels of fibrinogen are found in acquired and congenital hypo- and afibrinogenemias. Acquired fibrinogen deficient conditions occur especially as a result of intravascular proteolysis of fibrinogen by thrombin (consumption coagulopathy e.g. in obstetrics, following surgical intervention), snake venom or plasmin (primary hyperfibrinolysis following therapy with streptokinase, urokinase or tPA). Furthermore, moderate hypofibrinogenemias may occur in cases of diminished production (in acute or chronic liver diseases), loss into the intravascular space (e.g. in ascites or acute hemorrhage and burns) or increased degradation (in shock or carcinoma).Temporarily elevated levels of fibrinogen are found as a result of the behavior of fibrinogen as an ”acute-phase protein”

a. Transitory hyperfibrinogenemias may occur after operations, traumas, myocardial infarction and infections.

b. Persistent hyperfibrinogenemias may be found in patients with neoplasias and chronic inflammatory diseases.

Levels of fibrinogen are found to increase slightly with age.

Elevated fibrinogen levels are a risk factor for cardiovascular disease.

Blood samples collected from patients who have received heparin may not clot and units of plasma may need to be converted to serum prior to laboratory use. Thrombin can be used to accelerate clotting of anticoagulated samples and units of plasma for use in immunohematologic tests.

- SPECIMEN REQUIREMENTS

The only acceptable sample for routine coagulation is plasma, from whole blood drawn into blue-stoppered tubes containing 3.2% sodium citrate (nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant).

- 4.5 ml-draw tube must have a MINIMUM ACCEPTABLE VOLUME of 4.0 ml total (3.5 ml blood +0.5 ml anticoagulant).

- 2.7 ml-draw tube must have a MINIMUM of 2.5 ml (2.2 ml blood + 0.3 ml anticoagulant).

- A “micro method” sample (newborn PT/PTT/FIB only) must be a full 1.0 ml specimen (0.9 ml of whole blood drawn with a syringe added to 0.1 ml 3.2% sodium citrate anticoagulant.)

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing of blood and anticoagulant.

A syringe or evacuated tubes (blue top) may be used for collection. If multiple specimens are collected; the coagulation sample should be the second or third tube collected. If only coagulation testing is to be performed, a red-top tube, which has no additives, should be drawn first and discarded prior to drawing the blue-top coagulation tube.

The patient cannot be on anti-coagulants when the test specimen is collected. Sufficient time after discontinuance of heparin should be allowed for heparin to be cleared from the patient’s blood, usually 6 hours.

If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values above 55%.

Specimens that are clotted, collected in the wrong tube or serum, or have less than the 90% expected fill should be rejected. The floor must be notified, a comment entered into the computer.

Handling Condition and Stability

The whole blood specimen is checked for clot formation by gentle inversion and observation. Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x10 9/L for 10 minutes at 4000 g or 3 minutes at 8000 rpm (STAT centrifuge) for STAT samples and 7 minutes at 5200 RPM (DAS centrifuges) as soon as possible after collection. If immediate testing is to be done, the plasma may remain on the packed cells or separated.

Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells. For special coagulation testing, spin samples 20 minutes at 4000 g, separate plasma into plastic tubes, label and freeze all aliquots at –70C located in the Special coag area until ready to use. Always check for clots after aliquoting with applicator sticks. Always tracks aliquots in BEAKER under YH Coag Hold before freezing them. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested.

Specimen stability at ambient temperature: 4 hours; Frozen at -70° C 6 months.

Specimen Labelling Specimen should be properly labeled with at least 2 unique patient identifiers. Patient's full name and medical record number (MRN) and should have the date and time of collection. The patient's birth date may substitute for the MRN if the MRN is not available

- ENVIRONMENTAL OPERATING CONDITIONS:

The room temperature and humidity percent are monitored and documented on the Routine Coagulation checklist.

- EQUIPMENT AND MATERIALS:

A.	Equipment:

    - Nerl Water: pH 7.0

    - Gauze

    - Citrated blue top tubes

    - Frosted tubes for aliquots

    - Cuvettes

    - SLD Mini - cups

    - CS5100 sample cups – 4 mL

    - CA Clean I – 50 mL (On-board stability – 5 days)

    - CA Clean II – 5 L (On-board stability – 2 months)

    - OVB (Buffer) (On-board stability – 24 hours)

    - Trash bags

B.	Materials:

- Standard Human Plasma

- Dade Thrombin reagent

- Citrol 1

- Dade Abn Fibrinogen control

C.	Preparation:

Dade® Thrombin Reagent: Lyophilized bovine thrombin preparation (approx. 100 IU/mL) with stabilizers and buffers.

      - Reconstitute the Thrombin reagent with 1.0 mL NERL  water (according to label on vial). Close the vial and let stand until the contents have dissolved. Wait for 15-20 minutes. Carefully swirl the contents to mix. Do not shake. Mix carefully once more before using.

      - Stability after reconstitution:  72 hours on board (10 C) and 5 days @ 2-8 C.

      - Reconstitute lyophilized Standard Human Plasma with 1.0 mL     distilled water (As needed for calibration).

      - Reconstitute lyophilized Citrol Level 1 and Abnormal Fibrinogen control with 1.0 mL Nerl water.

        - Re-stopper vial and mix carefully to dissolve (without foam         formation).

        - Let stand at 15-25C for at least 15 minutes before use.

Note:

To perform QC or calibration, always transfer controls or calibrator into SLD Mini Cups.

D.	Stability after reconstitution:

Calibrators: As calibrators are intended for immediate use, no on-board stability data has been established.

Thrombin Reagent: Stability after reconstitution: 72 hours on board, 5 days at   2 - 8C.

Citrol 1 and Abn Thrombin Control: -- 24 hours on board.

- LOADING REAGENTS and QC

- Reagent table loading:

Note:  Always load reagent in the orange racks (A racks) and the control and calibrators in the yellow racks (C racks).

Reagents cannot be loaded while running the patients.

- Press Reagent on the toolbar.

- Verify system is in solid green ready status.

- Open light shield lid.

- Open reagent table cover.

- Remove expired or empty reagent vials by pressing Change/Add icon. It will bring the rack in front you want to remove reagents.

- Place new reagents in a rack with barcode facing out.

- Load the reagent rack.

- Press and lock reagent table cover.

- Press OK on screen to barcode read.

- Wait for barcode reading to read the barcode.

- On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

B.	Loading the Buffer

1. Press Reagent on the toolbar.
2. Verify buffer table cover LED is green.
3. Open buffer table cover.
4. Place OBB (Buffer) in an adaptor with barcode facing out.

5. Place into the buffer table and close the cover.
6. Press OK on the barcode reading screen.
7. On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

C.	Loading QC

C-rack and SLD mini cup

- Reconstitute Citrol 1 and  Abn Ctl P

- Aspirate entire content of QC vial (GW5) using transferring pipette.

- Slowly dispense the entire volume into a new SLD mini cup, avoiding any air bubble formation.

- After reagent transfer, carefully check for bubbles and, if required, remove using smaller pipette tip.

- Set SLD mini cup into corresponding GW5 vial.

- Set GW5 with SLD into yellow-handled C-rack. Push down on vial to ensure proper seating in position.

- Place the rack back in reagent table and close lid.

- Lock lever and Press OK to read barcode.

- On Reagent screen, highlight vial just loaded and press Change to update date/time.

Note the SLD mini cups are for controls and calibrators only.

- How to order QC:

Processing from the reagent table: Control lot

2. Select Order.

3. Select Switch Order.

4. Select Holder QC Order

5. Press Order Entry.

6. Select control material from the list.

7. Place cursor in the Lot No. field and select lot from list.

8. Select the test ‘Fib’ from the key pad.

9. Press the down arrow to order second level of QC-Abn ctl

10. Press OK once controls have been ordered.

11. Press Start.

- How to review qc and add a comment or initial to a point:

1. Press QC Chart.
2. Highlight Fib QC Chart.

3. Place cursor on point.
4. Press Confirm Data key.
5. Enter your initial or comment in field if there is any QC issue. Also log in the QC log in L: drive once done.
6. To exclude point: check box.

7. Press Update.

1.	Controls should be tested once each 8-hour shift, and after reagent changes / additions. Tech has to review shift control and placed an initial in QC Software under each control.

2.	Controls should be run in the same manner as the test samples and by the same technologists who work in the routine coagulation area.

3.	Control tolerance limits--the range is calculated based on +/- 2 SD

(Standard deviations) from the mean control value.

4.	Corrective action when tolerance limits are exceeded (results will be flagged with error message):

a.	Swirl reagent and QC well and rerun out-of-range control

b.	If still out, Check reagent/QC for bubbles and turbidity.

c.	Check reagents/ change if expired, ensure fresh OVB buffer is on-board and change if necessary (if not changed during set up). Prime the instrument.

d.	Make fresh reagent or QC if needed and rerun again

e.	Document actions taken to identify in QC Software and correct the problem before reporting any patient data.

NOTE: If QC is out of range after running twice, corrective action and an evaluation need to be performed. Up to 5 patient samples when QC was acceptable will need to be rerun to determine if there is a significant clinical difference in the patient results. Documentation is to print results from the Instrument of the repeated run of the patient samples, fill out corrective action log; print Beaker sample label and file in Corrective Action notebook. Acceptable agreement should be +/- 10% of the original results (See Appendix A for log). If significant clinical changes have been determined, there will be an amended report with corrected data and a caregiver will be notified promptly. In these circumstances, the analyzer will be put out of service and troubleshooting/service dispatched. Continue to run samples on the other functional instrument, if needed set up CS 5100-3 instrument for routine coag tests.

- If this occurs during off-shifts and the immediate concern is to determine if patient data has been compromised, the patient samples from the problem instrument should be rerun on the other functioning instrument and results compared. See corrective action log for acceptable criteria. Troubleshooting will then be performed timely on the problem instrument, QC run and evaluated and patient comparisons performed before instrument is put back into service.

- Calibration

- Identify the reagents

- Enter reagent and calibrator lot information in Reagent Lot Master. Refer to the Enter Reagent Information tab for more details.

- Load reagents, calibrator, and buffer.

- Order the calibration

- Select Order.

- Select Switch Order.

- Select Holder Calib Curve Order.

- Select the Fib test to be calibrated.

- Select Change and select the correct lot number.

- Select OK.

- Select the correct calibrator lot number from list. Assay value automatically displays.

- For manual entry of assay value: Place the cursor in the Assay Sheet Value field and enter value using keypad.

- Select OK.

- Select Start.

- To view the calibration status and progress, press Joblist.

Note: Calibrator plasma should be transferred into a SLD mini cup and placed in the vial. Place vial in a C-rack. Refer to Loading Reagents tab for more details.

- Calibration Validation

- View new calibration curve

- Select Calib. Curve.

- Select Change

- Select assay key.

- Select the lot number.

- Compare new/current calibration curve

- Select Detailed Display

- Select Select Compared Calib. Curve

- Select a calibration curve to compare and press Load.

- Compare calibration curves

- Select Close.

- Validate new calibration curve

- Select Validate to validate the calibration curve.

- Select OK.

- Select Print.

Notes:

- Validate calibration by performing QC.

- If Calibration fails, then recalibrate instrument following steps from the Calibration Procedure Sections. Refer to Calibration or recalibration frequency is based on Policy HEM 179 (Calibration and Analytical measurement Policy)

- If repeat Calibration fails again, Call Siemens Technical Service Assistance, If Instrument is unacceptable for operation, use back-up analyzer.

- Loading patient samples on the cs 5100

- LIS order processing

- Load bar-coded samples into any sample rack (7-20) with the bar-code label visible through the window of the rack position and Place rack on the sampler.

- Check host connection (HC) status. HC status icon must be green or orange.

- Press Start.

- After barcode reading, confirm sample order status and progress on the Joblist screen.

- Manual order processing

- Place rack with sample tubes on sampler.

- Press Order.

- Enter rack number

- Select tube position to input an order.

- Press Order Entry.

- Press Ordinary Sample.

- Place cursor in Sample No. and input sample ID if the sample does not contain a barcode.

- Select Fib test to be analyzed.

- Press the down arrow to order the next sample.

- Press OK.

- Press Start.

- Confirm sample order status on the Joblist.

- Micro mode

- For barcoded samples,

- Press in Mc column on Order screen.

- Load uncapped tube on to system.

- Select Start.

- Change to different dilutions if sample run in micro mode

- Press Detailed Settings button.

- Click box below Dilution value

- Select dilutions based on the Fib result. Select 4/1dilution if Fib is <80 and select 1/2if Fib is >450 mg/dL.

- Select OK.

- Select Start.

Note: if a sample is ordered in micro analysis mode, re-analysis rules will not be applied, and the high or low dilutions will need to be requested manually.

- Joblist information: Status

| Status | Meaning |
| --- | --- |
| Pending | Orders are registered but analysis is not yet started |
| Processing | Assay groups are being analyzed |
| On Hold | All assay groups have been analyzed but a validated calibration curve is missing |
| (Blank) | All assays are normal |
| Review | Analysis completed but needs a review of the results |
| Error | Hardware error has occurred. |

- REPORTING RESULTS:

Patient values auto verify in mg/dL (or can be posted by the tech) in the computer system. For auto posting or manual reporting results, refer to Sysmex 5100 Operational procedure for auto verification guidelines, reporting normal and flagged results.

Hemolyzed, or lipemic samples must be commented with the result.

- REFERNECE RANGE:

Reference range of fibrinogen assay: 194 - 448 mg/dL

A Reference interval value verified for each new lot number of Dade Thrombin reagent (+/- 2.0 SD) of the mean put in use and varied with the lot#. The current ranges are found in the reagent rollover data binder on supervisor’s desk and are updated in BEAKER by IT personnel prior to putting the new reagents in use.

Reference interval values are verified for new instrumentation and new lot number of dade Thrombin (+/- 2.0 SD) reagent of the mean put in use and varied with the lot#. Manufacturer recommended reference range is 194-448 mg/dL if the verified range mean falls in the manufacturer recommended range, lab will accept the manufacturer range. If not, follow the procedure to establish the reference range. The current ranges are found in the reagent rollover data binder on supervisor’s desk and are updated in BEAKER by IT personnel prior to putting the new reagents in use.

- Clinical Reportable Range (CRR): 50 mg/dL to 860 mg/dL

-Results below 80 automatically run at 4/1 dil and result above 450 run at ½ dilution if sample run in normal mode. For no coagulation flag, report “No Clot” result. Refer to CS 5100 Operational procedure for Reporting results.

- CRITICAL RESULTS:

Fibrinogen result < 100 (each occurrence) must be phoned to the patient’s nursing unit or physician as a critical value. This must be documented in the computer according to established protocol. For calling critical results and documentation refer to Critical Call Policy Hematology Laboratory, Doc# HEM 148.

- PROCEDURE NOTES:

The overall performance of Fibrinogen testing is dependent on the instrument 			performance as well as the reagent used. Acceptable variability (imprecision) 			should be such that the total day-to-day coefficient of variation (CV) of the 			analytic system is less than =7.5% on the same lot of control plasma.

A new reference curve should be established every 6 months, or with each change of reagent lot or instrument, or with any consistent deviation from control or proficiency testing limits.

- INTERFERENCES

Others

The results obtained may be influenced by the presence of heparin or fibrino(geno)lytic degradation products in patient plasma. Significant amounts of each of these substances may lead to a false low value for fibrinogen in the test.

Results of this test should always be interpreted in conjunction with the patient's medical history, clinical presentation and other findings

- LIMITATIONS OF THE PROCEDURE:

Levels of the following appear to interfere with Dade Thrombin Reagent on the

Up to

Bilirubin 6 mg/dL

Hemoglobin (free) 100 mg/dL

Triglycerides 284 mg/dL

Heparin (LMW) 0.4 U/mL

Heparin (unfractionated) 0.6 U/mL

The results obtained may be influenced by the presence of heparin or fibrino(geno)lytic

degradation products in patient plasma.

Lipemic samples may need to be ultra centrifuged in Chemistry lab before analysis.

- REFERENCES:

  - Sysmex CS2500 Reference Guide

  - Sysmex CS 2500 System Evaluation and Algorithm

  - Sysmex CS2500 Instruction for Use.

- HISTORY:

H-1	This procedure was written by P Bahel on 6/30/2020

| Appendix-A Clinical Hematology--PS5 Instrument: |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Date | Specimen ID (Label) | Test result with previous QC | test Result after accepting QC | Acceptability Criteria (See Note below) | Tech initial/ Supervisor initial |
| Note: Acceptability Criteria for |  |  |  |  |  |
| Coags /Fluid/ ESR: %age difference (old result-new result)/Old result*100; Acceptable within +/- 10% |  |  |  |  |  |
| Urine: +/- 1 grade of the original results for macroscopic and specific gravity within +/- 0.003, +/-10% or within 1 range for Microscopic |  |  |  |  |  |