|  | TITLE: THROMBIN TIME (Test Thrombin Reagent-CS5100) |  | DEPT OF LAB MEDICINE CLINICAL HEMATOLOGY Policy and Procedure Manual |
| --- | --- | --- | --- |
|  |  |  | DOCUMENT # HEM 222 - 1 - |
| WRITTEN BY: Parveen Bahel H(ASCP) | EFFECTIVE DATE: 7/14/2020 | REVIEW/REVISION H-1 (New) | Pages 1 of 8 |

I.	PRINCIPLE:

Thrombin converts the fibrinogen in the plasma sample into fibrin, thus producing a 	fibrin clot. The time taken for the fibrin clot to form is measured as the thrombin time.

II.	INTENDED USE: Reagents for the quantitative determination of the thrombin time in citrated human plasma.

III. 	SUMMARY AND EXPLANATION: The determination of thrombin time Test Thrombin Reagent is suitable for

- monitoring fibrinolysis therapy.

- screening for disorders of fibrin formation and in case of suspected severe fibrin deficiency states.

- differentiating between a heparin-induced prolongation of the thrombin time and fibrin formation disorders.

The thrombin time is found to be prolonged not only due to disorders in fibrin polymerization but also due to the presence of heparin.

IV. 	SPECIMEN REQUIREMENTS

Type:	The only acceptable sample for routine coagulation is plasma, from whole blood drawn into blue-stoppered tubes containing 3.2% sodium citrate (nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant).

- 4.5 ml-draw tube must have a MINIMUM ACCEPTABLE VOLUME of 4.0 ml total (3.5 ml blood +0.5 ml anticoagulant).

- 2.7 ml-draw tube must have a MINIMUM of 2.5 ml (2.2 ml blood + 0.3 ml anticoagulant).

- A “micro method” sample (newborn PT/PTT/FIB only) must be a full 1.0 ml specimen (0.9 ml of whole blood drawn with a syringe added to 0.1 ml 3.2% sodium citrate anticoagulant.)

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing of blood and anticoagulant.

A syringe or evacuated tubes (blue top) may be used for collection. If multiple specimens are collected; the coagulation sample should be the second or third tube collected. If only coagulation testing is to be performed, a red-top tube, which has no additives, should be drawn first and discarded prior to drawing the blue-top coagulation tube.

The patient cannot be on anti-coagulants when the test specimen is collected. Sufficient time after discontinuance of heparin should be allowed for heparin to be cleared from the patient’s blood, usually 6 hours.

If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values above 55%.

Specimens that are clotted, collected in the wrong tube or serum, or have less than the 90% expected fill should be rejected. The floor must be notified, a comment entered into the computer.

Handling Condition and Stability

The whole blood specimen is checked for clot formation by gentle inversion and observation. Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x10 9/L for 10 minutes at 4000 g or 3 minutes at 8000 rpm (STAT centrifuge) for STAT samples and 7 minutes at 5200 RPM (DAS centrifuges) as soon as possible after collection. If immediate testing is to be done, the plasma may remain on the packed cells or separated.

Patient plasma should be tested within 4 hours. If immediate testing is to be done, the plasma may remain on the packed cells. For special coagulation testing, spin samples 20 minutes at 4000 g, separate plasma into plastic tubes, label and freeze all aliquots at –70C located in the Special coag area until ready to use. Always check for clots after aliquoting with applicator sticks. Always tracks aliquots in BEAKER under YH Coag Hold before freezing them. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested.

Specimen stability at ambient temperature: 4 hours; Frozen at -70° C 6 months.

Specimen Labelling Specimen should be properly labeled with at least 2 unique patient identifiers. Patient's full name and medical record number (MRN) and should have the date and time of collection. The patient's birth date may substitute for the MRN if the MRN is not available.

V.	ENVIRONMENTAL OPERATING CONDITIONS:

The room temperature and humidity percent are monitored and documented on the Routine Coagulation checklist.

VI. 		EQUIPMENT AND MATERIALS:

A.	Equipment:

    - Nerl Water: pH 7.0

    - Gauze

    - Citrated blue top tubes

    - Frosted tubes for aliquots

    - Cuvettes

    - SLD Mini - cups

    - CS5100 sample cups – 4 mL

    - CA Clean I – 50 mL (On-board stability – 5 days)

    - CA Clean II – 5 L (On-board stability – 2 months)

    - OVB (Buffer) (On-board stability – 24 hours)

    - Trash bags

B.	Reagents:

Test Thrombin Reagent kit has 10x for 5 mL Thrombin reagent (1.5 IU/mL of bovine thrombin, bovine albumin),1 x 50 mL Buffer Solution (HEPES 25 mmol/l) pH 7.4

Control N (normal ctl)

Control P diluted 1:3 with OVB Buffer

VII.	REAGENT PREPARATION:

- Reconstitute the contents of a vial of Test Thrombin reagent with 5 mL of buffer solution. Mix gently before using. Prior to the assay warm the reagent solution to 37 °C in the original vial or in a plastic tube. Mix carefully once more before using.

DO NOT SPILL BUFFER—THERE ARE ONLY 50 ML OF BUFFER PROVIDED.

B.	After reconstitution, reagent is stable on board the CS 5100 for 72 hours (HOWEVER, after using reagent, re-cap the vial and store in refrigerator 	 (2- 8C), where it is stable for 7 days. Mark the expiration date on the vial.

C. 	Reconstitute the normal control N with 1 mL distilled water. Let it sit for 15 minutes. Load into Yellow rack. Log result in the Thrombin time QC notebook.

D.  		Reconstitute a Control P and let sit 15 minutes at room temperature. Swirl gently to mix; prepare 1:3 dilutions with buffer (0.4 mL OVB buffer + 0.2 mL Control P). Run abnormal ctl as a patient sample. Ensure that the thrombin time result falls within the acceptable range, repeating if necessary.

NOTE: After opening any vial to place onto the instrument, label that vial with the open and expiration date, referring to the stability information provided here.  Discard reagent when expired.

- LOADING REAGENTS and QC

- Reagent table loading:

Note:  Always load reagent in the orange racks (A racks) and the control and calibrators in the yellow racks (C racks).

Reagents cannot be loaded while running the patients.

- Press Reagent on the toolbar.

- Verify system is in solid green ready status.

- Open light shield lid.

- Open reagent table cover.

- Remove expired or empty reagent vials by pressing Change/Add icon. It will bring the rack in front you want to remove reagents.

- Place new reagents in a rack with barcode facing out.

- Load the reagent rack.

- Press and lock reagent table cover.

- Press OK on screen to barcode read.

- Wait for barcode reading to read the barcode.

- On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

B.	Loading the Buffer

1. Press Reagent on the toolbar.
2. Verify buffer table cover LED is green.
3. Open buffer table cover.
4. Place OVB in an adaptor with barcode facing out.

5. Place into the buffer table and close the cover.
6. Press OK on the barcode reading screen.
7. On Reagent screen, touch reagent position just loaded, and press Change to update the Set Time.

C.	Loading QC

C-rack and SLD mini cup

- Reconstitute QC for both levels.

- Aspirate entire content of QC vial (GW5) using transferring pipette.

- Slowly dispense the entire volume into a new SLD mini cup, avoiding any air bubble formation.

- After reagent transfer, carefully check for bubbles and, if required, remove using smaller pipette tip.

- Set SLD mini cup into corresponding GW5 vial.

- Set GW5 with SLD into yellow-handled C-rack. Push down on vial to ensure proper seating in position.

- Place the rack back in reagent table and close lid.

- Lock lever and Press OK to read barcode.

- On Reagent screen, highlight vial just loaded and press Change to update date/time.

- Note the SLD mini cups are for controls and calibrators only.

- How to order QC:

Control N is processed from the reagent table and abnormal control is processed from the sample rack

A. Processing from the reagent table: QC files

1. Load QC onto a C-rack.

Refer to the Loading Reagents tab to load QC

material in a C-rack using a SLD mini cup.

2. Select Order.

3. Select Switch Order.

4. Select Holder QC Orde.

5. Press Order Entry.

6. Select control N from the list.

7. Place cursor in the Lot No. field and select lot from list.

8. Select the Thrombin time – control N for TT

10. Press OK once controls have been ordered.

B. Processing from the sample rack

1. Place QC in cup and place in sample rack.

2. Press Order.

3. Enter rack number.

4. Select tube position to input the order

5. Select Order Entry.

6. Select QC Sample radio button.

7. Select control (Abn ctl P) from the list.

8. Place cursor in the Lot No. box and select the lot number from the list.

9. Select the Thrombin Time assay.

11. Press OK once control has been ordered.

12. For Micro Mode, check MC column of each control.

13. Press Start

- How to review qc and add a comment or initial to a point:

1. Press QC Chart.
2. Highlight TT QC Chart.

3. Place cursor on QC point.
4. Press Confirm Data key.
5. Enter your initial or comment in field if there is any qc issue and also enter initial in QC lig saved in L:Drive.
6. To exclude point: check box.

7. Press Update.

QC Corrective action when tolerance limits are exceeded:

- Verify reagent performance.

- Check instrument performance.

- Document actions taken to identify and correct the problem before reporting any patient data.

- If problem cannot be resolved.

- Call for service if necessary and properly document in troubleshoot log

- Notify supervisor

X.	CALIBRATION:

No calibration of the system is necessary for performing Thrombin Time assay.

XI.	PROCEDURE:

A.	Prepare reagent, let sit for 10 minutes, then place in an orange rack.

B.	Prepare control N and place on instrument in yellow rack and P run in a sample rack Run controls and assure 	that results are within acceptable ranges.

C.	Place patient samples on instrument and Prior to the assay warm the reagent solution to 37 °C in the original vial or in a plastic tube. Mix carefully once more before using. Order thrombin time test.

D.	Record the thrombin time and controls in the QC notebook; patient value will usually autopost, but can be posted manually.

XII.	NOTES:

A.	If the PTT is prolonged, repeat the PTT with added polybrene (heparin antagonist) as described in a separate procedure. If standard (unfractionated) heparin is present in the sample, the PTT should shorten to normal or near-normal with polybrene.  Cancel the thrombin time and include a comment against the thrombin time that the test is not performed in the presence of heparin. Abnormal thrombin time result auto reflexes to PTT if not ordered by physician.

Refer to Sysmex 5100 Operational procedure for auto verification guidelines, normal and flagged results

XIII.      REFERENCE RANGE:

Procedure for validating new lot of Thrombin Time for reference range and lot to lot comparisons refer to “New Reagent Lot Number --Correlation with Old Lot Number and Validation of Reference Range”.

Reference interval value determined for Thrombin Time is 16.7 – 18.3 seconds (Effective from 7/14/2020)

Analytical Measuring Range: 10.0 – 107 seconds.

All Results reflex to MD Interpretations.

XIV.	CRITICAL RESULTS: No critical result for the procedure.

XV. 	LIMITATIONS OF THE PROCEDURE: Results of this test should always be interpreted in conjunction with the patient’s medical history, clinical presentation and other findings. Carry-over of traces of thrombin may interfere with subsequent coagulation analyses. Therefore, only single-use plastic materials should be used.

XVI. 	INTERFERENCES

XVII.		REFERENCES:

  - Sysmex CS5100 Reference Guide

  - Sysmex CS 5100System Evaluation and Algorithm

  - Sysmex CS5100 Instruction for Use.

  - Siemens CS 5100 Operator’s Manual

XVIII. HISTORY H-1This procedure was written by P Bahel on 6/25/2020.