| Ridgeview Medical Center Laboratory Services |  |  | Blood Bank |  | Effective Date: 6/10 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose and or Principle When more than one instrument/method is used to test for a given analyte, the instruments/methods must be checked against each other at least twice each year for comparison of test results. Antibody screen- IgG gel card versus tube, immediate spin crossmatch – buffered gel card versus tube, group/Rh typing – ABD/ABD gel card, ABD/Reverse gel card versus tube, C and E and c and e antigen typing – gel antigen typing cards versus tube, K and Fya antigen typing – IgG gel card versus tube using typing antisera, and direct antiglobulin testing – IgG gel card versus tube will be performed twice a year. All results must agree for the comparison testing results to be considered acceptable. |  |  |  |  |  |  |  |
| Specimen Specimens for correlation testing shall include patient specimens, quality control material, and proficiency testing specimens. Refer to specific SOPs for specimen requirements. |  |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |
| Reagents |  | Supplies |  | Equipment |  |  |  |
| · See appropriate procedure |  | ● See appropriate procedure |  | ● See appropriate procedure |  |  |  |
| Quality Control Prior to routine testing, reagents must be attested with appropriate controls each day for use. Refer to Quality Control Procedure. Controls must exhibit expected results for testing to be valid before reporting any patient results. |  |  |  |  |  |  |  |
| Procedure Notes 1. Do not remove the foil seal until ready to use or within one hour of testing. 2. Inspect each gel card for a clear liquid layer on top of the opaque gel. 3. Do not touch the gel card during pipetting to prevent carryover leading to erroneous results. |  |  |  |  |  |  |  |
| Step | Action |  |  |  |  |  |  |
| 1 Group/Rh Typing | Gel ABD/ABD card, Gel ABD/Reverse card and tube typing with tube antisera Unit confirmation/patient recheck/Cord – ABD/ABD card versus tube Perform unit confirmation testing on A Pos, B Pos, AB Pos, O Pos, A Neg, B Neg, AB Neg and O Neg unit Perform patient recheck on as many patient specimens of different types as available Perform Cord typing on at least two specimens, one Rh negative and one Rh positive using both the clot and EDTA tubes Patient typing – ABD/Reverse card versus tube Perform patient typing on as many patient specimens of different types as available Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 2 K and Fya Typing | Gel IgG card with tube antisera and tube typing with tube antisera Unit/patient antigen typing – IgG gel card versus tube Perform K and Fya antigen typing on at least 4 unit segments and 4 patient specimens Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 3 Antibody Screens | Gel IgG card screens with LISS and Albumin enhancement tube screens Patient screens – IgG gel card versus LISS tube versus ALB tube methods Perform antibody screens on at least 5 specimens including positive and negative results Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 4 Immediate Spin Crossmatch | Buffered gel card crossmatches and tube immediate spin crossmatches Patient crossmatches–Buffered gel card versus tube immediate spin xmatch Perform immediate spin crossmatches on at least 5 specimens Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 5 DAT Testing | Gel card and tube DATs including Anti IgG and Coombs Control Cell phases Patient DATs – IgG gel card versus tube Perform DATs on a least 4 patient specimens and 4 unit segments Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 6 C, E, c and e Antigen Type | Gel C, E, c and e cards and tube typing Unit/patient antigen typing – C, E, c and e card versus tube Perform C, E, c, and e antigen typing on at least 4 unit segments and 4 patient specimens Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |
| 7 Weak D Typing | Gel ABD/Reverse , ABD/ABD card and tube weak D typing Patient typing – ABD/Reverse card, ABD/ABD card, versus tube weak D Perform Weak Ds on at least 4 patient specimens and 1 Cord Document results on appropriate form including tech initials and date performed Perform review for acceptability – resolve and investigate any failures File forms in Comparison Testing notebook |  |  |  |  |  |  |

Interpretation To be considered acceptable performance the results obtained using the different methodologies and/or instruments should be comparable to each other. The ABO/Rh interpretation should agree exactly between the tube and gel method. The degree of strength must agree within one. The K and Fya antigen typing interpretation should agree exactly between the tube and gel method. The degree of strength must agree within one. The strength of antibody screen reactions in the comparison of the tube method to the manual gel method should agree within one degree of strength. The immediate spin crossmatch interpretation should agree exactly between the tube and gel method. The degree of strength must agree within one. The tube IgG DAT and the gel DAT should agree within one degree of strength. The C, E, c and e antigen typing interpretation should agree exactly between the tube and gel method. The degree of strength must agree within one. The Weak D interpretation should agree exactly between the tube and gel method. The degree of strength must agree within one.

Result Reporting Results are recorded on the appropriate forms All results will be reviewed by the Lead Tech or the Medical Director of the Transfusion Service. Any unacceptable correlation results will be investigated by the Transfusion Service Lead Tech.

Restrictions/Limitations Correlation testing should end in the same results for the corresponding methods of testing except for the Backup Method for antibody screen comparisons of gel with tube enhancement media. In the backup method correlation testing, the Gel IgG card is more sensitive than the backup methods of LISS and Albumin tube screening. For the Screen correlations, the interpretation results should agree as positive for strengths of 1+ and more in gel. For strengths less than 1+ in gel, the interpretation results aren’t required to agree with the tube results because of the weakness of the reactions. The results should agree within one degree.

Safety Refer to Laboratory Procedure # 4002-1: Infection Control and Laboratory Safety Hygiene Procedure and all safety policies of the laboratory.

Related Documents Specimen Requirements, Labeling and Storage Cell Suspensions and Conversions Interpreting Reactions ABO Type-Forward and Reverse, Rh Gel Method Rechecks and Confirmation Testing Indirect Antiglobulin Test – Gel Method Crossmatch – Electronic, Immediate Spin & IgG (Gel) Antigen Testing Guide/Antisera IDs C, c, E e Antigen Typing – Gel Fya and K Antigen Tying – Gel Cord Blood Workup Quality Control – Reagent QC – Gel Method ABO Forward and Reverse, Rh-Weak D – Tube Rh Phenotyping – C, c, E, e – Tube Direct Antiglobulin Testing – Tube Fya and K Typing – Tube Alternate Methods – LISS, Albumin

References CAP Transfusion Medicine Checklist, 6/17/2010 Edition, Northfield, IL

| Date: |  | Author: |  |  |  |  | Date: |  |  | Validated by: |  |  |  | Date: |  |  | Approved by: |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 6/2010 |  | Jo Lohmann, MT (ASCP) |  |  |  |  | 6/2010 |  | Jo Lohmann, MT (ASCP) |  |  |  |  | 6/20/2010 |  |  | R.P. Weatherby, MD |  |  |  |
| File name and location (L:Blood Bank: Method Correlation) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 8/2011 | JL |  |  | Addition of additional gel method correlations |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Discontinued |  |  |
| --- | --- | --- |
| Date: | Reason: | Initials: |