| Ridgeview Medical Center Laboratory Services |  | Blood Bank | Effective Date: 12/15/2010 |  |  |
| --- | --- | --- | --- | --- | --- |
| Purpose This procedure provides instructions for performing daily reagent and instrument quality control, individual reagent testing, and the performance of a Performance Measure Quality Assessment. |  |  |  |  |  |
| Reagents | Supplies |  |  | Equipment |  |
| MTSTM A/B/D Monoclonal & Reverse Grouping Card MTSTM A/B/D Monoclonal Grouping Card MTSTM Buffered Gel Card MTS TM IgG Gel Card --- all cards stored at 2-25°C--- MTSTM Diluent 2 PLUS, stored at 2-8°C MTSTM Diluent 2, stored at 2-8°C Ortho Screening Cells, stored at 2-8°C Ortho 0.8% Affirmagen 3, stored at 2-8°C Ortho Confidence System, stored at 2-8° C 22% Albumin, stored at 2-8° C Anti-A Antisera, stored at 2-8° C Saline | - Pipette tips - 12 x 75 tubes |  |  | - BioHit Pipettor - Eppendorf Pipettor - MTS Incubator - MTS centrifuge |  |
| Preparation of 4% QC Cells 1 and 2. WEEKLY (Wednesday). 1.; Label 2-12x75 tubes – QC1, QC2. 2.; Place 2.0 mls of QC Cell 1 and QC Cell 2 into the appropriate tubes by the following method: - Use the Eppendorf Pipettor (set to 1.0 ml)– dispense 1.0 ml twice. 3.; Centrifuge both tubes for 60 seconds and decant the supernatant completely.. 8. 4.; Add 1.5 mls MTSTM Diluent 2 PLUS to each tube by the following method: - Use the Eppendorf Pipettor (1.0 ml, then reset to 500 ul). 5.; Label 2 transport vials – Reagent Labeling requirements - QC 1 – Lot # - Date made – tech initials – expiration date - refrigerate - QC 2 – Lot # - Date made – tech initials – expiration date - refrigerate. 6.; Mix gently and transfer cell solutions to appropriate transport vials for use and storage.. DAILY (if necessary). 1.; Using BioHit Program ‘P’ ▲200 - Pipette 200 µL of QC Cells 1 & 2 into labeled tubes. - Centrifuge 60 seconds - Decant supernatant completely Using BioHit Program ‘P’▼ 150 - Pipette 150 µL of MTSTM Diluent 2 PLUS into each tube - Mix to resuspend. |
| DAILY Preparation of 6% Albumin Control. 1.; Using BioHit Program ‘P’ ▲100 - Pipette 100 µL of 22% Albumin into a labeled and dated 12 X 75 tube Using BioHit Program ‘P’ ▲300 - Pipette 300 µL of saline into albumin tube - Mix gently. |

| Checking the Reagent Rack |  |  |  |
| --- | --- | --- | --- |
| Step |  | Prompt | Action |
|  |  | Patient/Order module |  |
| 1. |  | FileOpenReagent Rack |  |
| 2. |  | Rack ID field | Enter 1Query |
| Rack 1 reagents will display. If any are outdated,“Reagent rack or test batch contains expired reagent” displays. |  |  |  |
| 3. |  | OK | Click |
| 4. |  | Compare all reagents currently in use with Rack 1 lot numbers, open dates, and outdates. |  |
|  |  | Replace outdated reagents, new lots, and new bottles of the same lot as follows: |  |
| 5. |  | Scroll to desired reagent | Right click mouse and Insert a Row |
| 6. |  | Using Find button, scroll to desired reagent and lot number, or, using the current reagent as a guide, enter antisera name and lot number. Enter the open date. |  |
| 7. |  | Select replaced reagent row to delete | Right click mouse and Delete a Row |
| 8. |  | SAVE | Click and close |
|  If you Add a Row instead of Insert a Row in the Reagent Rack the order of reagents will change. - Reagents should be checked each day for abnormal appearance, cloudiness or turbidity when verifying the lot number, outdate, and the date opened. - Reagents should be refrigerated when not in use. QC Racks. Rack Name; Testing to be Done; Comments. 1; Daily QC; Gel Methods Weak D included with CELL1 ABD. 2; Antigen Testing; C, c, E, e, Fya, K – gel Jka, M, A1 – tube. 2 Backup; Antigen Testing; C, c, E, e, Fya, K by backup method (tube). ABD Tube; G/Rh; G/Rh by backup method (tube). ALB/LISS QC; ALB & LISS; ABSC, ABID, Crossmatch by backup method. Gel Card Quality Control - Do not remove the foil seal until ready to use or within 1 hour of testing. - Inspect each gel card for a clear liquid layer on top of the opaque gel. - Do not touch the gel card during pipetting to prevent carryover leading to erroneous results. |
| 1. | Bring samples and reagents to room temperature. |  |  |
| 2. | Five (5) cards are required for daily QC testing: 1- IgG, 2-A/B/D Monoclonal & Reverse, 1- A/B/D Monoclonal, 1-Buffered Gel Label as shown  IgG Card 2 - A/B/D Monoclonal & Reverse Cards A; B; D; Ctl; Buf; Buf. Cell 1; Cell 3. a1; b. I; II; I; II. IgG; IgG; IgG; IgG; IgG; IgG. QC Pos; QC Neg. A; B; D; Ctl; Buf; Buf. Cell 2; Cell 4. a1; b. A/B/D Monoclonal Card Buffered Gel Card A; B; D; A; B; D. Cell 1; Cell 2. Buf; Buf; Buf; Buf; Buf; Buf. PXM NXM. a1; b. |
| 3. | Start with the IgG Gel Card: |  |  |
|  | - Pipette 50 µL of 0.8% Screening Cells 1 and 2 into microtubes - Pipette 25 µL of Confidence Antibody into the QC Pos microtubes. - Pipette 25 µL of MTS Diluent 2 into the QC Neg microtubes. |  |  |
| 4. | Place in MTS incubator for 15 minutes. |  |  |
| Using the labeled A/B/D Monoclonal & Reverse, A/B/D Monoclonal, and Buffered Gel cards: |  |  |  |
| 5. | Using BioHit program P1 - Pipette 50 µL of 0.8% a1 Cells into the 3 microtubes (on 3 cards) labeled “a1”. - Pipette 50 µL of 0.8% b Cells into the 3 microtubes (on 3 cards) labeled “b”. |  |  |
| 6. | Using BioHit program P1 - Pipette 50 µL of Confidence Antibody into the 2 microtubes labeled Cell 3. - Pipette 50 µL of 6% Albumin into the 2 microtubes labeled Cell 4. |  |  |
| 7. | Using BioHit program P1 - Pipette 50 µL of Anti-A into the microtubes labeled NXM and PXM |  |  |
| 8. | Using BioHit program P9 - Pipette 10 µL of the prepared 4% QC Cell 1 into the 7 microtubes labeled Cell 1. - Pipette 10 µL of the prepared 4% QC Cell 2 into the 7 microtubes labeled Cell 2. |  |  |
| 9. | If centrifuging all cards together, place the IgG Gel card (after incubation) and the 4 others into the MTS centrifuge for 10 minutes. |  |  |
| 10. | Read reactions immediately following centrifugation. |  |  |
|  | While cards are spinning, perform Weak D testing on QC Cell 1 using the Tube method per ABO Type-Forward and Reverse, Rh-Weak D procedure |  |  |
| 11. | Add Optional Subtest WEAKD to Rh-negative QC Cell 1-ABD (ResultsAdd Optional SubtestWEAKD) |  |  |
| 12. | Grade reactions following Interpreting Reactions procedure. |  |  |

Interpretation Observe the front and back of each microtube for agglutination and/or hemolysis. Refer to Interpreting Reactions procedure. - Negative –No agglutination, no hemolysis, and complete sedimentation of all cells to the bottom of the microtube. - Positive - Red blood cells suspended on the top of the gel or dispersed throughout the gel in varying degrees, or hemolysis. Some positive reactions can have a small button of red blood cells remaining in the bottom of the microtube.

Resulting Quality Control Reagent Test Batch. Step; Prompt; Action. 1.; Patient/Order module; File  New  Reagent Test Batch. 2.; Rack ID field; Enter 1OK. 3.; Select and Go to Result Entry box; Click to √  OK. 4.; Grade reactions according to Interpreting Reactions procedure.. 5.; Record results of reactions, as well as all temperatures, rpms, timer start times, lot number checks, and visual inspection of the blood bank, gel cards, and diluent.. 6.; SAVE; Click and close. 7. If an INVALID or UNTESTED condition exists: ResultsCorrect Test: enter results Incorrect testing: enter INVALID Test ID; Passing Interpretation; Comments; Reaction Fields. CELL1 ABD; AB Neg; Add WEAKD subtest; Enter reaction strengths. CELL2 ABD; O Pos; Enter reaction strengths. CELL1 REV; AB Neg; Enter reaction strengths. CELL2 REV; O Pos; Enter reaction strengths. CELL3 REV; O; Enter reaction strengths. CELL4 REV; AB; Enter reaction strengths. NEG ABSC; Neg; Enter reaction strengths. POS ABSC; Pos; Enter reaction strengths. NEGISXM G; Comp; Enter reaction strengths. POSISXM G; Incomp; Enter reaction strengths. WKD QC2; Neg; Enter reaction strengths. BB Chart; PASS; Enter temperature BB recorder. BB Digital; PASS; Enter temperature digital readout. BB Top; PASS; Enter temperature top shelf. BB Lower; PASS; Enter temperature lower shelf. BB Visual; PASS; OK. FP Chart; PASS; Enter temperaturefreezer recorder. FP Digital; PASS; Enter temperature digital readout. FP Inside; PASS; Enter temperature upper drawer. FP Low; PASS; Enter temperature lower drawer with product. Card Lot; PASS. Dil Lot; PASS. Anti A Lot; PASS. Anti B Lot; PASS. FMBS Kit; PASS. Dil Clear; PASS; CLEAR. Dil 2+Clear; PASS; CLEAR. Well-IgG; PASS; OK. Well-ABD; PASS; OK. Well-Rev; PASS; OK. Well-Buf; PASS; OK. Well-Contr; PASS; OK. Well-Poly; PASS; OK. Temp36-38; PASS; Enter temperature dry bath. Thawer-U; PASS; Enter temperature of upper thermometer. Thawer-L; PASS; Enter temperature of lower thermometer. H2O Level; PASS; OK. Timer (1&2); PASS; Enter time upon starting centrifuge timers. RPM870-920 (1&2); PASS; Enter RPMs from centrifuge readouts. Washer 1; PASS; Enter saline volume. CompCt; PASS. PLT CHART; PASS; Enter temperature from recorder. PLTDIGITAL; PASS; Enter temperature from digital readout. PLTAGITAT; PASS; OK.. Restrictions The cells and antibody reagent contained in the Ortho Confidence Quality Control Kit are for use only as directed and are not suitable for patient testing. If quality control failsPatient results may not be reported until QC issues are resolved Notify Lead Tech if any failure cannot be resolved.. Step; Possible Sources of Error; Action. 1.; QC Kit is contaminated Cells 1 and 2 are contaminated (results unchanged with new antisera); Refer to Reagents Prepared In House for procedure. Cell 1=AB Negative Cell 2=O Positive Repeat testing using newly prepared cells. If acceptable, these cells may be labeled and used for the remainder of the kit shelf life.. Antibody suspension is contaminated; Prepare a new suspension as follows: Add 4 drops of Anti D antisera to 5 mls of saline. Repeat testing using new antibody suspension If acceptable, this suspension may be labeled and used for the remainder of the kit shelf life.. 2.; Gel cards show signs of drying, discoloration, bubbles, crystals, or other artifacts; Repeat testing with acceptable cards.. 3.; Improper storage of reagents; Lead tech will evaluate situation and acceptability of reagents after quality control testing is acceptable. 4.; Equipment failure; If the temperature of the MTS incubator, the timer, or rpms of the MTS centrifuge is not acceptable, contact the BioMed department.. In the event of equipment failure, use back up equipment as follows:. Equipment; Back up Equipment. MTS Centrifuge; Back up centrifuge. MTS Incubator; None (revert to tube methods). Cell Washer; Back up cell washer. Plasma Thawer; Back up plasma thawer. Platelet Agitator; None (return to supplier before 24 hours has elapsed). BioHit Pipettor; Back up BioHit pipettor. Plasma Freezer; Pharmacy freezer. Blood Bank; Back up refrigerator (bacti refrigerator – record temps every 4 hours). Scanner; Back up scanner from another station. Microscope; Hematology microscope. Horizon Blood Bank; Paper methods of result recording. Label Printer; Legible hand written labels on tubes. If a new lot of reagent or gel cards are opened after performing the daily reagent quality control, testing of this new lot must done in conjunction with subsequent patient or component testing. Ordering Single Reagent Quality Control Testing Step; Prompt; Action. 1.; File NewReagent Test Batch; Uncheck Include All Reagents. Determine reagent to be tested and using the following chart, enter corresponding Reagent ID and Test ID. 2.; Reagent ID Manufacturer Lot # Expiration Date Test ID; Enter exactly as indicated in chart Enter manufacturer, lot #, and expiration date of the item in the Reagent ID column, NOT the item you are QCing Enter exactly as indicated in chart. Reagent to be tested; REAGENT ID; TEST ID. Buffered Gel Card; A1 cells-G; NEGISXM G. B cells-G; POSISXM G. A/B/D Monoclonal Grouping Card (ABD/ABD card); MTS Dil2+; CELL1 ABD. IgG,C3d; CELL2 ABD. A/B/D Monoclonal & Reverse Card (ABD/Reverse card); ControlCar; CELL1 REV. BufferCard; CELL2 REV. ABD/ABD; CELL3 REV. ABD/Rev; CELL4 REV. IgG Gel Card; Anti D; NEG ABSC. A2 cells-G; POS ABSC. A1 or B Cells; ABD/ABD; CELL3 REV. ABD/Rev; CELL4 REV. MTS Diluent 2 PLUS; MTS Dil2+; CELL1 ABD. IgG,C3d; CELL2 ABD. MTS Diluent 2; Anti D; NEG ABSC. Screening Cells; Anti D; NEG ABSC. A2 cells-G; POS ABSC. Coombs Control Cells; Saline; WKD QC2. Tube; Anti A or Anti B or Anti D; Anti ATube; CELL 1. Anti BTube; CELL 2. Tube; Anti-IgG; Saline; WKD QC2..

Safety Refer to Laboratory Procedure #4002-1: Infection Control Laboratory Safety Hygiene Procedure and all other safety policies for the lab.

Tech Responsibilities. 1.; All techs must review patient test results for accuracy. Computer warnings and flags must be evaluated and resolved before releasing patient testing. Refer to Daily Operating Procedure.. 2.; If reagent quality control results are not acceptable, one of the instruments does not perform correctly, incorrect tests are reported, tubes are not labeled properly, deviations in transfusion policies occur, or other incidents of unsatisfactory performance occur, document the situation and notify the Lead tech. Remedial action should: - take place immediately after it has been identified - be consistent with defined quality control policies - be consistent with defined preventive maintenance, performance testing and equipment inspection policies - include documentation of repeat patient testing and quality control if indicated - adequately correct all the deficiencies implied in the problem, (for example if one patient’s results are discovered to be incorrect, other patients’ results from the same testing sequence are evaluated to ensure correctness) - review the adequacy of actions taken - document correction of individual results when necessary, notifying the ordering physician and station of the corrected results, issuing a corrected report immediately, retaining original and corrected copies of results.

| Supervisory Responsibilities and Duties |  |
| --- | --- |
| Immediate review: |  |
| - Unacceptable quality control values remain unacceptable after attempt at resolution - Unexplainable incongruent patient results - Patient results do not correlate with patient’s known condition, age, sex, diagnosis, or clinical data - Review of adequacy of corrective action taken - Reporting of incorrect test results - Critical values - Deviations from transfusion policy - Incidents of unsatisfactory performance |  |
| Monthly: |  |
| 1. | Review Blood Bank Daily checklist for completion |
| 2. | Review and document in computer: - Daily reagent QC - All overrides |
| 3. | Review Temperature Charts  Blood Bank Daily Temperature Chart records – temperatures from all sources for the refrigerator agree within 1C, freezer within 2C. - Temperature charts for the refrigerator/freezer – review for temperatures outside the range, temperatures outside the limits are explained, date and initial |
| 3. | Blood Bank Performance Measure – Quality Assessment ▪ Blood Bank Procedures ▪ Blood Bank Transfusions ▪ Blood Bank Phlebotomy |
| 4. | Blood Bank Utilization Reports |
| 5. | Monthly Quality Assurance Checks, Reports and Statistics |
| As necessary: |  |
| Supports timely response to inspections |  |

Related Documents

ABO-Forward and Reverse, Rh-Gel Card

Indirect Antiglobulin Test

Interpreting Reactions

Blood Bank Refrigerator and Freezer Failure

Performance Measure Quality Assessment form

Resulting Testing from the Pending Worklog

Daily Operating Procedure

Computer Downtime Procedure (manual recording)

Daily Temperature Log

Equipment Maintenance

Reagents Prepared in House (procedure and log)

References

MTSTM  Monoclonal Grouping Card Instructions for Use, Ref MTS081115

MTSTM  Monoclonal Control Card Instructions for Use, Ref MTS081001

MTSTM  Buffered Gel Card Instructions for Use, Ref MTS085014

MTSTM IgG Gel Card Instructions for Use, Ref MTS084024

MTSTM Diluent 2 Instructions for Use, Ref MTS9230

Ortho Clinical Diagnostics, Inc. ORTHO Confidence System Instructions for Use, January 2005.

MTSTM  Diluent 2 PLUS Instructions for Use, Ref MTS9230

RMC Quality Control Procedure, 6.1.36A.10

| Date: | Author: |  |  |  |  |  |  | Date: |  | Validated by: |  |  | Date: |  | Approved by: |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8/10/10 | Cheryl Webster, MLS(ASCP)CM |  |  |  |  |  |  | 11/23/10 |  | Jo Lohmann, MT(ASCP) Cheryl Webster, MLSCM(ASCP) |  |  |  |  |  |  |  |  |
| L: Blood Bank/Quality Control-Reagent Quality Control-Gel Method |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2/14/12 |  | CW |  |  | New plasma thawer replacing 2 older ones |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |