| Ridgeview Medical Center Laboratory Services | Blood Bank | Effective Date: 12/15/10 |  |  |
| --- | --- | --- | --- | --- |

Purpose

This procedure provides instructions for the performance of daily temperature recording and instrument monitoring.

Temperatures

| Step | Action |
| --- | --- |
| 1. | All temperatures (refer to Daily Temperature Chart) are taken daily and must meet targets as specified. Temperatures are recorded on Daily Temperature Chart for ease of entry into the reagent test batch in HBB. Any deviation is recorded on Corrective Action Log and referred to Lead Tech or Lab Director, if unresolved. |
| 2. | Chart recorders on the Plasma Freezer, Blood Bank, and Platelet Incubator are changed every Monday, as follows: 1) Open compartment door- push “C” 2) When the stylus moves off the chart, remove paper. 3) Record information on chart as follows: - “Stop” - Stop date and time - Tech initials 4) Record information on new chart: - “Start” - Start date and time - Tech initials 5) Replace chart. 6) Push “Change chart” or “C” and make sure that stylus returns to the correct position. 7) On the Blood Bank Plasma Freezer: Push “Change chart paper” button on digital display. 8) File charts in the appropriate box. |
| 3. | Sorvall Cell Washer CW2+ (#1) Daily Maintenance - Make sure tubing is connected and can drain freely - Remove guard bowl and inspect and wipe out interior bowl - Check saline delivery volume: - CW2+ hold graduated cylinder under delivery port, press ‘Calibrate’. - Volume delivered should be 53-55 mls - Tuesdays - perform the decontamination procedure - Record on log. |

Reagent Test Batch

| Step | Prompt |  | Action |
| --- | --- | --- | --- |
| 1. | Create a new reagent test batch to perform daily QC (this reagent test batch encompasses temperatures, cell washer checks and Daily Reagent Quality Control). |  |  |
| 2. | Patient/Order module | FileNewReagent Test Batch |  |
| 3. | Rack ID field | Enter 1 |  |
| 4. | OK | Click |  |
| 5. | Select and Go to Result Entry field | Click on box to  |  |
| 6. | OK | Click |  |
| 7. | Record all temperatures, rpms, timer start times, lot number checks, and visual inspection of the blood bank in PW (see table below) in conjunction with Daily Reagent Quality Control. |  |  |
| 8. | SAVE | Click and Close |  |
| 9. | Make sure that all entries are COMPLETED, not INVALID or UNTESTED | If an INVALID or UNTESTED exists ResultsCorrect Test and enter results Enter INVAL as interpretation of incorrect testing |  |

Resulting Quality Control

| Test ID |  | Passing Interpretation |  | Comments |  | Reaction Fields |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CELL1 ABD |  | AB Neg |  | Add WEAKD subtest |  | Enter reaction strengths |  |
| CELL2 ABD |  | O Pos |  |  |  | Enter reaction strengths |  |
| CELL1 REV |  | AB Neg |  |  |  | Enter reaction strengths |  |
| CELL2 REV |  | O Pos |  |  |  | Enter reaction strengths |  |
| CELL3 REV |  | O |  |  |  | Enter reaction strengths |  |
| CELL4 REV |  | AB |  |  |  | Enter reaction strengths |  |
| NEG ABSC |  | Neg |  |  |  | Enter reaction strengths |  |
| POS ABSC |  | Pos |  |  |  | Enter reaction strengths |  |
| NEGISXM G |  | Comp |  |  |  | Enter reaction strengths |  |
| POSISXM G |  | Incomp |  |  |  | Enter reaction strengths |  |
| WKD QC2 |  | Neg |  |  |  | Enter reaction strengths |  |
| BB Chart |  | PASS |  | Enter temperature BB recorder |  |  |  |
| BB Digital |  | PASS |  | Enter temperature digital readout |  |  |  |
| BB Top |  | PASS |  | Enter temperature top shelf |  |  |  |
| BB Lower |  | PASS |  | Enter temperature lower shelf |  |  |  |
| BB Visual |  | PASS |  | OK upon visual inspection |  |  |  |
| FP Chart |  | PASS |  | Enter temperaturefreezer recorder |  |  |  |
| FP Digital |  | PASS |  | Enter temperature digital readout |  |  |  |
| FP Inside |  | PASS |  | Enter temperature upper drawer |  |  |  |
| FP Low |  | PASS |  | Enter temperature lower drawer with product |  |  |  |
| Card Lot |  | PASS |  |  |  |  |  |
| Dil Lot |  | PASS |  |  |  |  |  |
| Anti A Lot |  | PASS |  |  |  |  |  |
| Anti B Lot |  | PASS |  |  |  |  |  |
| FMBS Kit |  | PASS |  |  |  |  |  |
| Dil Clear |  | PASS |  | CLEAR upon inspection |  |  |  |
| Dil 2+Clear |  | PASS |  | CLEAR upon inspection |  |  |  |
| Well-IgG |  | PASS |  | OK upon inspection |  |  |  |
| Well-ABD |  | PASS |  | OK upon inspection |  |  |  |
| Well-Rev |  | PASS |  | OK upon inspection |  |  |  |
| Well-Buf |  | PASS |  | OK upon inspection |  |  |  |
| Well-Contr |  | PASS |  | OK upon inspection |  |  |  |
| Well-Poly |  | PASS |  | OK upon inspection |  |  |  |
| Temp36-38 |  | PASS |  | Enter temperature dry bath |  |  |  |
| Thawer-U |  | PASS |  | Enter temperature of upper thermometer |  |  |  |
| Thawer-L |  | PASS |  | Enter temperature of lower thermometer |  |  |  |
| H2O Level |  | PASS |  | OK |  |  |  |
| Timer (1&2) |  | PASS |  | Enter times upon starting centrifuge timers |  |  |  |
| RPM870-920 (1&2) |  | PASS |  | Enter RPMs from centrifuge readouts |  |  |  |
| Washer 1 |  | PASS |  | Enter saline volume |  |  |  |
| CompCt |  | PASS |  |  |  |  |  |
| PLT CHART |  | PASS |  | Enter temperature from recorder |  |  |  |
| PLTDIGITAL |  | PASS |  | Enter temperature from digital readout |  |  |  |
| PLTAGITAT |  | PASS |  | OK |  |  |  |
| If quality control failsPatient results may not be reported until QC issues are resolved Notify Lead Tech if any failure cannot be resolved. |  |  |  |  |  |  |  |
| Step | Possible Sources of Error |  |  |  | Action |  |  |
| 1. | Antisera bottle is contaminated |  |  |  | Repeat testing with new bottle. Enter comment indicating that new bottle was used. If initial bottle is contaminated, discard and record in the Reagent Log and Reagent Rack in computer. |  |  |
| 2. | QC kit is contaminated (results unchanged with new antisera) |  |  |  | Refer to Reagents Prepared In House for procedure. Cell 1=AB Negative Cell 2=O Positive Repeat testing using newly prepared cells. If acceptable, these cells may be labeled and used for the remainder of the kit shelf life. |  |  |
| 3. | Antibody suspension is contaminated |  |  |  | Prepare a new suspension as follows: Add 4 drops of Anti D antisera to 5 mls of saline. Repeat testing using new antibody suspension If acceptable, this suspension may be labeled and used for the remainder of the kit shelf life. |  |  |
| 4. | Improper storage of reagents |  |  |  | Lead tech will evaluate situation and acceptability of reagents after quality control testing is acceptable |  |  |
| 5. | Equipment failure |  |  |  | If the temperature of the MTS incubator, the timer or rpms of the MTS centrifuge are not acceptable, contact the BioMed department. |  |  |
| In the event of equipment failure, use back up equipment as follows: |  |  |  |  |  |  |  |
| Equipment |  |  | Back up Equipment |  |  |  |  |
| MTS Centrifuge |  |  | None (revert to tube methods) |  |  |  |  |
| MTS Incubator |  |  | None (revert to tube methods) |  |  |  |  |
| Cell Washer |  |  | Back up cell washer |  |  |  |  |
| Plasma Thawer |  |  | Back up plasma thawer |  |  |  |  |
| Platelet Agitator |  |  | None (return to supplier before 24 hours has elapsed) |  |  |  |  |
| BioHit Pipetor |  |  | Back up MTS Pipettor |  |  |  |  |
| Plasma Freezer |  |  | Tissue freezer in surgery |  |  |  |  |
| Blood Bank |  |  | Back up refrigerator (bacteriology refrigerator – record temps every 4 hours) |  |  |  |  |
| Scanner |  |  | Back up scanner from another station |  |  |  |  |
| Microscope |  |  | Hematology microscope |  |  |  |  |
| Horizon Blood Bank |  |  | Paper methods of result recording |  |  |  |  |
| Label Printer |  |  | Legible hand written labels on tubes |  |  |  |  |

Safety

Refer to Laboratory Procedure #4002-1: Infection Control Laboratory Safety Hygiene Procedure and all other safety policies for the lab.

| Tech Responsibilities |  |
| --- | --- |
| 1. | All techs must review patient test results for accuracy. Computer warnings and flags must be evaluated and resolved before releasing patient testing. Refer to Daily Operating Procedure. |
| 2. | If reagent quality control results are not acceptable, one of the instruments does not perform correctly, incorrect tests are reported, tubes are not labeled properly, deviations in transfusion policies occur, or other incidents of unsatisfactory performance occur, document the situation and notify the Lead tech. Remedial action should: - take place immediately after it has been identified - be consistent with defined quality control policies - be consistent with defined preventive maintenance, performance testing and equipment inspection policies - include documentation of repeat patient testing and quality control if indicated - adequately correct all the deficiencies implied in the problem, (for example if one patient’s results are discovered to be incorrect, other patients’ results from the same testing sequence are evaluated to ensure correctness) - review the adequacy of actions taken - document correction of individual results when necessary, notifying the ordering physician and station of the corrected results, issuing a corrected report immediately, retaining original and corrected copies of results |
| Supervisory Responsibilities and Duties |  |
| Immediate review: |  |
| - Unacceptable quality control values remain unacceptable after attempt at resolution - Unexplainable incongruent patient results - Patient results do not correlate with patient’s known condition, age, sex, diagnosis, or clinical data - Review of adequacy of corrective action taken - Reporting of incorrect test results - Critical values are reported - Deviations from transfusion policy - Incidents of unsatisfactory performance |  |
| Monthly: |  |
| 1. | Review Blood Bank Daily checklist for completion |
| 2. | Review and document in computer: - Daily reagent QC - All overrides |
| 3. | Review Temperature Charts  Blood Bank Daily Temperature Chart records – temperatures from all sources for the freezer agree within 2C and the refrigerator/incubator within 1C. - Temperature charts for the refrigerator/freezer – review for temperatures outside the range, temperatures outside the limits are explained, date and initial |
| 3. | Blood Bank Performance Measure – Quality Assessment ▪ Blood Bank Procedures ▪ Blood Bank Transfusions ▪ Blood Bank Phlebotomy |
| 4. | Blood Bank Utilization Reports |
| 5. | Monthly Quality Assurance Checks, Reports and Statistics |
| As necessary: Supports timely response to inspections |  |

Related Procedures

ABO Type-Forward and Reverse, Rh-Weak D

Indirect Antiglobulin Test

Interpreting Reactions

Blood Bank Refrigerator and Freezer Failure

Performance Measure Quality Assessment form

Resulting Testing from the Pending Worklog

Daily Operating Procedure

Computer Downtime Procedure (manual recording)

Daily Temperature Log

Equipment Maintenance

Reagents Prepared in House (procedure and log)

Quality Control-Reagent Quality Control

References

Ortho Clinical Diagnostics, Inc. ORTHO Confidence system package insert, January 2005.

Standards for Blood Banks and Transfusion Services, 25th Edition, 2008, pages 13, 101.

AABB Technical Manual, 15th Edition, 2005, pages 8-28.

| Date: |  | Author: |  |  |  |  | Date: |  | Validated by: |  |  |  |  | Date: |  | Approved by: |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4/23/2008 |  | Mary Farlow MT(ASCP) SBB Jo Lohmann, MT(ASCP) Cheryl Webster, MT(ASCP) |  |  |  |  | 6/18/08 7/1/08 12/1/10 |  | Jo Lohmann, MT(ASCP) Cheryl Webster, MT(ASCP) Jo Lohmann, MT(ASCP) |  |  |  |  | 6/19/2008 |  | Dr. T. Hallin |  |  |
| L: Blood Bank/Quality Control-Daily Temperature Recording and Instrument Checks |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 5/10 | JL |  |  | Add Cell 4 to rack, known negative reactions for back typing cells Dr. Weatherby, MD 6/3/2010 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 10/10 | CW |  |  | Break out temps and instrument checks from reagent QC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 11/2/2012 | CW |  |  | Add platelet incubator temperature checks |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Discontinued |  |  |
| --- | --- | --- |
| Date: | Reason: | Initials: |