| Ridgeview Medical Center Laboratory Services |  |  |  |  |  | Blood Bank |  |  | Effective Date: 9/26/08 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose This procedure provides directions for preparing the following reagents when necessary due to spillage or contamination: Coombs Control Cells, back typing cells, and Cell I and Cell II in the Confidence Kit. In house prepared cells are used only until replacement reagents arrive. |  |  |  |  |  |  |  |  |  |  |
| Reagents |  | Supplies |  |  |  |  | Equipment |  |  |  |
| - Saline - A1 Lectin - Anti-D - 0.8% A1 cells - 0.8% A2 cells - MTS Diluent 2 Plus - ABD/ABD cards |  | - Pipets - Glass tubes - Pipet tips - Plasma vials |  |  |  |  | - MTS incubator - Cell washer, centrifuge - Eppendorf 100 l pipet - MTS BioHit pipettor |  |  |  |
| Coombs Control Cells |  |  |  |  |  |  |  |  |  |  |
| Step | Action |  |  |  |  |  |  |  |  |  |
|  | Remove one segment from 4-5 units of O Positive red cells. |  |  |  |  |  |  |  |  |  |
|  | Wash the cells three times by hand to remove the plasma and EDTA. |  |  |  |  |  |  |  |  |  |
| 3. | Measure approximately 1 ml of pooled cells into a glass tube. |  |  |  |  |  |  |  |  |  |
| 4. | Place 2 mls of saline in a glass tube. Add 10 drops of Anti D and mix well. |  |  |  |  |  |  |  |  |  |
| 5. | Add the saline /Anti-D mixture to the pooled cells and mix well. Divide this mixture into 2 glass tubes so that the liquid level is below the top edge of the 37C incubator well. |  |  |  |  |  |  |  |  |  |
| 6. | Incubate at 37C for 60 minutes. Mix the tube intermittently every 15 minutes during the incubation period. |  |  |  |  |  |  |  |  |  |
| 7. | Wash the antisera coated cells three times by hand, decanting completely after the last wash. |  |  |  |  |  |  |  |  |  |
| 8. | Prepare a 3% cell suspension. (1ml saline and 30 l of washed cells) |  |  |  |  |  |  |  |  |  |
| 9. | Confirm the reactivity of the prepared cell suspension by performing quality control as indicated on page 7 of Quality Control procedure (6.1.33.10). Record the date, initials, interpretation, reactions, unit numbers of source of cells units, and the acceptability on the Prepared Reagents Testing Log. |  |  |  |  |  |  |  |  |  |
| 10. | Once the quality control is acceptable, the cell may be placed in use. Place the cell suspension in a clean plasma vial. Write the name of the cell suspension, lot number (RMC+date made), date prepared, expiration date, and tech initials. |  |  |  |  |  |  |  |  |  |
| A1 Back Typing Cells |  |  |  |  |  |  |  |  |  |  |
| 1. | Find A1 positive A Negative cells by checking MBC packing slips for A1 typing. Pull one segment from each A1 positive unit. If no typing is indicated, pull one segment from each available A Negative unit. |  |  |  |  |  |  |  |  |  |
|  | Controls |  |  |  |  |  |  |  |  |  |
|  | Positive Control: 0.8% A1 cells Negative Control: 0.8% A2 cells These need to be concentrated to 3% as follows: |  |  |  |  |  |  |  |  |  |
| 2. | Label one 12 X 75 tube for a1 cells and one for a2 cells. |  |  |  |  |  |  |  |  |  |
| 3. | Place 4 volumes (e.g. 200 l) of each 0.8% cell suspension into the appropriate tube. |  |  |  |  |  |  |  |  |  |
| 4. | Spin for 60 seconds using Cell Washer/Centrifuge. |  |  |  |  |  |  |  |  |  |
| 5. | Decant the supernatant completely. |  |  |  |  |  |  |  |  |  |
| 6. | Add one volume (e.g. 50 l) of saline to each tube and mix gently to resuspend. |  |  |  |  |  |  |  |  |  |
|  | Units to be tested |  |  |  |  |  |  |  |  |  |
| 7. | Make a 3% red cell suspension of each unit in saline, using program P7 on the BioHit pipettor. |  |  |  |  |  |  |  |  |  |
| 8. | Label tubes – unit(s), positive control (a1), negative control (a2) |  |  |  |  |  |  |  |  |  |
| 9. | Add 1 drop of each cell suspension to the appropriate tube. |  |  |  |  |  |  |  |  |  |
| 10. | Add 1 drop of A1 Lectin to each tube. |  |  |  |  |  |  |  |  |  |
| 11. | Mix tubes gently and centrifuge for 15 seconds at 3400 RPM. |  |  |  |  |  |  |  |  |  |
| 12. | While observing in the mirror, resuspend the cell button completely and observe for agglutination macroscopically. |  |  |  |  |  |  |  |  |  |
| 13. | Interpret the reactions as negative, 1+, 2+, 3+, or 4+. Record results on log. A1 positive units are appropriate for use. |  |  |  |  |  |  |  |  |  |
|  | Preparation of cell suspension |  |  |  |  |  |  |  |  |  |
| 14. | Pull one segment from each unit and empty segments into a glass tube. Wash 3 times by hand. |  |  |  |  |  |  |  |  |  |
|  | Prepare approximately 6 mls of a 0.8% cell suspension as follows: |  |  |  |  |  |  |  |  |  |
| 15. | Place 2 mls of MTS Diluent 2 Plus into 3 tubes. |  |  |  |  |  |  |  |  |  |
| 16. | Make a 0.8% solution in each tube by using program P6 on the BioHit pipettor – twice in each tube. |  |  |  |  |  |  |  |  |  |
| 17. | Mix all tubes gently and combine into one vial. |  |  |  |  |  |  |  |  |  |
| 18. | Label : 0.8% ______ Cells Lot #: RMC + date made Date prepared/Tech initials Expiration date: one week from date made Storage temperature: refrigerate |  |  |  |  |  |  |  |  |  |
| 19. | Perform Quality Control – see page 7 of Quality Control procedure(6.1.33.10) |  |  |  |  |  |  |  |  |  |
| 20. | Record all information on the Prepared Reagents Testing Log. |  |  |  |  |  |  |  |  |  |
| B Back Typing Cells |  |  |  |  |  |  |  |  |  |  |
| Using 1 to 2 B Negative units  proceed to Steps 14-20 above |  |  |  |  |  |  |  |  |  |  |
| Confidence Kit Cell 1 |  |  |  |  | Confidence Kit Cell 2 |  |  |  |  |  |
| Using segments from 1 to 2 AB Negative units |  |  |  |  | Using segments from 1 to 2 O Positive units  |  |  |  |  |  |
| 1. | Place cells from segments into a glass tube and wash 3 times by hand. |  |  |  |  |  |  |  |  |  |
| 1. | Make two 3% solutions – one with saline, one with MTS Diluent 2 Plus. |  |  |  |  |  |  |  |  |  |
| 2. | To confirm acceptability of pooled cells, test the MTS Diluent 2 Plus cell suspension in ABD/ABD card. |  |  |  |  |  |  |  |  |  |
| 3. | Record all information on the Prepared Reagents Testing Log. |  |  |  |  |  |  |  |  |  |
| Bring Reagent into Inventory |  |  |  |  |  |  |  |  |  |  |
| 1. | Inventory module  File  New  Delivery |  |  |  |  |  |  |  |  |  |
| 2. | Supplier drop down box |  | Enter RMC |  |  |  |  |  |  |  |
|  | Delivery window |  |  |  |  |  |  |  |  |  |
| 3. | Type field |  | Enter G (reagent) |  |  |  |  |  |  |  |
| 4. | Item ID field |  | Use binoculars  OK |  |  |  |  |  |  |  |
| 5. | Quantity |  | 1  OK |  |  |  |  |  |  |  |
| 6. | Place cursor in Item ID field |  | Double-click |  |  |  |  |  |  |  |
|  | Reagent Receipt window |  |  |  |  |  |  |  |  |  |
| 7. | - Manufacturer - Lot. No. - Expiration Date - Rec’d Directions for Use - Rev. to Package Insert? - Visual Inspection - Quantity Received |  |  | Enter RMC RMC + date made (i.e. RMC53110) Enter 1 week from date made Y N Y 1 |  |  |  |  |  |  |
| Placing Reagents on the QC Rack: |  |  |  |  |  |  |  |  |  |  |
| 8. | File  Open  Reagent Rack |  |  |  |  |  |  |  |  |  |
| 9. | Rack ID field |  |  | Enter 1  Query |  |  |  |  |  |  |
| 10. | Scroll to desired reagent |  |  | Right click and choose Insert a Row |  |  |  |  |  |  |
| 11. | Use FIND |  |  | Scroll to desired reagent and select |  |  |  |  |  |  |
| 12. | Highlight row to delete |  |  | Right click and choose Delete a Row |  |  |  |  |  |  |
| 13. | SAVE icon |  |  | Click and close. |  |  |  |  |  |  |

| Interpretation The acceptability and specificity of the in house prepared cell suspensions will be determined according to the results of testing with known, manufacturer approved antisera. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Limitations Reagents prepared in house will not be used for longer than one week and only until replacement manufacturer’s approved cells can be obtained, with the exception of the Confidence Cells which may be used until gone. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety Refer to Laboratory Procedure: Infection Control, Laboratory Safety Hygiene Procedure and all other safety procedures of the laboratory. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Related Documents Prepared Reagents Log Quality Control Procedure A1 Lectin Procedure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| References AABB Technical Manual, 15 Edition, 2005, page 11, 715. Standards for Blood Banks and Transfusion Services, 25th Edition, 2008, page 13. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Author: |  |  |  |  |  | Date: |  | Validated by: |  |  |  |  | Date: |  | Approved by: |  |  |  |
| 7/5/08 |  | Jo Lohmann, MT(ASCP) |  |  |  |  |  | 7/8/09 |  | Cheryl Webster, MT(ASCP) |  |  |  |  | 12/08 |  | T. Hallin, MD |  |  |  |
| Reagents Prepared In House |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 6/2/10 | CW |  |  |  | Addition of directions for placing reagent on QC rack Dr. R. P Weatherby, MD |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 7/26/11 | CW |  |  |  | Revision of A1 section to include Lectin procedure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |