| C.difficile – Illumigene/C.diff/Clostridium difficile |  |
| --- | --- |
| Ridgeview Medical Center Laboratory Services | Department of Origin: Microbiology |

Intended Use

Principle:

Toxigenic Clostridium difficile is a major cause of antibiotic associated diarrhea and colitis and is the causative agent for virtually all cases of pseudomembranous colitis.  About 2% of normal healthy adults are colonized with C. difficile.  Exposure to most antibiotics is thought to allow proliferation of toxigenic C.l Difficile by distrupting the normal intestinal flora.  Two large toxin proteins (TcdA [or toxin A] and TcdB [toxin B]) are thought to be the primary virulence factors of C. difficile.  These toxins are encoded by two separate genes, named tcdA and tcdB, respectively.  Together, with three additional genes, they form a 19.6kb pathogenicity locus called PaLoc.

The illumigene® C. difficile DAN molecular assay is based on loop-mediated amplification technology, which uses specifically designed primers to the PaLoc pathogenicity locus to provide for specific and continuous isothermal DNA amplification.  A by-product of this amplification is the formation of magnesium pyrophosphate, which forms a white precipitate leading to a turbid reaction solution.  This presence of turbidity signifies a positive reaction while the absence of turbidity represents a negative reaction.  The illumigine C. difficile assay contains primers that specifically amplify a 204 pb region of the conserved 5’ sequence of the tcdA gene within the PaLoc of toxigenic C. difficile in diarrheal stool samples from patients suspected of having C. difficile-associated disease.

The illumigene C. difficile DNA molecular assay kit includes a Sample Preparation Apparatus and a Test Device.  The illumigene C. difficile Test Device contains two chambers; a TEST Chamber with C. difficile specific primers and a CONTROL Chamber with S. aureus specific primers.  The patient specimen and the S. aureus Control Target are processed in parallel through the Specimen Preparation steps and through the Test Procedure.  The internal Control Monitors amplification inhibition, assay reagent performance and sample processing effectiveness.  The Control S. aureus DNA target must be amplified and detected in the final reaction or the test is considered invalid and patient results are not reported.

Specimen:

Preferred Sample Types:

- Human stool samples, unpreserved

- Human stool samples, preserved in Cary-Blair-based media

- Unformed samples indicative of CDAD

Criteria for Rejection:

- Formed stool specimens that do not conform to the shape of the container or that can support an

applicator stick (flag pole test).  Patients with CDAD will produce 3 or more diarrheal stools per day.

- Replicate testing within 7 days of a negative test.

- Replicate test request within 10 days of a positive result.

Cancel the test in the computer with the appropriate cancelation comment.  If the specimen is a formed stool inform the physician that the test is being cancelled due to consistency.  Save the specimen in the large blue “Specimens to save” bin.

Cancelation comments:

- Formed stool specimen: 2 Cancelled: formed stool not acceptable for CDiff, plus Code: \FORMED: Unacceptable specimen received for testing.  Only liquid stools that form to the shape of the container are acceptable unless ileus is present.  Contact Lab within 3 days to request variance.

- Repeat testing - 1 Cancel: Only one CDIFF test allowed per 7 days. Auto-cancel on order.

- Previous Positive: 1 Cancel other; see comment - Code: DUPCDIFF  -  If a patient is positive for C.difficile, no additional specimens will be accepted for 10 days.  The assay is not approved to test for cure, since nucleic acids may persist after effective treatment.

Collection and Storage:

Human stool samples, unpreserved:  Samples should be transported and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 24 hours at 21-27 C or 5 days at 2-8 C.  Samples that will not be tested within these times should be frozen immediately upon receipt and stored at ≤ -20 C until tested.  Specimens may be frozen and thawed once.

Human stool samples, preserved in Cary-Blair-based media:  Samples should be transported and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 5 days at 2-8 C.  Samples that will not be tested with this time should be frozen immediately upon receipt and stored at ≤ -20 for seven days prior to testing.  Specimens may be frozen and thawed once.

Materials

Materials Provided:

- illumigene® Sample Preparation Apparatus:

- illumigene® Reaction Buffer:

- illumigene® difficile Test Device:

- Sample Collection Brushes

- illumigene® Heat Treatment Tubes

Materials not provided:

- illumigene® External Control Kit – Negative and Positive Control

- Disposable latex gloves, powder free

- DNAse/RNAse free, aerosol resistant pipette tips.

Equipment:

- Dry-bath with 12mm heat block capable of 95 C

- Vortex Mixer

- Interval Timer

- 50 µL micropipette

- 200 µL micropipette

- illumipro-10™ Incubator/Reader

- Digital Thermometer with Max/Min Temperature Memory

Reagent Preparation:

Ensure kit reagents are at room temperature (21 – 27 C) before use.  Incorrect results may be obtained if reagents are not brought to room temperature prior to use.

Precautions:

- Follow Biosafety Level 2 and Good Laboratory Practices during testing.  Treat all specimens and used test devices as capable of transmitting infectious disease.

- Wear disposable gloves while handling specimens and thoroughly wash hands afterwards.

- illumigene® Sample Dilution Buffer contains formalin inactivated organism and should be treated as potentially infectious.

- The illumigene® C. difficile Test Device contains lyophilized reagents.  The protective pouch should not be opened until ready to perform the assay.

- The illumigene® C. difficile Test Device includes a latch feature that is designed to prevent contamination of the test area with amplification produce.  Do NOT use Test Devices with broken latches.

- Dispose of used illumigene® Test Devices immediately after processing, leaving the device latch securely in place.  Opening the device after amplification may result in contamination of the test area with amplification product.

- When not in use, the Illumipro-10 should be stored with the lid closed.

Shelf Life and Storage:

The expiration date is indicated on the kit label.  Store the kit at 2 – 27 C.

Kits are marked with date received at the time of receipt.

Do not use the devices or reagents after their expiration dates.

Quality Control

- Each device contains an internal control well that controls for amplification inhibition, assay reagents and sample processing effectiveness.

- The heat treatment step is monitored with an external thermometer and interval timer.  Use the Max/Min temperature memory of the thermometer to ensure that a temperature of 95 +/- 5 C is maintained.  Use the interval timer to ensure that heat-treatment duration is 10 +/- 2 minutes.

- Good laboratory practice recommends the use of control materials.  Users should follow the appropriate federal, sate and local guidelines concerning the running of external quality controls.

- illumigene®  C. difficile External Control Reagents are supplied separately (Catalog 279920). It is recommended that the reactivity of each new lot and each new shipment of Illumigene® C. difficile be verified on receipt and before use.

External Control Procedure

  - Meridian External controls are plasmid DNA and heat treatment is not necessary.

  - Remove and label one Red Capped Reaction Buffer Tube for the positive control and one for the negative control.

  - Using a precision pipette, transfer 50µL of the positive control to the reaction buffer tube labeled positive control.

  - Using a new pipette tip, transfer 50µL of the negative control to the reaction buffer tube labeled  negative control.

  - Remove one Illumigene C.difficile Test Device for each control from its pouch.

  - proceed with step 5 of Test Procedure.

QC Testing Frequency and Documentation:

- External QC is run monthly (The first day of the month) and with each new lot number and/or each new shipment of reagents.

- Results of External QC and actions taken when control results are unacceptable are:

  - Do not report patient results.

  - Repeat controls, if still unacceptable contact the Lead Technologist and/or Meridian Technical Service.

  - Document all actions taken on C.Diff Log Sheet with control results.

Quality Assurance

Quarterly

Send 1 Positive patient sample and 1 Negative patient sample to Reference Laboratory for comparison

testing. Part of the specimen should be retained to resolve discrepancies. Environmental contamination must

be ruled out if false positive is suspected.  Consult with pathologist and Meridian if results discrepant.

Environmental screening for DNA contamination is performed Monthly on the first Tuesday of each month.

- Using a polyester swab (swab from rapid strep kit), swab the following in order:

- Inside of illumipro-10

- Outside of illumopro-10

- White testing rack

- Heat block

- Illumipro bench

- BSC (hood) work surface

- Break off swab in Illumigene Sample Preparation Apparatus.

- Recap and vortex 10 seconds.

- Continue to process and test according to the Illumigene procedure.

Results

Negative – No further action required.

Positive – Wipe down all testing areas with 10% bleach followed by Sani-Cloth.

Retest to assure all areas are free of contamination, testing hood work surface separately

before resuming patient testing.  Notify Lead Technologist.  Review recent results for increased

positive tests. Notify patient care provider if necessary.

Positive CDiff  rates will be monitored every 6 months for any large increase in positives that cannot be explained.  Corrective action: An investigation to determine if false positives are occurring because of contamination.

Safety

Follow Biosafety Level 2 and Good Laboratory Practices during testing.  Treat all specimens and used test devices as capable of transmitting infectious disease.

Wear gloves while handling specimens and thoroughly wash hands afterwards.

Dispose of test materials in the large Biohazard container or the Blood culture Biohazard container.

Procedure

NOTE:  Ensure that the illumipro-10™ instrument is powered on and required performance verifications have been completed prior to initiation of SPECIMEN PREPARATION.  Refer to the illumipro-10™ Operator’s

Manual for further information regarding instrument set-up and operation.

Specimen Preparation (Performed in the Biological Safety Cabinet):

- Mix stool sample thoroughly.

- Collect mixed sample specimen using Sample Collection Brush					.

  - Liquid Stool:  Immerse Sample Collection Brush completely into specimen.

  - Semi-solid Stool:  Rotate Sample Collection Brush over specimen surface, lightly coating the    surfaces of the brush bristles.  Approximately one-half of the brush should be coated; over-collection of stool may lead to clogging of the sample collection device or cause an Invalid result.

- Add the Sample Collection Brush to the Sample Preparation Apparatus containing Sample Diluent.  Break-off the Sample Collection Brush at the ‘Break’ mark on the handle.  Replace and secure the Sample Preparation Apparatus cap.  For stool in Cary-Blair media, transfer 200 µL of specimen to the illumigene® C. difficile Sample Preparation Apparatus and secure the cap.  Vortex the Sample Preparation Apparatus for a minimum of 10 seconds.  Specimen in illumigene® C. difficile Sample Preparation Apparatus prior to heat treatment may be held at 2-27 C for up to 24 hours prior to testing.

- Remove the tip cap from the Sample Preparation Apparatus and squeeze five to ten drops of sample into a clean illumigene® Heat Treatment Tube.

- Repeat Sample Preparation Steps for all samples to be processed.

- Heat the illumigene® Heat Treatment Tubes containing mixed sample in a dry-bath/heat block at 95 C +/- 5 C for 10 +/- 2 minutes.  Use the digital thermometer and the interval timer to monitor temperature and time.

- Remove the tubes from the dry-bath/heat block and vortex for approximately 10 seconds.  Heat treated Samples may be held at 2-8ºC  for 4 hours OR may be frozen and stored at ≤ -20ºC for 1 day prior to addition to Reaction Buffer.   Heat treated samples may be frozen and thawed once.

- Decontaminate the specimen preparation area with 10% bleach when finished.  Remove the bleach with a wet paper towel.

Test Procedure (Perform at Illumigene bench):

This test should be performed by qualified personnel per local regulatory requirement.

Note:  A maximum of ten samples can be processed in a single illumipro-10™ run.

- Transfer 50µL of heat treated sample to an appropriately labeled illumigene® Reaction Buffer tube.

- Vortex the Reaction Buffer Tube containing heat treated sample for approximately 10 seconds.

- Repeat steps 1 and 2 for all the samples to be analyzed before proceeding.

- Remove one illumigene® C. difficile Test Device from its protective pouch per sample.  Carefully open the device, holding the tubes such that the lyophilized reagent will not fall out upon opening.  Place device on a flat surface or in a rack that can accommodate the device.

- Using a new pipette tip, transfer 50 µL from the Reaction Buffer tube containing heat treated sample to the TEST chamber (White Bead) of the illumigene Test Device.  Do not introduce air bubbles.  Using a new pipette tip, transfer 50 µL from the reaction buffer tube containing heat treated sample to the CONTROL chamber (Yellow Bead) of the illumigene device.  Do not introduce air bubbles.  Close the illumigene Assay Device and fasten the latch securely.

- Gently mix the device by tapping or flicking being careful to not introduce air bubbles or get the reaction mixture into the device lid.  Carefully examine the reaction tubes for air bubbles or liquid in the lid.  If air bubbles or liquid is noted, tap the device on the bench top and repeat visual inspection.

- Initiate C.Diff run.  See Amplification and Detection section below.   Results will be displayed at the conclusion of the run.

- Decontaminate the Illumigene rack after each use with 10 % bleach.

Amplification and Detection

Initiate C. difficile Run

- To run specimen on the illumipro-10, press#1 on the main menu for Block A or press#2 for Block B.

- Press#1 “C.difficile”

- Press down and lift the lid.

- Remove any sample tubes in the instrument.

- Press enter.

- Close lid.

- Empty Well Test is performed.  Instrument will click once complete.

Amplification

- Press down and lift the lid.  User has two minutes to insert devices and initiate run.

- Check each device to ensure no bubbles are present and no liquid in the cap.

- Insert samples.

- Close lid.

- Press RUN.

- The timer will begin counting down.

Entering Sample IDs

- Once run is initiated, “Edit sample ID” is displayed on the screen of the instrument.

Sample IDs may be edited at any time prior to the completion of the incubation.

- Select the well of the sample ID to be entered.

- Enter sample ID by using keypad, keyboard, or barcode scanner.

- Press enter.

- Repeat steps 2-4 for each sample tested.

Print Results

- Once the run is completed the illumipro-10 screen will display zeros, 00:00.

- Press #1 for Block A or press #2 for Block B and the results will then print.  The status of each block will now be --:--, indicating the instrument is ready for another run.  Attach the print out tape to the Illumigene result log sheet for the run.  Also record the run temperature minimum and maximum

- Results can also be viewed on the screen by well number.  Press the enter key to scroll down to see results in the last few runs.

- To review and print previous result, Press #3 “Results” on the main menu.

- Press #1 “View results”.

- Select the number from the keypad of the run that you want to print or scroll the cursor up or down to the run and press enter.

- Selected results will automatically print.

Deleting Results

- From the Main Menu Screen Press #3 for “Results”

- From the Stored Results Menu Press #2 for “Delete All”.

Note: This will delete all results stored on the illumipro-10.

3.   From the Delete All Menu select “Cancel” or “Delete”.

Interpretation of Results:

| Sample ID | Reported Result | Interpretation |
| --- | --- | --- |
| Patient Specimen | POSITIVE | Sample contains toxigenic C. difficile strain with the pathogen locus (PaLoc). |
|  | NEGATIVE | No toxigenic C. difficile detected. |
|  | INVALID | No reportable result. Repeat the test using the original stool sample. Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure. See Reporting of Results section |

| Sample ID | Reported Result | Interpretation |
| --- | --- | --- |
| Positive Control | POSITIVE | Valid positive control result. Reagents active at time of use, illumipro-10 performing correctly. |
|  | NEGATIVE | Incorrect control results. Repeat control testing as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | INVALID | No reportable result. Repeat entire assay run using original stool samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |

| Sample ID | Reported Results | Interpretation |
| --- | --- | --- |
| Negative Control | POSITIVE | Incorrect control results. Repeat control testing as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | NEGATIVE | Valid negative control result. Reagents active at the time of use, illumipro-10 performing correctly. |
|  | INVALID | No reportable result. Repeat entire assay run using original stool samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |

| Sample ID | Reported Results | Interpretation |
| --- | --- | --- |
| Empty Well | NONE | No illumigene Test Device in the illumipro-10 Well. OR The illumigene Test Device present is compromised due to sample preparation failure, dirty device or improperly seated device. Repeat the test using original sample. |

Reporting of Results::

Positive Test:  Sample contains toxigenic C. difficile strain with the pathogen locus (PaLoc)

Negative Test:  No toxigenic C. difficile detected.

Invalid Test:  Cancel test.  Add comment Code:\CDINVALID - Invalid due to inhibitory substances.

Test referred to Reference Laboratory (to credit our test).

Order a CDiff by PCR to Reference Laboratory.

Call physician or nurse and request another specimen if invalid test is on inpatient.

Invalid Results may be caused by:

- Over-inoculation – Repeat test with less sample.  If invalid result persists, the specimen may be frozen and then repeated or put the specimen in C&S (Cary Blair) before repeating.

- Excess turbidity in the initial reading can lead to invalid results or an “empty well” result.

- If the heat block is not at 90-100C, cells will not lyse releasing DNA preventing amplification

- If the heat block temperature is too high and/or the specimens are heated for a prolonged period of time, the DNA will begin to denature and fail to replicate.

- If the pipetting steps are not completed in the correct order and into the correct vial, an invalid result will occur.

- If the control well was over or under inoculated, the turbidity will be over or under diluted causing an invalid result.

Results are entered into the Horizon Lab System using the Illumigene C.Diff worksheet.

Expected Results:

The frequency of antibiotic-associated diarrhea caused by C. difficile is dependent on several factors, including patient population, type of institution, and epidemiology.  The incidence of CDAD in patients suspected of having antibiotic-associated disease is 15-20%, although different facilities may find positive rates above or below this range.  The incidence of toxigenic C. difficile during the 2010 period of this study was 15.5% and 16.1%.

Limitations of the Procedure:

- Assay does not distinguish between viable and non-viable organisms.

- Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

- This test is for use with unformed stools preserved or unpreserved stools only.  Performance characteristics of other clinical specimen types have not been established.

- Two distinct groups have been identified that can harbor C. difficile asymptomatically a very high rates.  Colonization at rates up to 50% and higher have been reported in infants and rates up to 32% in cystic fibrosis patients.

- This assay does not identify antimicrobial susceptibility.

- This test detects but does not differentiate the NAP1 (Ribotype 027) strain from other toxigenic strains of C. difficile.

- The detection of bacterial nucleic acid is dependent on proper specimen collection, handling (including transportation and storage) and preparation (dilution and heat treatment).  Failure to follow instructions for collection, handing and preparation may cause incorrect results.

Performance Characteristics::

Refer to Directional Insert-Meridian Bioscience illumigene C. difficile.

References:

Refer to Directional Insert-Meridian Bioscience illumigene C. difficile.  SN11176  Rev. 04/11

Illumipro-10 Operator’s Manual, Meridian Bioscience, Inc. SN11007 Rev. 07/2010

Appendix

Daily Maintenance Tasks

General cleaning

System self check for illumipro-10

Refer to Illumipro-10 Maintenance Procedure

Monthly Maintenance Tasks

Optical verification

Refer to Illumipro-10 Maintenance procedure

CDiff Illumigene Molecular Assay Workflow

- Test all in-patients when received on all shifts.  Out-patients will be tested daily in batch.

- Specimen labels will print when ordered.  Micro labels will print when received.  Only the large specimen labels can be scanned on the instrument.

- Workflow

- Check to see if daily maintenance has been done before test run.  Record that you have done daily maintenance.

- Testing steps up heating will be performed under the hood.

- Clean the hood, Illumigene bench, instrument and heat block with 10% bleach followed by Sani-Cloth before starting and when done. White testing racks are cleaned with 10%bleach followed by Sani-Cloth after each use.

- Clean any spills immediately with 10% bleach followed by Sani-Cloth.

- Change gloves immediately if they become soiled.

- Fill out a Specimen Log Sheet for each Run.

Record kit lot #/Exp. Date

Date/tech

Heat block MAX/MIN temperature

- Discard pipette tips in Safe-keeper container to be changed daily.

- Discard testing supplies in Chemistry biohazard (preferred) or large Micro biohazard container.

- Staple print out tape to log sheet.

| Date: |  | Author: |  |  |  |  |  | Date: |  |  | Validated by: |  |  |  |  | Date: |  |  |  | Approved by: |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 11/30/2011 |  | Elaine Krueger |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 08/15/2013 | EK |  |  | Addition of QA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |