| Ridgeview Medical Center Laboratory Services |  | Blood Bank |  |  | Effective Date: 9/26/08 |  |
| --- | --- | --- | --- | --- | --- | --- |
| Principle This procedure provides directions for the determination of the presence or absence of the A1 antigen by testing red cells with Anti A1 Lectin. Anti A1 Lectin is a purified extract of the seeds of Dolichos biflorus and agglutinates only red blood cells containing the A1 antigen. Anti A1 Lectin can be used to resolve discrepancies between front and back types, namely an A2 or weaker subgroup of A with Anti A1 formed in their serum. A1, A2 or weaker subgroups are all classified as Group A and differ with respect to the amount of A antigen carried on the red cells and in secretors with respect to the amount in their saliva. Patients with Anti A1 should receive red blood cells lacking the A1 antigen. |  |  |  |  |  |  |
| Specimen The specimen of choice is fresh EDTA red blood cells or segments from donor units. Refer to Specimen Requirements, Labeling, and Storage Procedure |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |
| Reagents | Supplies |  |  | Equipment |  |  |
|  Anti A1 Lectin stored at 2-8C.  0.8% A1 cells (positive control) and 0.8% A2 cells (negative control) stored at 2-8C.  Saline | - 12 x 75 mm test tubes - Pipet tips - “Antigen Negative” stickers |  |  | - Pipettor - Centrifuge/cell washer - Blood bank mirror |  |  |
| Quality Control Prior to routine testing, reagents must be tested with appropriate controls each day of use. To confirm the specificity and reactivity of the antisera, it is tested with a known antigen positive control cell (A1 cells) and a known antigen negative control cell (A2 cells). These cells must be converted to 4% before use as follows: Step; Action. 1.; Label one 12 X 75 tube for a1 cells and one for a2 cells.. 2.; Place 4 volumes (e.g.200 ul) of each 0.8% cell suspension into the appropriate tube.. 3.; Spin tubes for 30 seconds using Cell Washer/Centrifuge.. 4.; Decant the supernatant completely.. 5.; Add one volume (e.g. 50 ul) of saline to each tube and mix gently to resuspend.. ***Testing of control cells must exhibit expected results for testing to be valid and reportable.*** Procedure Step; Action. 1.; Label 3 tubes – Patient and/or unit, “a1” (positive control), “a2 “(negative control). 2.; Prepare a 3-5% red cell suspension of patient/unit cells visually with saline and add 1 drop to the patient/unit tube.. 3.; Add 1 drop of the prepared a1 reagent red cells to “a1”.. 4.; Add 1 drop of the prepared a2 reagent red cells to “a2”.. 5.; Add 1 drop of Anti A1 Lectin to all tubes.. 6.; Mix the tubes gently and centrifuge for 15 seconds at 3400 rpm.. 7.; While observing in the mirror, resuspend the cell button completely and examine macroscopically for agglutination.. 8.; Interpret the reactions as negative, 1+, 2+, 3+, or 4+ and record the test results as follows:. |
| Step; Prompt; Action. 1.; Order G/RH in LIS; Refer to ABO Type-Forward and Reverse, Rh-Weak D Procedure to perform and result testing.. 2.; If there is a discrepancy between the front and back typing during ABO testing, enter reactions in testing fields for the ABO typing, but do not enter an interpretation.. 3.; With an ABO entry field highlighted; Results  Add Test. 4.; Test Battery ID  Find; Enter A1 (or scroll)  OK. 5.; OK; Click. 6.; SAVE icon; Click and close. Ordering QC Testing for A1 Antigen Step; Prompt; Action. 1.; File  Open  Reagent Rack. 2.; Rack ID field; Enter “2” and Query. 3.; Check reagent lot numbers; Save and close. 4.; File  New  Reagent Test Batch. New Reagent Test Batch window. 5.; Batch ID window; Create a Test Batch #  R(for Reagent), date, tech initials e.g. R030607JZ This number will be recorded at patient testing. 6.; Rack ID field; Enter 2. 7.; Include All Reagents box; Click to uncheck. 8.; OK; Click. Reagent Test Batch window. 9.; Reagent ID field; Enter appropriate reagent from reagent bottle label. 10.; Manufacturer ID field; Enter reagent manufacturer from reagent bottle label. 11.; Lot. No. field; Enter reagent lot # from reagent bottle label. 12.; Exp. Date field; Enter expiration date from reagent bottle from reagent bottle. 13.; Test ID field→ FIND; Click. 14.; Test ID field; Scroll to A1 and highlight. 15.; OK; Click. 16.; OK; Click to save Reagent Test Batch (Do not go to result entry yet). |

To result QC and patient testing only, proceed to Resulting Antigen Testing.

To antigen test units at the same time, continue to Ordering Unit Antigen Testing.

Ordering Unit Antigen Testing (Testing Units for A1)

(This method charges the unit antigen testing to the patient automatically.)

| Step | Prompt | Action |
| --- | --- | --- |
| 1. | Pending Worklog |  |
| 2. | RBC order for patient | Highlight  Click OP on tool bar |
| 3. | Time | Enter “T” for current time |
| Items grid |  |  |
| 4. | Right click and Add a Row |  |
| 5. | Test Type field | Enter T |
| 6. | Test ID field | Enter test code for unit antigen testing to be performed |
| 7. | Quantity | Enter number of units to be tested |
| 8. | SAVE |  |
| 9. | On PW  | Highlight antigen tsting  PS |
| 10. | Scan units to be tested | ACCEPT  OK |

RESULTING TESTING – QC, Patients, Units Step; Prompt; Action. 1.; PW  √ Inc. Work  Query (Red?). 2.; Scroll to the first test batch; Click. 3.; Hold the CTRL key down and find the subsequent product test batch and/or patient; Click. 4.; RE on the tool bar; Click. Result control testing first . 5.; Reaction fields; Enter reaction strengths tab to next field. 6.; Interpretation field Binoculars; Choose correct interpretation → tab to next field. 7.; Comment field; Enter antisera lot #, A1 cells lot #/A2 cells lot #, cell out date e.g. AL195A, A2365/A2445, 041507. Result patient testing. 8.; Verify the specimen; Scan barcode on patient specimen. 9.; Reaction field; Enter reaction strengths tab to next field. 10.; Interpretation window Binoculars; Choose correct interpretation tab to next field. 11.; Comment field; Enter Reagent Test Batch # from Step 2 of Ordering QC Testing for A1 Antigen e.g.R030607JZ. 12.; SAVE icon; Click. If front and back type on GRH testing agree  go to step 13 If front type and back type on GRH testing do not agree (ex. A2 with Anti A1)go to step 25. Front type and back type agree. 13.; Interpretation  Binoculars; Choose correct type. 14.; OK; Click tab to comment field. 15.; Comment field; Enter complete ABO/Rh interpretation (A1/Rh, A2/Rh, A1B/Rh, or A2B/Rh). 16.; Save; Click and close  proceed to enter patient comment. Entering Patient Comment. 17.; File  Open  Patient. 18.; MR# or Pt. ID; Enter informationQuery. 19.; OK; Click. 20.; Edit  Edit Patient Comment. 21.; Description box; Click. 22.; Enter complete interpretation; Enter the ABO/Rh interpretation (A1/Rh, A2/Rh, A1B/Rh, A2B/Rh, or A2/Rh with Anti A1, A2B/Rh with Anti A1). 23.; OK; Click. The Comment “C” field on the Patient-at-a-Glance bar will be highlighted.. 24.; Go to LIS  Result Entry  Comment (on GRH) Enter A1 antigen result. Front and back type do not agree. 25.; Refer to ABO Discrepancies Procedure. Resulting Unit Testing 26.; Repeat Steps 9-12 for units to be tested Enter Reagent Test Batch # in the comment field of the first unit in the batch.. 27.; SAVE icon; Click and close. 28.; When unit testing is complete, place “Antigen Tested” stickers on appropriate units..

Interpretation

Refer to Interpreting Reactions Procedure

Agglutination is a positive test result, indicating the presence of the A1 antigen.

Absence of agglutination is a negative result, indicating the antigen is not demonstrable.

Result Reporting

| Sample Cells | A1 pos ct. cells | A2 neg ct. cells | Interpretation | Result |
| --- | --- | --- | --- | --- |
| 0 | + | 0 | A2 or weaker subgroup | Yes |
| + | + | 0 | A1 | Yes |
| 0 | + | + | Invalid – repeat | No |
| + | 0 | 0 | Invalid – repeat | No |
| + | + | + | Invalid – repeat | No |

Discrepancies between forward and back typing may indicate an A2 or weaker subgroup of A

with an Anti-A1 in their serum with these possible reaction patterns:

| Anti A | Anti B | A1 cells | B cells | Type |
| --- | --- | --- | --- | --- |
| + | 0 | + | + | A2 with Anti A1 |
| + | + | + | 0 | A2B with Anti A1 |

Anti A1 Lectin may be used to resolve these discrepancies

Limitations

- Newborn or infant types are not recommended because the A1 antigen is not fully expressed on red cells of newborn infants and false negatives may occur.

- Determination of A1 type can be difficult if the patient has been transfused within the preceding

three months.

- Interpretations of current testing must be compared with the patient’s past records to detect a possible error or a potentially dangerous situation.

- Aged red blood cells may yield weaker reactions than those obtained with fresh cells.

- Testing with A1 Lectin should never be incubated at 37°C.

| Notes - Frequencies of Blood Types: - 80% of Group A people are A1 - 20% of Group A people are A2 or a weaker subgroup - A1, A2 or weaker subgroups are all classified as Group A and differ with respect to the amount of A antigen carried on the red cells and in secretors with respect to the amount in their saliva. Anti A1 Lectin is a purified extract of Dolichos biflorus seeds and it will only agglutinate red blood cells containing the A1 antigen. Weaker subgroups of A cells or B cells are not agglutinated with Anti A1 Lectin. Both A1 and A1B cells will react positively with Anti A1 Lectin. Patients with the Anti A1 should receive components containing red blood cells lacking the A1 antigen. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Safety Refer to Laboratory Procedure #4002-1: Infection Control, Laboratory Safety Hygiene Procedure and all other safety policies for the lab. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Related Documents ABO Type: Forward and Reverse, Rh(Weak D) ABO Discrepancies |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| References Ortho Anti A1 Lectin, Ortho package insert, May, 1999. AABB Technical Manual, 15th Edition, 2005, pages 293-294. Spectra Biologicals Applied Blood Group Serology, Second Edition, July 1977, pages 79-82. Standards for Blood Banks and Transfusion Services, 23rd Edition, 2005, page 38. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Author: |  |  |  |  |  | Date: |  | Validated by: |  |  | Date: |  | Approved by: |  |  |  |
| 5/06/08 1-15-07 |  | Cheryl Webster, MT(ASCP) Jo Lohmann, MT |  |  |  |  |  | 6/17/08 |  | Jo Lohmann, MT(ASCP) |  |  | 6/18/2008 |  | Dr. T Hallin |  |  |  |
| L: Blood Bank: A1 Antigen Type (Lectin) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |