| Ridgeview Medical Center Laboratory Services |  |  | Blood Bank | Effective Date:12/08 |
| --- | --- | --- | --- | --- |
| Procedure This procedure provides abbreviated directions for antigen testing and detailed instructions for entering results of antigen testing on patients, units, and QC into HBB. Procedure Notes - Before selecting units for testing, check Blood and Blood Component Packing List for historical antigen typing. If unit is available in bank, proceed to confirm antigen typing (Historical typing cannot be accepted as the type of record) - If there is a unit in the bank with an Antigen Negative tag from the supplier, the unit does not need to be retyped. - If necessary, antigen negative or historically typed units can be ordered per Ordering Products from Supplier - Positive control should be heterozygous - Perform DAT on all patients in conjunction with antigen testing - Refer to specific Antigen Testing procedures for cell suspension percentages and diluents K,Fya; IgG card50 µL cells (0.8%) + 25 µL antiseraIncubate 15 minCentrifugeRead. C,E,c,e; Appropriate antigen testing card & control card 4% in MTS Diluent2Plus10 µLCentrifuge. Jka; 1 drop cells + 1 drop antisera Incubate RT 5-30 min Spin Read. M; Wash cells once2 drops 2% cells + 1 drop antiseraIncubate 30 min Do Not SpinRead. These steps allow all testing to be resulted on the same Result Entry Screen |
| Step | Prompt | Action |  |  |
| 1. | Patient/Order module |  |  |  |
| ORDERING PATIENT ANTIGEN TESTING |  |  |  |  |
| 2. | File Open Patient |  |  |  |
| 3. | MR# or Pt. ID | Enter Information |  |  |
| 4. | Query | Click OK |  |  |
| 5. | File NewOrder  Regular | OK |  |  |
| 6. | Priority Field Choose priority | Click OK |  |  |
| 7. | Time field | Enter time specimen was collected |  |  |
| 8. | Test Type Field | Enter T |  |  |
| 9. | Item ID field → Binoculars | Scroll to RCA |  |  |
| 10. | OK | Click |  |  |
| 11. | Tab to Quantity field | Enter 1 |  |  |
| 12. | SAVE icon | Click and close |  |  |

ORDERING ANTIGEN QC TESTING

| Step | Prompt | Action |
| --- | --- | --- |
| 1. | File  Open  Reagent Rack |  |
| 2. | Rack ID field | Enter “2” and Query |
| 3. | Check reagent lot numbers | Save and close |
| 4. | File  New  Reagent Test Batch |  |
| New Reagent Test Batch Window |  |  |
| 5. | Batch ID window | Create a Test Batch #  R (for Reagent), date, tech initials e.g. R030607JZ This number will be recorded at patient and unit testing. |
| 6. | Rack ID field | Enter 2 |
| 7. | Include all Reagents box | Click to uncheck |
| 8. | OK | Click |
| Reagent Test Batch Window |  |  |
| 9. | Reagent ID Field | Enter appropriate reagent |
| 10. | Manufacturer ID field | Enter reagent manufacturer |
| 11. | Lot. No. field | Enter reagent lot # |
| 12. | Exp. Date field | Enter expiration date from reagent bottle |
| 13. | Test ID field → FIND | Click |
| 14. | Test ID field | Scroll to appropriate test ID and highlightOK |
| 15. | OK | Click to save Reagent Test Batch (Do not go to result entry yet). |

To result QC and patient testing only, proceed to Resulting Antigen Testing.

To antigen test units at the same time, continue to Ordering Unit Antigen Testing.

ORDERING UNIT ANTIGEN TESTING

See Procedure Notes before selecting units

(This method charges the unit antigen testing to the patient automatically)

| Step | Prompt | Action |
| --- | --- | --- |
| 1. | Pending Worklog |  |
| 2. | RBC order for patient | Highlight  Click OP on tool bar |
| 3. | Time | Enter “T” for current time |
|  | Items grid |  |
| 4. | Right click and Add a Row |  |
| 5. | Test Type field | Enter T |
| 6. | Test ID field | Enter test code for unit antigen testing to be performed |
| 7. | Quantity | Enter number of units to be tested |
| 8. | SAVE |  |
| 9. | On PW | Highlight antigen testing  PS |
| 10. | Scan units to be tested | ACCEPT OK |

RESULTING ANTIGEN TESTING – QC, PATIENTS, UNITS

| Step | Prompt | Action |
| --- | --- | --- |
| 1. | PW √ Inc. WorkQuery (Red?) |  |
| 2. | Scroll to the first test batch | Click |
| 3. | Hold the CTRL key down and find the subsequent batch, patient, and/or unit. | Click |
| 4. | RE on the tool bar | Click |
| Result CONTROL testing first  (e.g. “Anti e”) |  |  |
| 5. | Reaction fields | Enter reaction strengths  tab to next field |
| 6. | Interpretation field Binoculars | Choose correct interpretation  tab to next field |
| 7. | Comment field | Enter antisera lot #, +/- cell numbers, cell lot#, cell out date – e.g. KC143N, +5/-6, 8RA199, 041507 |
| Resulting PATIENT testing  |  |  |
| 8. | Verify Specimen ID | Scan Accession # on tube |
| 9. | Test field Binoculars | Scroll to appropriate testing and select |
| 10. | OK | Click |
| 11. | Reaction field | Enter reaction strengths  tab to next field |
| 12. | Interpretation field  Binoculars | Choose correct interpretation  tab to next field |
| 13. | Comment field | Enter Reagent Test Batch # from Step 2 of Ordering Antigen QC Testing e.g. R030607JZ |
|  | Resulting UNIT testing  |  |
| 14. | Repeat Steps 11-12 for units to be tested |  |
| 15. | Comment field | For the first unit tested: enter lot number of the antisera or card used in testing |
| 15. | SAVE icon | Click and close |

| Date: | Author: |  |  |  |  |  |  | Date: |  |  | Validated by: |  |  |  | Date: |  | Approved by: |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4/20/07 | Cheryl Webster, MT(ASCP) |  |  |  |  |  |  | 4/9/08 |  | Cheryl Webster, MT(ASCP) Jo Lohmann, MT(ASCP) |  |  |  |  | 6/18/2008 |  | Dr. T. Hallin |  |  |  |
| Antigen Testing – Quick Guide |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 12/22/10 |  | CW |  |  | Changed C ,c, E, e testing to gel method |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3/10/11 |  | CW |  |  | Changed antigen testing of units to bill the patient automatically |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |