| Ridgeview Medical Center Laboratory Services |  |  | Blood Bank | Effective Date: 9/26/08 |  |
| --- | --- | --- | --- | --- | --- |
| Purpose This procedure provides instructions for the preparation of cell suspensions for gel and tube testing. Suspensions. Test; Gel Testing Concentration; Diluent; Tube Testing Concentrations; Diluent. ABO, Rh; 4%; MTS Diluent2 PLUS; 4%; Saline. Crossmatch-IS; 0.8%; MTS Diluent2 PLUS; 4%; Saline. Crossmatch-IgG; 0.8%; MTS Diluent2; ---; ---. DAT; 0.8%; MTS Diluent 2; 3-4%; Saline. Auto Control; 0.8%; MTS Diluent 2; ---; ---. C, c, E, e; 4%; MTS Diluent 2 PLUS; 4%; Saline. FMBS; ---; ---; 2-3%; Saline. Conversions. Panel Cells (0.8%); Must be converted to 3-4% for use in tube methods for antigen testing, or albumin, LISS, and prewarm procedures. Screening Cells (0.8%). Coombs Control Cells (3%); Must be converted to 0.8% for use as positive control for DAT testing.. Complement Control Cells (4%). |
| Materials |  |  |  |  |  |
| Supplies |  | Equipment |  |  |  |
| - Saline - MTS Diluent 2 - MTS Diluent 2 Plus - 12 x 75 tubes - Disposable glass pipettes |  | - Cell washer/centrifuge - BioHit Pipettor & tips - Eppendorf Pipettor & tips |  |  |  |
| Preparing a 0.8% Cell Suspension from packed cells Step; Ortho BioHit pipettor  P6. 1.; Dispense 1.0 ml diluent into a labeled tube using the Eppendorf pipettor.. 2.; Aspirate 100ul of diluent from this tube.. 3.; Press Start again to create an air gap. 4.; Place tip in packed cells, and press Start to aspirate 10 µL of sample.. 5.; Place tip into tube of diluent and press Start to purge contents and mix automatically.. Preparing a 3-4% Cell Suspension from packed cells Step; Ortho BioHit pipettor  P7. 1.; Aspirate 200ul of appropriate diluent.. 2.; Press Start to create an air gap.. 3.; Place tip in packed cells and press Start to aspirate 10 µL of sample.. 4.; Place tip into a labeled tube and press Start to purge contents and mix automatically.. Preparing a 3-4% Cell Suspension from whole blood Step; Ortho BioHit pipettor  P8. 1.; Aspirate 200 µL of appropriate diluent.. 2.; Press Start to create an air gap.. 3.; Place tip in whole blood and press Start to aspirate 20 µL of sample.. 4.; Place tip into a labeled tube and press Start to purge contents and mix automatically.. Preparing a 2% Cell Suspension from whole blood Step; Action. 1.; Aspirate 200 µL of appropriate diluent.. 2.; Press Start to create an air gap.. 3.; Place tip in whole blood and press Start to aspirate 20 µL of sample.. 4.; Place tip into a labeled tube and press Start to purge contents and mix automatically.. 5.; Using P1, aspirate 50 µL of this suspension and and place in an appropriately labeled tube.. 6.; Using P1, aspirate 50 µL of appropriate diluent and add to tube. Mix gently.. Converting 0.8% to 4% |
| Step | Action |  |  |  |  |
| 1. | Place 4 volumes (e.g.200 ul) of 0.8% cell suspension in a labeled test tube. |  |  |  |  |
| 2. | Using Cell Washer/Centrifuge, spin tube for 30 seconds. |  |  |  |  |
| 3. | Decant the supernatant completely. |  |  |  |  |
| 4. | Add one volume (e.g. 50 ul) of appropriate diluent to the tube and resuspend. |  |  |  |  |

| Converting 3-4% to 0.8% |  |
| --- | --- |
| Step | Action |
| 1. | Place 1 volume (e.g.50ul) of a 4% cell suspension in a labeled test tube. |
| 2. | Add 3 volumes (e.g.150ul) of appropriate diluent and mix gently. |

3-4% cell suspensions in saline used in tube typing may be made using the following visual method.

| Step | Action |
| --- | --- |
| 1. | Place approximately 1 ml of saline in an appropriately labeled tube. |
| 2. | Add approximately 40 µL of packed cells with a disposable glass pipette. |
| 3. | Empty and fill pipette to rinse and mix the suspension. |

References

Conversion of 0.8% to 4%: Procedure 8225A

Ortho BioHit Pipettor Instruction Card

Related Documents

Ortho BioHit Pipettor Instruction Card

Safety

Refer to Laboratory Procedure #4002-1:  Infection Control and Laboratory Safety Hygiene Procedure

| Date: |  | Author: |  |  |  |  |  | Date: |  | Validated by: |  |  | Date: |  | Approved by: |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 5/16/07 |  | Cheryl Webster, MT(ASCP) Jo Lohmann, MT(ASCP) |  |  |  |  |  | 6/17/08 |  | Jo Lohmann, MT(ASCP) |  |  | 6/18/08 |  | Dr. T. Hallin |  |  |  |
| Cell Suspensions and Conversions |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 12/2/10 | CW |  |  |  | Addition of directions for suspensions used in gel testing. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2/14/12 | CW |  |  |  | Addition of Complement Control Cells |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |