| Ridgeview Medical Center Laboratory Services |  |  |  |  | Blood Bank |  |  |  | Effective Date: 12/15/2010 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose This procedure provides instruction for routine crossmatch of red blood cells for transfusion. Including: - Electronic - Immediate Spin- MTS Buffered Gel - MTS (IgG) Gel |  |  |  |  |  |  |  |  |  |  |
| Specimen The specimen of choice is plasma from EDTA anticoagulated whole blood. Refer to Specimen Requirements, Labeling, and Storage Procedure The pretransfusion sample used for testing must be not more than three days old at the time of the intended transfusion if the patient has been pregnant or transfused within the previous three months, or if the pregnancy history and transfusion history are uncertain. The day of collection is counted as day zero. |  |  |  |  |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |  |  |  |
| Reagents |  |  |  | Supplies |  |  | Equipment |  |  |  |
| - MTSTM IgG Gel Card, stored at 2-25°C. - MTSTM Buffered Gel Card, stored at 2-25°C. - MTSTM Diluent 2, stored at 2-8°C. - MTSTM Diluent 2 PLUS, stored at 2-8°C |  |  |  | - Pipet tips - 12 x 75 test tubes |  |  | - MTS incubator - MTS centrifuge - MTS pipettor |  |  |  |
| Quality Control Prior to routine testing, reagents must be tested with appropriate control each day of use. Refer to the Quality Control Procedure. Controls must exhibit expected results for testing to be valid before reporting any patient results. |  |  |  |  |  |  |  |  |  |  |
| General Crossmatch Criteria Before performing Electronic, Immediate Spin (Buffered Gel) or MTS (IgG) Gel crossmatches, the following steps must be followed. Step; Action. 1.; Check patient history for current specimen and antibody screen (within 3 days). File Open Patient. 2.; If there is no current specimen and/or antibody screen, perform testing according to Indirect Antiglobulin Test procedure and ABO-Forward and Reverse, Rh – Gel procedure.. 3.; Result all testing in computer using Resulting Testing from the Pending Worklog procedure prior to unit selection.. 4.; All antibody identification and antigen typing for units and patient should be completed before crossmatching units for patient.. 5.; Add-on RBC orders require a RECORD and RECKG, added as follows:. |
| Step | Prompt |  |  |  |  | Action |  |  |  |  |
| 1. | Pending Work log |  |  |  |  |  |  |  |  |  |
| 2. | RBC order for the patient you are crossmatching |  |  |  |  | Highlight to select |  |  |  |  |
| 3. | OP on the tool bar |  |  |  |  | Click |  |  |  |  |
|  | Order Profile window |  |  |  |  |  |  |  |  |  |
| 4. | Time field |  |  |  |  | Enter time |  |  |  |  |
| 5. | Test Type field |  |  |  |  | Enter T |  |  |  |  |
| 6. | Item ID field |  |  |  |  | Enter RECKG tab |  |  |  |  |
| 7. | Quantity field |  |  |  |  | Enter 1  tab |  |  |  |  |
| 8. | Repeat Steps 5-7 for RECORD |  |  |  |  | SAVE and close window |  |  |  |  |
|  | Pending Work log |  |  |  |  |  |  |  |  |  |
| 9. | RECORD RECKG |  | Perform and result according to Resulting Testing from the Pending Worklog procedure |  |  |  |  |  |  |  |
| ELECTRONIC Crossmatch Criteria - Use of a validated Blood Bank Information System that prevents selection of ABO incompatible red blood cells. - No history of a clinically significant antibody or a positive antibody screen - Two different specimens and two different visits - Two ABO/RH types on separate specimens. One must be a current specimen - A type and negative antibody screen on a current specimen. |  |  |  |  |  |  |  |  |  |  |
|  | Pending Work log √ ”Show Electronic Crossmatch” Q=Qualifies A=Almost qualifies(pending current testing) N=Does Not Qualify |  |  |  |  |  |  |  |  |  |
| 1. | RBC order for the patient you are crossmatching |  | Select  PS on tool bar |  |  |  |  |  |  |  |
|  | Product Selection window |  |  |  |  |  |  |  |  |  |
| 2. | Select units for crossmatch, preferably type specific. See Leukoreduced Red Blood Cells and Product Selection in HBB procedures. Remove appropriate units from the Blood Bank. |  |  |  |  |  |  |  |  |  |
| 3. | Bar code field |  | Scan product code and unit number |  |  |  |  |  |  |  |
| 4. | Accept |  | Click |  |  |  |  |  |  |  |
|  | If patient is eligible for an Electronic Crossmatch the following message will display: |  |  |  |  |  |  |  |  |  |
| EXM has been run with COMPATIBLE interpretation for (product code/unit number). Post results? |  |  |  |  |  |  |  |  |  |  |
| 5. | YES | Click |  |  |  |  |  |  |  |  |
|  | Scan number of units ordered by following steps 3-5 above. |  |  |  |  |  |  |  |  |  |
| 6. | OK | Click |  |  |  |  |  |  |  |  |
|  | Unit tags containing recipient information print. XM on PW will automatically cancel. |  |  |  |  |  |  |  |  |  |
| 7. | Place stickers from unit tags on units, attach tag, and return unit to blood bank refrigerator. |  |  |  |  |  |  |  |  |  |

immediate Spin Crossmatch Criteria

- No previous history of clinically significant antibody

- ABO/RH and negative antibody screen on current specimen

| Step | Prompt | Action |
| --- | --- | --- |
|  | Pending Worklog |  |
| 1. | RBC for the patient you are crossmatching | Select PS (product select)on the tool bar |
|  | Product Selection window |  |
| 2. | Select units for crossmatch, preferably type specific. See Leukoreduced Red Blood Cells and Product Selection in HBB procedures. |  |
| 3. | Remove correct number of ordered type specific units from the Blood Bank refrigerator. |  |
| 4. | Bar Code field | Scan product code and unit number |
| 5. | Accept | Click |
|  | Repeat Steps 4-5 for number of units ordered. |  |
| 6. | OK | Click |
| Procedure Notes - Do not remove the foil seal until ready to use or within 1 hour of testing. - Inspect each gel card for a clear liquid layer on top of the opaque gel. - Do not touch the gel card during pipetting to prevent carryover leading to erroneous results. Procedure: |  |  |
| 1. | Bring samples and reagents to room temperature. |  |
| 2. | Using a Buffered Gel Card, label microtube with patient name and unit number and remove the foil seal. |  |
| 3. | Label 12 x 75 tubes with unit numbers. Remove one segment and place in appropriate tube. Return units to the blood bank refrigerator. |  |
| 4. | Using a segment device, place one or two drops into the appropriately labeled tube. |  |
| 5. | Using BioHit Pipettor P6, make a 0.8% solution of the red cells to be tested in MTSTM Diluent 2 PLUS. |  |
| 6. | Add 50 µL of this cell suspension to the appropriately labeled microtube. |  |
| 7. | Add 50 µL of patient plasma to the same microtube. |  |
| 8. | Centrifuge for 10 minutes. |  |
| 9. | Read reactions immediately following centrifugation. See “Interpretation” section of this procedure. |  |
|  | Pending Worklog |  |
| 10. | XM for same patient | Highlight to select  RE (result entry) |
| 11. | Specimen ID field | Scan accession # from patient’s current specimen. |
| 12. | Crossmatch field | Enter XMISG or click on binoculars to select correct method. |
| 13. | Reaction field | Record reactions as 0, 1+, 2+, 3+, 4+ |
| 14. | Interpretation field | Enter COMP or click on binoculars for appropriate result. |
| 13. | SAVE | Click and close. |
| 14. | Unit tags containing recipient information print. |  |
| 15. | Place stickers from unit tags on units, attach tag, and return unit to blood bank refrigerator. |  |

Interpretation:

Observe the front and back of each microtube for agglutination and/or hemolysis.

Refer to Interpreting Reactions procedure.

- Compatible: Complete sedimentation of all red blood cells in the bottom of the microtube and the absence of hemolysis.

- Incompatible: Red blood cells remain suspended on top of the gel or are dispersed throughout the gel and/or hemolysis.  Some positive reactions can have a small button of red blood cells remaining in the bottom of the microtube

Procedure Notes

- Electronic and Immediate Spin crossmatch only detects ABO incompatibility and should not be performed on patients with a history of, or a current, clinically significant antibody.

- Components can be added on to a current Type and Screen within 3 days. Use of an Electronic or Immediate Spin Crossmatch may be used where appropriate.

- Immediate Spin testing is performed at room temperature.  Antibodies that react at room temp may be detected ex. Anti-M, N and P.

- A unit tag will print when resulting an incompatible unit; destroy this unit tag.

Limitations

- False positive or false negative results may occur from bacterial or chemical contamination of test materials, aged blood specimens, inadequate incubation time or temperature, improper centrifugation, improper storage of materials or omission of test samples.

- False positive results may occur if a card is used that shows signs of drying.

- Antibodies to preservatives, medications, disease states, Wharton’s jelly, and/or cross contamination of reaction microtubes may cause false positive reactions.

- False positives or hazy reactions may occur in patients with rouleaux caused by high plasma concentrations of protein. Samples exhibiting rouleaux should be washed several times with warm saline and retested.

- Anomalous results may be caused by fibrin or other particulate matter that may stick to the sides of the microtube.

MTS IgG GEL Crossmatch Criteria

- Patient has a  clinically significant antibody

- Patient has a history of a clinically significant antibody.

| Step | Prompt | Action |  |
| --- | --- | --- | --- |
|  | Pending Worklog |  |  |
| 1. | RBC for the patient you are crossmatching | SelectPS(product select) on the tool bar |  |
|  | Product Selection window |  |  |
| 2. | Select units for crossmatch, preferably type specific. See Leukoreduced Red Blood Cells and Product Selection in HBB procedures. Select units that are antigen negative for the patient’s antibody, if necessary. |  |  |
| 3. | Remove the correct number of type specific units from the Blood Bank refrigerator. |  |  |
| 4. | Bar Code field | Scan product code and unit number |  |
| 5. | Accept | Click |  |
| 6. | Repeat steps 4 and 5 for number of units ordered. |  |  |
| 7. | OK | Click |  |
| 8. | Label test tubes with unit numbers and remove segment from each unit and place in appropriate tube. |  |  |
| 9. | Return units to Blood Bank refrigerator. |  |  |
| 10. | Using a Hematype Segment Device pierce segment, drip two to three drops of red cells into appropriately labeled tube. |  |  |
| 11. | Label 2 tubes for each unit to be crossmatched: IS and IgG Prepare 2 different 0.8% cell suspensions from these segments as follows: “IS” For Immediate Spin Crossmatch; “IgG” For IgG Crossmatch. Using BioHit Pipettor P6, make a 0.8% solution of the red cells to be tested in MTSTM Diluent 2 PLUS; Using BioHit Pipettor P6, make a 0.8% solution of the red cells to be tested in MTSTM Diluent 2. |
|  | Procedure Notes - Do not remove the foil seal until ready to use or within 1 hour of testing. - Inspect each gel card for a clear liquid layer on top of the opaque gel. - Do not touch the gel card during pipetting to prevent carryover leading to erroneous results. |  |  |
| 12. | Using an Anti-IgG Gel card, label microtube with patient name and unit number and remove the foil seal. Using a Buffered Gel Card, label microtube with patient name and unit number and remove the foil seal. |  |  |
| 13. | Pipette 50 µL of the “IgG” cell suspension into the appropriate microtube of the Ig G card. Add 50 µL of the “IS” cell suspension to the appropriately labeled microtube of the Buffered Gel card. (Cells must be pipetted first to allow for proper mixing with addition of plasma) |  |  |
| 14. | Add 25 µL of patient plasma to the microtube of the IgG card. Add 50 µL of patient plasma to the microtube of the Buffered Gel card. |  |  |
| 15. | Incubate the IgG card at 37°C for 15 minutes (incubation may be extended to 30 minutes). Centrifuge the Buffered Gel card for 10 minutes. Read reactions immediately following centrifugation. Go to Step 17. |  |  |
| 16. | After 15 minute incubation of IgG card, centrifuge for 10 minutes. Read reactions immediately following centrifugation. Go to Step 17. |  |  |
| 17. | Read macroscopically the front and back of each microtube. Referring to Interpreting Reactions procedure, read the reactions as 0, 1+, 2+, 3+ or 4+ and record results as follows. |  |  |
|  | Pending Worklog |  |  |
| 18. | XM for the patient being crossmatched |  | Select  RE |
|  | Result Entry window |  |  |
| 19. | Specimen ID field |  | Scan accession # from patient’s current specimen. |
| 20. | Crossmatch field |  | Enter XMG or use binoculars to select the correct method |
| 21. | Result field |  | Record reaction strengths for both IS and IgG methods. |
| 22. | Interpretation field |  | Enter COMP or use binoculars to select appropriate result. |
| 23. | SAVE |  | Click and close |
| 24. | Unit tags containing recipient information print. |  |  |
| 25. | Place stickers from unit tags on units, attach tag, and return unit to blood bank refrigerator. |  |  |
| 26. | Discard the card if all wells are used, or place the card with foil covered wells into rack holder to be used later. |  |  |

Interpretation of Results

Refer to Interpreting Reactions procedure

- Compatible: Complete sedimentation of all red blood cells in the bottom of the microtube

- Incompatible: Red blood cells that remain suspended on the top of the gel or are dispersed throughout the gel. Some positive reactions can have a small button of red blood cells remaining in the bottom of the microtube.  See Interpreting Reactions.

- Incompatible units will not be issued and will be followed up with additional testing to resolve the problem if necessary. Additional units should be crossmatched to determine compatibility.

- Use Alternate Methods- Prewarm, Albumin, LISS, if necessary.

Limitations

- Use of MTS Gel may not detect antibodies that bind complement. These antibodies can be clinically significant, especially Jka and Jkb.

- MTS Gel may not detect weak antibodies to Kell and E.

- A unit tag may print when resulting an incompatible unit; destroy this unit tag.

- Clots, particulates, platelets or other artifacts may cause some red blood cells to be entrapped at the top of the gel and may cause a false positive result.

- False positives may occur if antibodies to components of the gel are present in the sample tested.

- Donor units with positive DATs will be positive using gel methodology.

| Least Incompatible Crossmatches This procedure provides directions for the performance and interpretation of crossmatches where a compatible result will not occur using appropriate testing methods, due to the presence of clinically insignificant antibodies to antigens present on donor units. Patient antibody identification should be completed at the reference laboratory and their interpretation should indicate the proper procedure for the selection of units for compatibility testing. If that recommendation indicates using Least Incompatible units, the following procedure should be used. Step; Action. 1.; Open the patient and order the number of RBCs that will be crossmatched as follows: - A and O patients – order and select 8 units - B and AB patients – order and select entire available supply. 2.; Pending Worklog  RBC  PS (on the tool bar). 3.; Barcode field; Scan product code and unit number. 4.; Accept; Click. Repeat Steps 3 and 4 for all units to be crossmatched. 5.; Perform testing using the method recommended by the reference laboratory (i.e. Gel, LISS, albumin) See appropriate testing method procedures.. 6.; Pending Worklog  XM  RE(on the tool bar). 7.; Specimen ID Field; Scan accession # from patient’s specimen. 8.; Crossmatch field; Select correct testing method. 9.; Result field; Record reaction strengths. 10.; Comparing reaction strengths of all units tested, choose the weakest reactions and result those as Least Incompatible. - If all reactions are greater than 2+, select more units and repeat the crossmatching - If necessary, order more units from the blood center to find least incompatible units. 11.; Once all units have been tested and evaluated, result the most incompatible units as Incompatible.. 12.; Unit tags will print for all testing resulted  - Pull Incompatible tags and discard.. 13.; Place labels from unit tags on units, attach tags, and return units to blood bank refrigerator.. 14.; Units that were resulted as Incompatible and those that are not tagged for the patient must be returned to available inventory for general use. Proceed as follows: 1.; Transfusion: Inventory module Components (on tool bar) Release to Available. Release to Available window. 2.; Barcode field; Scan unit number and SPC. 3.; Accept; Click to move unit to Components to Release field. 4.; Repeat Steps 2 & 3 for all units. 5.; OK; Click and close. Units are now available for other patients... Note: Least Incompatible units will require Supervisory Override at Issue.. Safety Refer to Laboratory Procedure #4002-1: Infection Control and Laboratory safety Hygiene Procedure and all safety policies of the laboratory. |
| --- |
| Related Documents Alternate Methods-Prewarm, Albumin, LISS ABO-Forward and Reverse, Rh Gel Card Cell Suspensions and Conversions Crossmatch-Routine Indirect Antiglobulin Test Interpreting Reactions Quality Control Resulting Testing from the Pending Worklog Specimen Requirements, Labeling, and Storage |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| References AABB Technical Manual, 15th Edition, 2005 Standards for Blood Banks and Transfusion Services, 25th Edition, 2008, pages 8,41-54. Ortho Clinical diagnostics, Inc. Anti-Human Globulin Anti-IgG (Rabbit) MTS Anti-IgG Card package insert, December 2004. Ortho MTSTM Buffered Gel Card, Technical Reference Document RMC Procedure: Crossmatch-Routine |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Author: |  |  |  |  |  | Date: |  | Validated by: |  |  |  | Date: |  | Approved by: |  |  |  |
| 09/07/08 11/21/10 |  | Mary Farlow,MT(ASCP) SBB Jo Lohmann, MT(ASCP) Cheryl Webster, MLSCM(ASCP) |  |  |  |  |  | 12/1/10 |  | Jo Lohmann, MT(ASCP) |  |  |  | 12/9/10 |  | R.P.Weatherby, MD |  |  |  |
| L: Blood Bank: Crossmatch-Electronic, Immediate Spin(Gel),IgG(Gel) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Initials: |  |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 11/21/10 | CW |  |  |  | Conversion from IS tube to IS gel & slight revision of computer steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2/18/11 | CW |  |  |  | Addition of Least Incompatible Crossmatch |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |