| Ridgeview Medical Center Laboratory Services |  |  | Blood Bank |  |  | Effective Date: |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Principle ABO forward grouping and Rh (D) grouping determine the presence or absence of antigens. Red cells are tested with Anti-A, Anti-B, and Anti-D including Weak-D (tube) reagents which are incorporated into the MTS gel for Infant Typing. After centrifugation, microtubes are examined for agglutination. |  |  |  |  |  |  |  |
| Specimen Fresh red blood cells from EDTA whole blood, centrifuged and separated according to Specimen Requirements, Labeling and Storage procedure; cord blood (EDTA whole blood and clot tube). |  |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |
| Reagents |  | Supplies |  |  | Equipment |  |  |
| - MTSTM A/B/D Monoclonal Card, stored at 2-25°C. - MTSTM Diluent 2 PLUS, stored at 2-8°C. - Anti-D reagent, stored at 2-8˚ C - Anti-IgG reagent, stored at 2-8˚ C. - 6-8% Albumin, prepared daily, stored at 2-8˚C. - Coombs Control Cells stored at 2-8˚C. - Saline |  | - Glass test tubes – 12x75 - Pipette tips |  |  | - MTS BioHit Pipettor - MTS Incubator - MTS Centrifuge - Blood bank mirror |  |  |
| Quality Control Prior to routine testing, reagents must be tested with appropriate controls each day of use. Refer to Quality Control Procedure. Controls must exhibit expected results for testing to be valid before reporting any patient results. Procedure Notes - Do not remove the foil seal until ready to use or within 1 hour of testing. - Inspect each gel card for a clear liquid layer on top of the opaque gel. - Do not touch the gel card during pipetting to prevent carryover leading to erroneous results. |  |  |  |  |  |  |  |
| Procedure: |  |  |  |  |  |  |  |
| Step | Action |  |  |  |  |  |  |
|  | Bring samples and reagents to room temperature. |  |  |  |  |  |  |
|  | Label appropriate cells with a patient sticker and remove the foil seal. |  |  |  |  |  |  |
|  | Label a 12 x 75 glass tube. Using BioHit Pipettor program P7, make a 4% solution of the red cells to be tested in MTSTM Diluent 2 PLUS. If using whole blood, use BioHit program P8. |  |  |  |  |  |  |
|  | Using BioHit Pipettor program P9, add 10 µL of this solution to the Anti-A, Anti-B, and Anti-D microtubes. |  |  |  |  |  |  |
|  | Centrifuge the gel card for 10 minutes. |  |  |  |  |  |  |
|  | Read reactions immediately following centrifugation. |  |  |  |  |  |  |
|  | If infant has not been previously typed a recheck (RECKI) must be performed. Repeat steps 1 – 6 using a new cell suspension. |  |  |  |  |  |  |
| If all microtubes are positive (AB Pos) a control must be performed as follows: |  |  |  |  |  |  |  |
| Step | Action |  |  |  |  |  |  |
| 1. | Using the Monoclonal Control Card, label one microtube with patient name. |  |  |  |  |  |  |
| 2. | Using BioHit Pipettor program P9: Pipette 10µL of the patient’s 4% cell suspension into the appropriately labeled microtube. |  |  |  |  |  |  |
| 3. | Centrifuge the card for 10 minutes. |  |  |  |  |  |  |
| 4. | Control must be negative for the testing to be valid. |  |  |  |  |  |  |
| All Rh Negative infants need to be tested for WeakD Weak D-Tube Method testing is as follows |  |  |  |  |  |  |  |
| 1. | Results Add Test  Battery: WeakD  OK |  |  |  |  |  |  |
| 2. | Label a tube for “D” and one for “Ct” |  |  |  |  |  |  |
| 3. | Add one drop Anti-D antisera and one drop from infants 4% cell suspension in tube labeled D |  |  |  |  |  |  |
| 4. | Add one drop 6-8% albumin and one drop from infants 4% cell suspension in tube labeled Ct |  |  |  |  |  |  |
| 5. | Mix tubes gently and centrifuge for 15 seconds at 3400 rpm |  |  |  |  |  |  |
| 6. | While observing in the mirror, completely resuspend the cell button and examine macroscopically for agglutination. If the control is positive, repeat testing. |  |  |  |  |  |  |
| 7. | Incubate both tubes at 37˚C for 15-60 minutes. |  |  |  |  |  |  |
| 8. | Mix the tubes gently and centrifuge for 15 seconds at 3400 rpm |  |  |  |  |  |  |
| 9. | While observing in the mirror, completely resuspend the cell button and examine macroscopically for agglutination. * Control must be negative for testing to be valid. Resolve all discrepancies. |  |  |  |  |  |  |
| 10. | If reaction is positive after the 37˚C incubation, the patient is Rh positive. Result WeakD as INV and IAT/CC as NT. Immediate spin reading on initial testing will be resulted as “1”. Patient Comment will be entered “Rh positive after incubation at 37”. |  |  |  |  |  |  |
| 11. | If reaction is negative after 37˚ C incubation, wash tubes at least 3 times. |  |  |  |  |  |  |
| 12. | Add 2 drops Anti-IgG to both tubes. |  |  |  |  |  |  |
| 13. | Mix gently and centrifuge for 15 seconds at 3400 rpm. |  |  |  |  |  |  |
| 14. | While observing in the mirror, completely resuspend the cell buttons and examine macroscopically for agglutination. |  |  |  |  |  |  |
| 15. | Interpret the reactions as negative , weak + , 1+, 2+, 3+, 4+, and record the test results per Resulting Testing from the Pending Worklog. |  |  |  |  |  |  |
| 16. | Add one drop of Coombs Control Cells to all negative reacting tubes. |  |  |  |  |  |  |
| 17. | Mix gently and centrifuge for 15 seconds at 3400 rpm. |  |  |  |  |  |  |
| 18. | While observing in the mirror, completely resuspend the cell buttons and examine macroscopically for agglutination. |  |  |  |  |  |  |
| 19. | Interpret the reactions as negative, W+, 1+, 2+, 3+, 4+, and record the test results per Resulting Testing from the Pending Worklog. - A reaction less than 2+ is an invalid result and testing must be repeated. |  |  |  |  |  |  |

| Interpretation Observe each microtube for agglutination and/or hemolysis. Refer to Interpreting Reactions procedure. - Negative –No agglutination, no hemolysis, and complete sedimentation of all cells to the bottom of the microtube. - Positive - Red blood cells suspended on the top of the gel or dispersed throughout the gel in varying degrees, or hemolysis. Some positive reactions can have a small button of red blood cells remaining in the bottom of the microtube. A very weak reaction on one or both sides of the microtube is not an expected result. This may indicate that a false positive or a very weak/partial expression of the antigen is present. Further investigation of this cell should be performed before the ABO and Rh status is determined. Rh Negative Infant with a Positive DAT Positive DAT  WeakD testing is invalid Result as follows: 1.; Result Invalid (INV) in the interpretation field of WeakD Result Invalid (INV) in the interpretation field of CORD.. 2.; Add the following comment in the GRHI Comment field: “Infant Type (A, O, B or AB), Rh Unknown due to positive DAT”. See Example 1.. 3.; Save. 4.; In LIS: Result Group & Rh-Infant as “No Type”. 5.; Enter following under test comments: See Example 2 - Infant Type (enter group A,O,B or AB) / Rh Unknown due to positive DAT - Return in three months for retest - Rh-negative mothers should have Rh-Immune Globulin administered. 6.; Save. |
| --- |
| Interpretation. Anti-A microtube; Anti-B microtube; Anti-D microtube; Group and Rh. 0; 0; +; O Positive. 0; 0; 0; O Negative. +; 0; +; A Positive. +; 0; 0; A Negative. 0; +; +; B Positive. 0; +; 0; B Negative. +; +; 0; AB Negative. +; +; +; AB Positive. Result Reporting Record results according to Resulting Testing from the Pending Worklog Procedure Crossmatch patients: testing ordered will be INFANT Cord Blood Workups, GRHI: testing ordered will be GRHI Rechecks: RECKI |
| Restrictions/Limitations - Current test results must be compared to the patient’s past record to detect possible errors. - False positive or false negative results may occur from bacterial or chemical contamination of test materials, aged blood specimens, inadequate incubation time or temperature, improper centrifugation, improper storage of materials or omission of test samples. - False positives or hazy reactions may occur in patients with rouleaux caused by high plasma concentrations of protein. - Antibodies to preservatives, mediations, disease states, Wharton’s jelly, and/or cross contamination of reaction microtubes may cause false positive reactions. - Weak expressions of A,B,and D antigens may not be detected, as well as weak subgroups of A and B. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety Refer to Laboratory Procedure#4002-1: Infection Control and Laboratory Safety Hygiene Procedure and all safety policies of the laboratory |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Related Documents ABO Discrepancies ABO Type: Forward and Reverse, Rh-Gel Card Cell Suspensions and Conversions Cord Blood Workup Interpreting Reactions Recheck Testing Resulting Testing from the Pending Worklog Specimen Requirements, Labeling, and Storage |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| References MTSTM A/B/D Monoclonal Grouping Card Reference Document RMC Procedure: ABO Type- Forward and Reverse, Rh-Weak D- Tube Method RMC Procedure: Recheck Testing – Gel Card RMC Procedure: Infant Direct Antiglobulin Test – Gel Method |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: | Author: |  |  |  |  |  |  | Date: |  | Validated by: |  |  |  |  | Date: |  | Approved by: |  |  |  |
| 11/9/11 | Mary Farlow MT(ASCP) SBB |  |  |  |  |  |  | 3/29/12 |  | Cheryl Webster, MLSCM(ASCP) |  |  |  |  |  |  |  |  |  |  |
| L: Blood Bank: Infant ABO Forward Typing, Rh – Gel Card |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Initials: |  |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Discontinued |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  | Reason: |  |  |  |  |  |  |  |  |  |  |  |  |  | Initials: |  |  |