| Ridgeview Medical Center Laboratory Services |  |  | Blood Bank |  | Effective Date: 9/26/08 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Principle This procedure provides directions for determining the presence or absence of the Jka antigen on red cells. - The Kidd system is a clinically significant antibody because it can be produced by stimulation and may cause hemolytic transfusion reactions, mainly delayed reactions, and hemolytic disease of the newborn. - Once an antibody has been identified in the patient’s serum, their red blood cells must be tested and found negative for the corresponding antigen. - Red blood cell components selected for transfusion must be negative for the corresponding antigen. |  |  |  |  |  |  |  |
| Specimen The specimen of choice is fresh EDTA red blood cells or segments from donor units. Refer to Specimen Requirements, Labeling, and Storage. |  |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |
| Reagents |  | Supplies |  |  |  | Equipment |  |
| - Anti Jka-Bioclone stored at 2-8 C - Heterozygous antigen positive and antigen negative panel cells stored at 2-8 C - Saline |  | - 12 x 75 tubes - Pipet tips - “Antigen Negative” stickers |  |  |  | - Centrifuge/cell washer - Pipettor - Blood Bank mirror |  |
| Quality Control To confirm the specificity and reactivity of antisera, reagents must be tested with appropriate controls each day of use. Positive and negative controls must be chosen from an antibody panel identification kit: - Negative control should be a red cell lacking the appropriate antigen. - Positive control should be a red cell which exhibits heterozygous expression of that antigen. Jka; Jkb. +; +; Heterozygous. +; 0; Homozygous. ***Testing of control cells must exhibit expected results for testing to be valid and reportable.*** Procedure A direct antiglobulin test must be performed in conjunction with the antigen typing Refer to Direct Antiglobulin Testing Red cells which demonstrate a positive DAT cannot be accurately tested. Procedure Note – Unit Antigen Testing - Check Blood and Blood Component Packing List for historical typing. If unit is available in bank, proceed to confirm antigen type (Historical typing is NOT acceptable as the type of record) - Once units are determined to be antigen negative or positive, fill out the white “Antigen Tested” label and place on corresponding units. - If unit is tagged from Hospital Services with an Antigen Negative tag, unit does not need to be retyped. - If historical typing is not available, choose and type units as required. If necessary, order antigen negative units from supplier, as detailed in Ordering Products from Supplier. |
| Step | Action |  |  |  |  |  |  |
| 1. | Prepare a 4 % suspension visually of patient and or unit red blood cells in saline. |  |  |  |  |  |  |
| 2. | Prepare a 4% suspension of heterozygous positive and negative control cells according to Cell Suspensions and Conversions procedure. |  |  |  |  |  |  |
| 3. | Label test tubes for each patient and/or unit, negative control cell and positive control cell. |  |  |  |  |  |  |
| 3. | Add one drop of Anti-Jka to all labeled tubes. |  |  |  |  |  |  |
| 4. | Add one drop of the positive control cell and negative control cell suspensions to the appropriate tubes. |  |  |  |  |  |  |
| 5. | Add one drop of the 4 % patient and/or unit cells to the appropriately labeled tubes. Mix gently. |  |  |  |  |  |  |
| 6. | Incubate at room temperature (15 to 30C) for a minimum of 5 minutes up to a maximum of 30 minutes. |  |  |  |  |  |  |
| 7. | Centrifuge for 15 seconds at 3400 rpm. |  |  |  |  |  |  |
| 8. | Resuspend the cells by gentle agitation and examine macroscopically for agglutination in the mirror. |  |  |  |  |  |  |
| 9. | Grade the reaction as negative, 1+, 2+, 3+, or 4+ and record the test results. |  |  |  |  |  |  |

Interpretation Refer to Interpreting Reactions. ▪ Agglutination of the red cells constitutes a positive test result and indicates the presence of Jka, providing the test cells do not have a positive direct antiglobulin test. - No agglutination constitutes a negative result and indicates the relevant antigen is not demonstrable. Result Reporting Refer to Antigen Testing-Quick Guide to enter results of QC, patient, and unit testing in HBB.

Limitations - Aged red blood cells may yield weaker reactions than those obtained with fresh red blood cells. - Red blood cells which are DAT positive may give weak reactions. In extreme cases, false negative reactions may occur. - Determination of patient phenotype can be difficult if the patient has been transfused within the preceding three months. - Interpretations of current testing must be compared with the patient’s past records to detect a possible error or a potentially dangerous situation.

Safety Refer to Laboratory Procedure #4002-1: Infection Control and Laboratory Safety Hygiene Procedure and all other safety policies of the laboratory.

Related Documents Specimen Requirements, Labeling, and Storage Cell Suspensions and Conversions Direct Antiglobulin Testing Interpreting Reactions Resulting Testing from the Pending Worklog

References AABB Technical Manual, 15th Edition, 2005, pages 345-346. Blood grouping Reagents, Anti Jka Bioclone, Ortho package insert, revised May 19, 1999. HBB Quick Reference Guide Standards for Blood Banks and Transfusions Services, 23rd Edition, 2005, page 38-39.

| Date: | Author: |  |  |  |  | Date: |  | Validated by: |  |  |  | Date: |  |  | Approved by: |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3/7/07 | Jo Lohmann, MT(ASCP) Cheryl Webster, MT(ASCP) |  |  |  |  | 4/09/08 |  | Cheryl Webster, MT(ASCP) |  |  |  | 6/20/2008 |  |  | Dr. T. Hallin |  |  |
| Jka Antigen Typing-Bioclone |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Annual Review |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date: |  | Initials: |  | Revision and Reason |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Discontinued |  |  |
| --- | --- | --- |
| Date: | Reason: | Initials: |