|  | LABORATORY DEPARTMENT POLICIES AND PROCEDURES | Department: CHEMISTRY |
| --- | --- | --- |
|  |  | Number: 356.0-Chem-gu-rev12/12 |

LIPOCLEAR

- PURPOSE

Lipemia interferes with various assays on the COBAS 6000. LipoClear is a non-toxic, non-ionic solution that is used to remove lipoproteins in grossly lipemic specimens, yielding a clear specimen suitable for analysis.

SPECIMEN REQUIREMENTS

Fresh serum or plasma (Lithium Heparin) at room temperature is used.

REAGENT REQUIREMENTS

LipoClear LC 40

- PROCEDURE

- Allow reagent microtainers to come to room temperature.

- Add 0.5 mL of plasma to the LC 40 microtainer.

- Vortex for 10 seconds.

- Allow sample to sit at room temperature for 5 minutes.

- Using the STATSPIN EXPRESS 2 (coagulation centrifuge in hematology), centrifuge the specimen on the hard spin setting which is the 180 sec setting.

- When specimen has been spun, using a disposable pipette, go through the pellicle at the top and aspirate the cleared specimen and transfer to a properly labeled Roche sample cup for testing.

- Proceed with testing.

- All results are to be multiplied by 1.2. This value was verified by VVMC laboratory.

LIMITATIONS

Specimens treated with LipoClear can only be used for the following assays: ALBUMIN, TOTAL BILIRUBIN, AST, ALT, GLUCOSE and BUN. These are the only assays that are affected by lipemia. Do not use for any other assays on the COBAS 6000, including electrolytes. Use the straight, lipemic specimen for assays unaffected by lipemia. Specimens treated with LipoClear are not acceptable for any other types of testing (coagulation, serology or hematology).

If specimen is not completely cleared with the first treatment, treat the specimen again as described above and multiply the results by 1.44.

REFERENCES  LipoClear Package Insert 54-004752-001 Rev. C

| Prepared by/Title/Date: Robert Laferriere, MT 12/5/2012 Approved by/Title/Date: Signature on file in lab | Committee Approval/Date: Policy & Procedure ______ | Dates Reviewed/Revised: |
| --- | --- | --- |