|  | GENERAL LABORATORY POLICIES AND PROCEDURES | Department: MICROBIOLOGY |
| --- | --- | --- |
|  |  | Number: 1501.0-Micr-qc-rev09/13 |

QUALITY CONTROL OF TEST REAGENTS, STAINS, PANELS AND KITS

DAY OF USE:

India Ink					+   C. neoformans ATCC 34877

-    C. albicans 10231

WEEKLY:

Gram Stain  (included in Gram Slide QC)	+  S. aureus 29213

*(gst rgt. Exp date for iodine changes		-   E. coli 25922

to 3 months when open)

Positive Phoenix Panels       	 	            E. faecalis 29212

MIC/ID				E. faecalis 51299

S. aureus 29213

S. aureus 25923

S. aureus BAA-976

S. aureus BAA-977

Negative Phoenix Panels			E. coli 25922

MIC/ID				P. aeruginosa 27853

K. pneumoniae 700603

E. coli 35218

Strep Phoenix Panels				S. pneumoniae 49619

MIC

NEW LOT OR NEW SHIPMENT:

Beta Lactamase (Cefinase)		            +  S. aureus 29213

*(exp.date changes to 60 days			-  H. influenzae 10211 or

when open)					+  M. catarrhalis  8176

-  S. epidermidis 12228

Bile Solubility					+  S. pneumoniae 6303

-   S. salivarius 13419

PYR disk					+  S. pyogenes 19615

-   S. agalactiae 12386

Bacitracin (A disk)				+  S. pyogenes 12384

-   S. agalactiae 12386

Optochin (P disk)				+  S. pneumoniae 6305

-   S. pyogenes 12384

Hippurate 				            +  G. vaginalis 14018 or

+  S. agalactiae 12386

-   S. pyogenes 19615

Ninhydrin                                                        + G. vaginalis 14018 or

+ S. agalacticai 12386

-   S. pyogenes 19615

Catarrhalis Disk				+  M. catarrhalis 25238

-   N. lactamica 23970

Catalase                                                           +  S. aureus 29213

-   S. agalactiae 13813

Oxidase                                              	            +  P. aeruginosa 27853

-   E. coli 25922

Coagulase (Staphaurex)		            +  S. aureus 29213

-   S. epidermidis 12228

Candida albicans kit			            +  C. albicans 60193

-   C. krusei 14243

PBP2a Kit                                                       +  S.aureus 43300

-  S.aureus 25923

RapID ANA					C. sordellii 9714

B. distasonis 8503

B. uniformis 8492

RapID NH					H. influenzae 9006

H. parainfluenzae 7901

O. urethralis 17960

M. catarrhalis 8176

RapID STR					E. faecalis 29212

E. durans 11576

S. bovis 9809

S. pyogenes 19615

Positive Phoenix Panels 	 	            S. aureus 29213

ID					E. faecalis 29212

Negative Phoenix Panels		            E. coli 25922

ID					P. aeruginosa 27853

Strep Phoenix Panels				S. agalactiae 13813

ID

E Test (use for Strep isolated from sterile sites)			MIC’s (ug/ml)

a.)  Vancomycin (VA)			S. aureus 29213	(0.5-2.0)

(use for VRE confirmation)        S. pneumoniae 49619	(0.125-0.5)

E. faecalis 29212	(1.0-4.0)

b.)  Penicillin G (PG)			S. aureus 29213	(0.25-1.0)

E. faecalis 29212	(1.0-4.0)

S. pneumoniae 49619	(0.25-1.0)

c.)  Ceftriaxone (TX)			H. influenzae 49247	(0.06-0.25)

S. pneumoniae 49619	(0.03-0.12)

Streptex kit 					Use organisms included in kit

*(exp date of extraction enzyme changes to

3 months after reconstitution)

C. Difficile kit					Use organisms included in kit

Cryptosporidium Rapid assay kit	            Use separate Pos/Neg QC

Giardia Rapid assay kit			 Use separate Pos/Neg QC

Clo Test					Use separate Positive QC

Shiga Toxin Test				Use organisms included in the kit

When using a commercial product, QC is preformed precisely according to the manufacturer’s recommendations.  This includes, but is not limited to, Antimicrobial Susceptibility Testing/Identification (AST/ID) systems.

All biochemical tests in each new lot number and shipment must be evaluated with a known positive and negative organism, to assure appropriate reactivity.  Record results on the Daily QC logs for appropriate tests.

REFERENCES: CLSI (2005). Dade Behring Microscan Panel (2005).

| Prepared by/Title/Date: Dawn Guay/MT(ASCP)/08/05 Approved by/Title/Date: | Committee Approval/Date: Policy & Procedure ______ | Dates Reviewed/Revised: 06/09 09/13 02/11 06/13 |
| --- | --- | --- |