|  | LABORATORY DEPARTMENT POLICIES AND PROCEDURES | Department: Chemistry |
| --- | --- | --- |
|  |  | Number: 355.0-Chem-gu.rev10/13 |

TITLE: PROCEDURE FOR DILUTIONS FOR COBAS 501/6000

PURPOSE There are select assays that may require off-board dilution when the results exceed linearity despite use of the automated dilution protocol. These assays will be diluted manually and programmed manually on the Cobas 501/6000. Results must be multiplied by the dilution factor. Maximum manual dilution is 1:1 unless specifically requested by physician. Any test with an automatic dilution, that dilution is the maximum for that test. PROCEDURE - Refer to each assay package insert to determine the proper diluent. - Prepare selected dilution in the following manner: a. 1:1 Add 0.2mL of specimen to 0.2mL diluent b. 1:5 Add 0.1mL of specimen to 0.4 mL diluent c. 1:10 Add 0.1mL of specimen to 0.9mL diluent d. 1:20 Add 0.1mL of specimen to 1.9mL diluent e. 1:100 Add 0.1mL of specimen to 9.9mL diluent - Manually program the test to be performed as a STAT which will yield a printout. Indicate the dilution which has been used in the sample ID box. This can follow the accession number and will negate auto-verification. - Once assay is complete multiply the results by the dilution factor. - Enter manually into Laboratory Information System. - Attach dilution canned comment (F2, ‘D’), ‘Result verified by dilution’, to result. ATTACHMENTS N/A REFERENCES Cobas 6000 Operator’s Manual

| Prepared by/Title/Date: Robert Laferriere MT/ASCP 2/9/2009 Approved by/Title/Date: Signature on file in Lab | Committee Approval/Date: Policy & Procedure ______ | Dates Reviewed/Revised: 2/11 RL 10/13 RL |
| --- | --- | --- |