|  | LABORATORY DEPARTMENT POLICIES AND PROCEDURES | Department: MICROBIOLOGY |
| --- | --- | --- |
|  |  | Number: 1330.0-Micr-gu-rev10/13 |

C.Diff by PCR (Illumigene)

- POLICY

Toxigenic Clostridium difficile is a major cause of antibiotic associated diarrhea and colitis and is the causative agent for virtually all cases of pseudomembranous colitis.  Although about 2% of normal healthy adults are colonized with C. difficile, many patients acquire this organism through nosocomial infection.  Exposure to most antibiotics is thought to allow proliferation of toxigenic C. difficile by disrupting the normal intestinal flora. Two large toxin proteins (TcdA [or toxin A] and TcdB [toxin B]) are thought to be the primary virulence factors of C. difficile.  These toxins are encoded by two separate genes, named tcdA and tcdB, respectively.  Together, with three additional genes, they form a 19.6kb pathogenicity locus called PaLoc.

The illumigene C. difficile DNA molecular assay is based on loop-mediated amplification technology, which uses specifically designed primers to the PaLoc pathogenicity locus to provide for specific and continuous isothermal DNA amplification.  A by-product of this amplification is the formation of magnesium pyrophosphate, which forms a white precipitate leading to a turbid reaction solution.  This presence of turbidity signifies a positive reaction while the absence of turbidity represents a negative reaction.  The illumigene C. difficile assay contains primers that specifically amplify a 204 bp region of the conserved 5’ sequence of the tcdA gene within the PaLoc of toxigenic C. difficile in diarrheal stool samples from patients suspected of having C. difficile-associated disease (CDAD). The results of the assay are determined using the Meridian illumipro-10 Incubator / Reader.

Due to the sensitivity of the assay, the Microbiology Lab will accept only one specimen during each five day period.  Specimens received by the laboratory are batched and testing is performed daily.

SPECIMEN:

Preferred Sample Types:

1.   Human stool samples, unpreserved

2.   Human stool samples, preserved in Cary-Blair-based media

3.   Unformed samples indicative of C. difficile-associated disease (CDAD)

Undesirable samples: Formed stools with no indication of C. difficile-associated disease (CDAD)

Collection and Storage:

Samples should be transported in a clean, airtight container and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 24 hours at 21-27 C or 5 days at 2-8 C

MATERIALS AND EQUIPMENT:

- MATERIALS PROVIDED:

- illumigene Sample Preparation Apparatus:  Sampling unit consisting of sample preparation chamber, dropper tip, cap and  Sample Dilution Buffer (Phosphate Buffered Saline and formalin treated Staphylococcus aureus, with Sodium Azide (0.09%)as a preservative).

- illumigene Reaction Buffer: Tris-buffered solution containing sodium azide (0.09%)as a preservative.

- illumigene C. difficile Assay Device: Two separate chambers containing dry reagent lyospheres comprised of DNA  polymerase, dinucleotide tri-phosphates (dNTPs), and either C. difficile specific primers (TEST Chamber) or S. aureus primers (CONTROL Chamber).

- illumigene Sample Collection Brushes

- illumigene Sample Extraction Tubes

- MATERIALS NOT PROVIDED:

- Materials:

- illumigene External Control Kit

a)  Available, Catalog Number: 279920

b)   illumigene Negative Control: Tris-buffered solution containing non-infectious Plasmid DNA (S. aureus insert)

with azide as a preservative.

c)   illumigene C. difficile Positive Control: Tris-buffered solution containing non-infectious Plasmid DNA (S.

aureus and C. difficile inserts) with azide as a preservative.

- Disposable latex gloves, powderfree

- DNAse/RNAse free, aerosol resistant pipette tips

- Known positive clinical sample or a negative sample spiked with toxigenic C. difficile

a)  Available, Catalog Number: 500618

Equipment:

- Dry-bath with 12mm heat block capable of 95 C

- Vortex Mixer

- Interval Timer

- Micropipette capable of dispensing 50 μL

- Micropipette capable of dispensing 200 µL

- illumipro-10  Incubator/Reader

- REAGENT PREPARATION:

Ensure kit reagents are at room temperature (21 - 27 C) before use.  Incorrect results may be obtained if reagents are not brought to room temperature prior to use.

- PRECAUTIONS:

- All reagents are for in vitro diagnostic use only.

- Follow Biosafety Level 2 and Good Laboratory Practices during testing.  Treat all specimens and used test devices as capable of transmitting infectious disease.  Do not eat, drink or smoke in areas where specimens or kit reagents are handled.

- Wear disposable gloves while handling specimens and thoroughly wash hands afterwards.

- Quality Control Programs for Molecular Testing Laboratories should be employed.

- illumigene Sample Dilution Buffer contains formalin inactivated organisms and should be treated as potentially infectious.

- The illumigene C. difficile Test Device contains lyophilized reagents.  The protective pouch should not be open until ready to perform the assay.

- The illumigene C. difficile Test Device includes a latch feature that is designed to prevent contamination of the test area with amplification product. Do NOT use Test Devices with broken latches.

- Dispose of used illumigene Test Devices immediately after processing, leaving the device latch securely in place. Opening the device after amplification may result in contamination of the test area with amplification product.

- SHELF LIFE AND STORAGE:

The expiration date is indicated on the kit label.  Store the kit at 2 - 27 C.

Do not use the Test cards or reagents after their expiration dates.

At this facility, kits are stored:   Room Temperature (21 – 27 C)

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

  - Each device contains an internal control well that controls for amplification inhibition, assay reagents and sample processing effectiveness.

  - Good laboratory practice recommends the use of control materials.  Users should follow the appropriate federal, state and local guidelines concerning the running of external quality controls.

  - Illumigene C. difficile External Control Reagents are supplied separately (Catalog 279920).  It is recommended that the reactivity of each new lot and each new shipment of Illumigene C. difficile be verified on receipt and before use.  External control tests should be performed in accordance with appropriate federal, state, and local guidelines.  The illumigene C. difficile test kit should not be used in patient testing if the external controls do not produce the correct results.  Controls are stored at 2-8 C when not in use.

  - A separate device must be used for each control.

  - DO NOT extract the Positive or Negative Control samples.

QC Testing Frequency and Documentation:

For this facility, External QC and a known Positive and Negative patient are run:   With each new lot and each new shipment or every 30 days, whichever is more frequent.

Results of External QC and action(s) taken when control results are unacceptable are documented:  on the ‘QUALITY CONTROL DISCREPANCY FORM/LOG OF UNUSUAL RESULTS’ FORM

Procedure:

NOTE:  Ensure that the illumipro-10 instrument is powered on and required performance verifications have been completed prior to initiation of SPECIMEN PREPARATION. Refer to the illumipro-10 Operator’s Manual for further information regarding instrument set-up and operation.

- Specimen Preparation (Extraction):

1.	Mix stool sample thoroughly.

2.	Collect mixed sample specimen using Sample Collection Brush.

a. Liquid Stool:  Immerse Sample Collection Brush completely into specimen.

b. Semi-solid Stool: Rotate Sample Collection Brush over specimen surface, lightly coating the surfaces of the brush bristles. Approximately one-half of the brush should be coated; over-collection of stool may lead to clogging of the sample collection device.

3.	Add the Sample Collection Brush to the Sample Preparation Apparatus containing Sample Diluent. Break-off the Sample Collection Brush at the ‘Break’ mark on the handle.  Replace and secure the Sample Preparation Apparatus cap.  . For stool in Cary-Blair media, transfer 200 µL of specimen to the illumigene® C. difficile Sample Preparation Apparatus and secure the cap.  Vortex the Sample Preparation Apparatus for a minimum of 10 seconds.  Specimen in illumigene® C. difficile Sample Preparation Apparatus prior to heat treatment may be held at 2-27 C for up to 24 hours prior to testing.

4.    Remove the tip cap from the Sample Preparation Apparatus and squeeze five to ten drops of sample into a clean illumigene Extraction Tube.

5.	Repeat Sample Preparation Steps for all samples to be processed.

6. 	Heat the illumigene Extraction Tubes containing mixed sample in a dry-bath/heat block at 95 (+/- 5) C for 10 (+/- 2) minutes.

7.   Remove the tubes from the dry-bath/heat block and vortex for approximately 10 seconds. Extracted Samples may be held at 2-8˚C for 4 hours OR may be frozen and stored a ≤-20˚C for 1 day prior to addition to Reaction Buffer.

TEST PROCEDURE:

- This test should be performed by qualified personnel per local regulatory requirements.

NOTE: A maximum of ten samples can be processed in a single illumipro-10 run.

- Transfer 50 µL of extracted sample to an appropriately labeled illumigene Reaction Buffer tube.

- Vortex the Reaction Buffer Tube containing extracted sample for approximately 10 seconds.

- Repeat steps 1 and 2 for all the samples to be analyzed before proceeding.

- Remove one illumigene C. difficile Assay Device from its protective pouch per sample. Carefully open the device, holding the tubes such that the lyophilized reagent will not fall out upon opening.  Place device on a flat surface or in a rack that can accommodate the device.

- Using a new pipette tip, transfer 50 µL from the Reaction Buffer tube containing extracted sample to the TEST chamber (White Bead) of the illumigene Assay Device.  Do not insert air bubbles. Using a new pipette tip, transfer 50 µL from the reaction buffer tube containing extracted sample to the CONTROL chamber (Yellow Bead) of the illumigene device.  Do not insert air bubbles. Close the illumigene Assay Device and fasten the latch securely.

- Tap device on the bench top to mix and remove bubbles.  Carefully examine the reaction tubes to ensure that there are no air bubbles left in the tube.

- Insert the illumigene Assay Device into the illumipro-10 and initiate amplification reaction and detection.  Results will be displayed at the conclusion of the run.

INTERPRETATION OF RESULTS:

| Sample ID | Reported Result | Interpretation |
| --- | --- | --- |
| Patient Specimen | POSITIVE | Sample contains toxigenic C. difficile strain with the pathogen locus (PaLoc). |
|  | NEGATIVE | No toxigenic C. difficile detected. |
|  | INVALID | No reportable result. Repeat the test using the original stool sample. Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure. |
| Extraction Control | POSITIVE | Valid Extraction Control. Extractions performed correctly. |
|  | NEGATIVE | Extraction Failure. Do not report test result. Repeat the extraction process for all stool samples. |
|  | INVALID | No reportable result. Do not report test result. Repeat the extraction process for all stool samples. Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure. |
| Positive Control | POSITIVE | Valid positive control result. Reagents active at time of use, illumipro-10 performing correctly. |
|  | NEGATIVE | Incorrect control result. Repeat control testing as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | INVALID | No reportable result. Repeat entire assay run using original stool samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |
| Negative Control | POSITIVE | Incorrect control result. Repeat control testing as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | NEGATIVE | Valid negative control result. Reagents active at time of use, illumipro-10 performing correctly. |
|  | INVALID | No reportable result. Repeat entire assay run using original stool samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |
| EMPTY WELL | NONE | No illumigene Test Device in the illumipro-10 Well. OR The illumigene Test Device present is compromised due to sample preparation failure, dirty device or improperly seated device. Repeat the test using original sample. |

CALCULATIONS:

There are no calculations associated with this procedure.

REPORTING OF RESULTS:

Positive Test: Sample contains toxigenic C. difficile strain with the pathogen locus (PaLoc)

Negative Test: No toxigenic C. difficile detected.

EXPECTED RESULTS:

The frequency of antibiotic-associated diarrhea caused by C. difficile is dependent on several factors, including patient population, type of institution, and epidemiology.  The incidence of CDAD in patients suspected of having antibiotic-associated disease is 15-20%, although different facilities may find positive rates above or below this range.  The incidence of toxigenic C. difficile during the 2010 period of this study was 15.5%.

PERFORMANCE CHARECTERISTICS:

- Refer to Directional Insert- Meridian Bioscience illumigene C. difficile

LIMITATIONS OF THE PROCEDURE:

- Assay does not distinguish between viable and non-viable organisms.

- Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

- This test is for use with unformed stools preserved or unpreserved stools only.  Performance characteristics of other clinical specimen types have not been established.

- Two distinct groups have been identified that can harbor C. difficile asymptomatically at very high rates.  Colonization at rates up to 50% and higher have been reported in infants and rates up to 32% in cystic fibrosis patients.

- This assay does not identify antimicrobial susceptibility.

- This test detects but does not differentiate the NAP1 (Ribotype 027) strain from other toxigenic strains of C. difficile.

- The detection of bacterial nucleic acid is dependent on proper specimen handling (including transportation and storage) and preparation (dilution and extraction).  Failure to follow instructions for collection, handling, and preparation may cause incorrect results.

REFERENCES:

- Refer to Directional Insert- Meridian Bioscience illumigene C. difficile

| Prepared by/Title/Date: Dawn Lavash/MT(ASCP)/02-11 Approved by/Title/Date: Signature on file in Lab | Committee Approval/Date: Policy & Procedure ______ | Dates Reviewed/Revised: 06/11 07/11 10/11 10/13 |
| --- | --- | --- |