PARKER INDIAN HEALTH CENTER

COLORADO RIVER SERVICE UNIT

POLICY AND PROCEDURE

| ESTABLISHED BY: Unknown | MANUAL CODE: POCT ADMIN CODE: | JCAHO Standard(s) WT.05.01.01 |
| --- | --- | --- |
| SUBJECT: ACCUCHECK INFORM PROCEDURE |  |  |
| EFFECTIVE DATE: 3 February 1997 | REVISION DATE: 13 September 2012 | SUPERSEDES: 10 December 2009 |
| DISTRIBUTION: Colorado River Service Unit, POCT locations |  | REVIEW DATES/INITIALS: See last page |

I.	POLICY

Whole blood glucose levels are used for screening and monitoring patients with diabetes mellitus. Whole blood glucose testing is performed at point-of-care testing (POCT) locations using the Accucheck Inform Meter.

II.	PRINCIPLE

The enzyme glucose dehydrogenase converts the glucose in the blood sample to gluconolactone.  This reaction liberates an electron that reacts with a coenzyme electron acceptor, the oxidized form of the mediator hexacyanoferrate (III), forming the reduced form of the mediator, hexacyanoferrate (II).  The Accu-Chek Comfort Curve test strips employ the electrochemical principle of biamperometry.  The meter applies a voltage between two identical electrodes, which causes the reduced mediator formed during the incubation period to be reconverted to an oxidized mediator.  This generates a small current that is read by the meter.

III. 	 PURPOSE

The Accu-Chek Advantage consists of a lightweight, portable monitor, a base unit, a download interface and an Information Management system.  The portable monitor provides whole blood glucose results within seconds at the patient's bedside, enabling the medical care provider to initiate medical intervention immediately.  The interface and information management system allow for automated tracking of patient, quality control, linearity and proficiency testing data, as well as instrument maintenance and workload statistics.  The entire system is interfaced to a laboratory or hospital information system to further automate record keeping.

IV. 	REAGENT PREPARATION

ACCU-CHEK COMFORT CURVE TEST STRIPS (CAT. NO. 2030365)

Each test strip has a YELLOW Target Area containing the following chemicals:

Potassium ferricyanide  			     	43.7%

Glucose dehydrogenase **			      	 1.2%

Buffer										 	24.7%

Stabilizer									 	19.4%

Nonreactive ingredients				 		11.0%

** (from A.calcoaceticus,recombinant from E. coli)

Test strips must be stored at room temperature.  Do not freeze.  Test strips are stored in the same tightly capped vial in which they are packaged.  The vial cap is immediately replaced after removal of a test strip.  Opened and unopened test strips are stable until the expiration date on the vial.  Test strips are used at temperatures between 14º and 40º C (57º and 104º F) and <85% humidity.

# V. 	ACCU-CHEK INFORM SYSTEM

The Inform System consists of an AccuChek Inform Meter, an Inform Base Unit, download cable attached to a computer with the remote unit download software installed and the DataCare Information Management System or RALS dataport with the RALS Information Management System.

A. 	AccuChek Inform Meter

- The Accu-Chek Inform Meter is handled with care. Sudden shocks caused by dropping or rough treatment may affect performance. If the Accu-Chek Inform Meter is dropped, performance must be verified by quality control testing.  The Accu-Chek Inform Meter is stored away from direct sunlight and extreme temperatures.

- B. 	Cleaning and maintenance of the Accu-Chek Inform Meter

- Cleaning and maintenance of the Accu-Chek Inform Meter is performed EVERY 24  hours by the certified operator as follows:

- Wipe the surface of the Accu-Chek Inform Meter with a soft cloth slightly dampened (not wet) with one of the following solutions.

- A freshly mixed solution of 1:10 bleach in water (1 part bleach in 9 parts water), 70% isopropyl alcohol, full strength, Warm soapy water.

- (Note:  Do not spray the Accu-Chek Inform Meter directly with solution, as this could cause solution to enter the case and damage the electronic components.)

- Caviwipes are also used to clean meters.

- Disposable latex gloves are worn when performing preventive maintenance and cleaning on the Accu-Chek Inform Meter and blood glucose testing equipment. Used disposable latex gloves should be discarded in the biohazard waste according the infection control policy.

- Use only the battery pack available from Roche Diagnostics (Cat. No. 3034844) in the Accu-Chek Inform Meter. Using any other type of battery pack may damage the system.  If the Accu-Chek Inform Meter is to be stored for a long period of time, the battery is removed to avoid leakage or damage.  When storing or disposing of batteries, keep or replace in manufacturer’s packing material. Dispose of used batteries appropriately.  Incorrect storage or disposal of batteries could result in a hazardous condition.

#### C. 	Transferring Data from the Accu-Chek Inform Meter

- Data is transferred from an Accu-Chek Inform Meter to a computer with specialized software immediately upon docking the meter in the base unit.  The default prompt is for docking every 24 hours.

- Transferring data from an Accu-Chek Inform Meter is the responsibility of each certified operator.

- Result information can only be cleared from the Accu-Chek Inform Meter after results have been transferred

#### D. 	Documentation and Replacement

- Any maintenance or repair to an Accu-Chek Inform meter is documented according on the maintenance log or entered in the Inform meter.  Each certified operator is responsible for the performance and documentation of preventive maintenance.

- If certified operator’s are unable to correct a problem with the Accu-Chek Inform meter, call the 800 number for help. If still unable to correct the problem ,then the meter must be removed from service . Call and inform the lab for a replacement and change to the location in the system .  The Accu-Chek Inform meter must be cleaned and disinfected before it is sent out for repair or replacement.

- The unit supervisor or an appointed individual will review the preventive maintenance log for completion and any corrective action monthly.

- The Laboratory  oversee record keeping for compliance with acceditiation standards and laws.

VI. 	CALIBRATION

The meter is “calibrated” when the instrument is turned on with the Code Key inserted.

The operator must verify test strip code information in the Accu-Chek Inform System whenever a patient or quality control test is performed.  Coding is always verified by matching the code on the Accu-Chek Inform display screen with the code number printed on the side of the vial of test strips.  The test strip code displayed by the Accu-Chek Inform System must match the code of the test strips in use.  If not, the instrument will not work and the meter must be recoded (recalibrated) and the new code information must be entered in the Accu-Chek Inform System.

It is recommended that the Code Key be changed with each new vial of test strips.  Place the new Code Key in the meter and discard the old Code Key.

Code Key Procedure:

- Remove the Code Key from the test strip box.

- Verify the three-digit number on the Code Key matches the number on the test strip vial.

- Remove old Code Key from Accu-Chek Inform meter.

- Snap the new Code Key (slots facing towards the meter) into the Code Key slot with the printed side facing up.

- Leave the Code Key in the meter.

VII. 	QUALITY CONTROL

A. 	Accu-Chek Comfort Curve Glucose Control Solutions - 2-level (Cat. No. 2030390)

Each 4 mL bottle of glucose control solution contains:

Glucose

Buffer

Biological Salt

Microbial Dilution

Surfactant

Coloring Agent

Glucose control solutions must be stored at room temperature (<92oF or <32oC). Do not freeze. Glucose control solutions are stable for three months after opening or until the expiration date, whichever comes first. The date the vial is opened should be written on the vial label.  Any outdated glucose control solutions will be discarded.  Control solutions should be clear.  Do not use if cloudy or shows other evidence of contamination.

- The test strip lot number and the acceptable glucose control ranges are found on the label of each vial of test strips.

- Control tests are performed at the following times:

    - Each time a new vial of test strips is opened

    - When a vial of strips has been left opened

    - If the Accu-Chek Inform System has been dropped

    - Each  day of use

    - When the test strips have been exposed to extreme heat, humidity, or cold

    - When test results contradict clinical symptoms

    - After the battery in the Accu-Chek Inform System has been replaced or after the Accu-Chek Inform System has been recoded.

    - When a new batch of control test liquid with a lot # is received .

- B. 	Quality Control Procedure

- Press power ON button.

- Enter (or scan) your operator ID, then press the forward arrow button.

- Select Control Test.

- Select the desired control level:  Level 1 or Level 2.

- Verify the lot number of control solution displayed on the Accu-Chek Inform System, which has been                                               preprogram

- Select YES if the lot number is correct.

- Verify that the strip code number on the test strip vial matches the code number on the Accu-Chek Inform System, which has been preprogram.

- Select YES if the code numbers match.

- Select NO to enter a new code.

- Remove a test strip from the vial and replace the vial cap immediately.

- When the flashing strip icon appears on the meter display, gently insert test strip with the yellow target area or test window facing up. (Insert the end with the silver bars.)

- Note:  Insert test strip BEFORE dosing.

- Using the Accu-Chek Comfort Curve test strip.  Touch and hold drop of glucose control solution to the curved edge of the yellow target area.

- The glucose control solution is drawn into the test strip automatically.

- Do not allow the tip of the control solution bottle to touch the test strip.

- An hourglass will be displayed on the Accu-Chek Inform meter while waiting for the result.

- Review the result and enter the appropriate comment(s). Then press the forward arrow button to record the test and to test the next level of control (if other levels are required) or to proceed to patient testing.

- Remove the used test strip(s) and disposable latex gloves and discard them in the biohazard trash.

- Document the result(s):  The date, time, initials or ID of the operator, meter serial number, and quality control result are transferred into the computer database by docking the Inform meter.

If a quality control test result falls within the acceptable control range, it is acceptable to proceed with patient testing.

If a quality control test result falls outside of the acceptable control range, repeat the quality control test if it still falls outside the acceptable control range or the Troubleshooting section of the Accu-Chek Inform System Operator’s Manual is referenced, if needed, and the problem is corrected before proceeding with patient testing.

Any quality control result that falls outside the acceptable control range, along with any corrective action to restore that result to acceptable range, is recorded on the troubleshooting log for the individual instrument or entered in the Inform meter.

Each certified operator is responsible for the performance and documentation of quality control.

The laboratory supervisor or designee will review the quality control records for completion, as well as note any trends that may indicate potential problems. These trends include gradual drifting of values, sudden shifts in glucose control values while using the same lot of strips, and operator performance. This should be completed at least quarterly.  Quality control records will be retained for a minimum of two years.

VIII. 	Linearity Testing

A. 	Accu-Chek Comfort Curve Linearity Test Kit (Cat No. 2030357)

Six levels of glucose solution containing:

Glucose

Buffer Agents

Surfactants

Preservative

FD and C Blue #1

Stabilization Agents

Store Glucose Linearity Solutions at room temperature (<92oF or <32oC).  Do not freeze.  Unopened vials are stable at room temperature (<90oF or <32oC) until the expiration date on the label.  The Linearity solutions are stable for three months after first opening the bottles, or until the expiration date on the label, whichever comes first.

Control solutions should be clear.  Do not use if cloudy or shows other evidence of contamination

- B. 	Linearity is required

- Before a blood glucose meter is put into use

- With each new lot of test strips

- With each new shipment of test strips, even if it is the same lot number as a previous shipment

- At least every six months

- Anytime the Accu-Chek Inform System has been repaired

- When controls begin to reflect an unusual trend or are consistently out of range

- The manufacturer’s stated linear reporting range of each Accu-Chek Inform System is 10 mg/dL to 600 mg/dL.  The reportable range of each instrument is verified by the laboratory POCT co-ordinator or a qualified operator for each test strip lot number prior to release for patient testing.

- If a patient test result falls outside of the linear range, it is verified by the laboratory by an alternative method and is reported as less than (<) or greater than (>) the linear limits.

- The linearity results of each Accu-Chek Inform System are filed with each instruments records in the laboratory.  Linearity records are retained for two years.

To record a linearity test in the Accu-Chek Inform System:

- Press power ON button.

- Enter (or scan) your operator ID, then press the forward arrow button.

- From the Main Menu screen, press the forward arrow button.

- Select Admin….

- Select Linearity Test.

- Type in linearity solution lot # from bottle/box.

- Scan or type strip code # from vial.

- Select the linearity solution number for the first test.

- Perform linearity test.

- Select comment(s) and press the forward arrow button to return to the Linearity Test screen.

- Press the forward arrow button to record the test.

- Select the linearity solution number for the next test or press power OFF to turn the Accu-Chek Inform System off.

- Test each Level of Linearity Control Solution a minimum of 2 times each.  Dock the Inform meter to download the data for graphing and printing.

#### IX. 	Proficiency Testing

The Accu-Chek Inform System stores the following information about each proficiency test:

- Test Result

- Operator ID

- Sample ID

- Test Time and Date

- Strip Information

- Comment(s)

- Meter Serial Number

This facility is enrolled in an external proficiency program. Proficiency tests are performed 2-3 times per year, by a qualified operator.

To record a proficiency test in the Accu-Chek Inform System:

- Press the power ON button.

- Enter (or scan) your operator ID, then press the forward arrow button.

- Press the forward arrow button to display the Main Menu 2 screen.

- Select Proficiency.

- Enter (or scan) the sample ID and press the forward arrow button.

- Verify the strip code information:

- Select YES if the code is correct.

- Select NO to enter a new code.

- Perform the proficiency test.

- Enter comment(s), if necessary.

- Press the forward arrow button to return to the Main Menu 2 screen to run the next sample. Or press the power OFF button to turn the Accu-Chek Inform System off.

- Remove the strip and discard it according to your facility’s infection control policy.

X. 	PERSONAL PROTECTIVE EQUIPMENT

Because of the hazardous nature of handling blood products, it is recommended that disposable latex gloves be used when collecting specimens, performing test procedures, and cleaning blood glucose meter equipment.

Gloves are to be removed and hands washed thoroughly with soap and water after completing the test procedure and prior to handling equipment not related to the procedure. Used disposable latex gloves should be discarded according to the facility’s infection control policy.

Universal precautions should be observed for all blood specimens. They should be handled at Biosafety Level 2 as recommended for any potentially infectious material in the Centers for Disease Control/National Institutes of Health manual, Biosafety in Microbiological and Biomedical Laboratories, 1988 or in the National Committee for Clinical Laboratory Standard Document M29, Protection of Laboratory Workers from Infectious Disease Transmitted by Blood, Body Fluids, and Tissue, 1991.

XI. 	SPECIMEN COLLECTION AND HANDLING

- Capillary, venous, neonatal (including cord blood), and arterial whole blood specimens may be used for testing on the Accu-Chek Inform System with Accu-Chek Comfort Curve Test Strips (Cat. No. 2030365).

- The capillary sample must be tested immediately after collection.

- Blood glucose determinations using venous and arterial blood specimens should be performed within 30 minutes of specimen collection to avoid glycolysis. Mix samples thoroughly.

- For best results with arterial and venous blood, the following anticoagulants/preservatives are recommended: heparin and EDTA.

- Serum separator tubes and red-topped tubes are acceptable if blood is used immediately before the clotting process begins.

- Iodoacetate or fluoride/oxalate should not be used as a preservative.

- Caution should be taken to clear arterial lines before blood is drawn and dosed on the test strip.

- Sufficient sample size is required to ensure accurate results.

- Follow Service unit policy and procedures for blood collection by capillary, venous or arterial methods.

XII. 	TEST PROCEDURE

The Accu-Chek Inform System is set up to store the following information about each patient test:

- Test Result

- Operator ID

- Patient ID

- Test Strip Information

- Test Time and Date

- Comment(s)

- Meter Serial Number

Blood glucose tests with the Accu-Chek Inform System must be ordered by a physician unless the patient is experiencing symptoms of hypoglycemia or hyperglycemia, and quality care dictates a STAT test per policy.

A physician will be notified according to parameters specifically ordered.

The Provider determines what test(s) performed by the laboratory is ordered whenever the blood glucose result is less than 50 mg/dL or greater than 500 mg/dL.

Only a certified operator may perform a blood glucose test on the Accu-Chek Inform System.

NOTE: Any patient result that exceeds the critical range established by the facility is followed-up by the operator by repeating the test and notifying the physician.

The appropriate comment(s) is/are entered in the Accu-Chek Inform System by the operator.

A.	Patient Preparation

- The purpose of the test and the steps of the procedure will be explained to the patient prior to performing the test.  The operator’s hands must be washed before and after testing.

- Universal precautions must be observed and disposable latex gloves must be worn when the operator is handling blood products.

- Fresh capillary whole blood samples for use with the Accu-Chek Inform System are to be taken from the fingertip or heel (neonates).

- Wash your hands and put on disposable latex gloves prior to testing.

- If the patient is able, ask the patient to wash his/her hands with warm water and soap, rinse and dry well prior to testing capillary samples. If the patient is unable, cleanse the puncture site with an alcohol swab and allow it to thoroughly dry. (Alcohol at the puncture site must be dry or an error code/inaccurate result may occur.)

B.	Patient Test Procedure

The following equipment should be at the patient’s bedside prior to testing:

- Accu-Chek Inform System

- Accu-Chek Comfort Curve test strips (Cat. No. 2030365

- Single-use, disposable lancets

- Alcohol swab

- Gauze for wiping finger after stick

- Disposable latex gloves

- Press power ON button.

- Enter (or scan) your operator ID. Press the forward arrow button.

- Select Patient Test.

- Enter (or scan) the patient ID. Press the forward arrow button.

- Verify that the code number on the test strip vial corresponds to the code number on the Accu-Chek Inform System ,which has preprogram

- Select YES if the code numbers correspond.

- Remove a test strip from the vial. Immediately replace the cap on the vial.

- When the flashing strip icon appears on the meter display, gently insert test strip with the yellow target area or test strip window facing up. (Insert the end with the silver bars.)

- Note:  Insert test strip BEFORE dosing.

- When the flashing drop icon appears on the meter display, obtain a blood sample. You may use a whole blood capillary, venous, arterial or neonatal (including cord) blood sample.

- If using the Accu-Chek Comfort Curve test strip:

- Touch and hold drop of blood to the curved edge of the yellow target area.

- The blood is drawn into the test strip automatically.

Important:  If you see any yellow color in the target area or test strip window after you have applied the initial drop of blood, a second drop of blood may be applied to the strip within 15 seconds of the first drop.  If more than 15 seconds have passed, the test result may be erroneous, and you should discard the test strip and repeat the test.

- An hourglass will appear on the display while waiting for the result.

- Enter up to three preprogrammed comments and one custom comment, as necessary. Then press the forward arrow button to record the test and return to the Main Menu screen in order to run the next test.

- Remove the test strip from the meter and discard it according to the facility’s infection control policy.

- Press the power OFF button to turn the Accu-Chek Inform System off.

- Remove gloves and dispose of them according to the facility’s infection control policy. Wash hands thoroughly with soap and water.

- Document the blood glucose result according to facility policy.

XIII. 	CALCULATIONS

None

XIV. 	REPORTING RESULTS

The date, time, initials or ID of the operator, patient name/patient ID number, meter serial number, and glucose value are recorded on/in:

- Accu-Chek Inform System

- Patient’s chart

- Diabetes Flow Sheet

- In-Patient Log

Patient records in the Accuchek Inform Information Management System will be retained for a minimum of 2 years.

When the Inform Information Management system interface is down, send results to the laboratory and the laboratory staff will manually enter all patient test results into RPMS.

A. 	Expected Values

The normal fasting blood glucose range for a non-diabetic adult is 70-110 mg/dL.

One to two hours after meals, normal blood glucose levels for a non-diabetic adult should be less than 140 mg/dL.

B.	Abnormal Results

The Nursing staff has standing orders to notify the requesting Provider if the patient’s Whole Blood Glucose is

- Less than 60 mg/dL.

- Greater than 400 mg/dL.

Notify the physician immediately and document the notification on the PCC.  Follow medical orders for instituting appropriate therapy, including referring or transferring the patient to the Emergency Room.

The Inform meter and AccuChek comfort curve test strip combination have been found to loose accuracy at values greater than 500 mg/dL.  The Provider may wish to validate the result greater than 500 mg/dL by laboratory testing.

XV. 		PROCEDURE NOTES

- If HI is displayed, the blood glucose result may be higher than the reading range of the meter.  If this contradicts the patient’s condition, perform a quality control check with glucose control solution and a new test strip.  If the control result is within the acceptable range, review proper testing procedure and repeat the blood glucose test with a new test strip.  If HI still appears on the patient test, notify the Provider for specific follow up testing.  If the control result is not within the acceptable range, refer to the Accu-Chek Inform System Operator’s Manual before proceeding with patient testing.

- If the error message “testing error-133 A glucose overflow error has occurred, type 71” appears on the display, the blood glucose result may be extremely high and above the meter’s reading range.  If this contradicts the patient’s condition, perform a quality control check with glucose control solution and a new test strip.  If the control result is within the acceptable range, review proper testing procedure and repeat the blood glucose test with a new test strip.  If the error still appears on the patient test, notify the Provider for specific follow up testing.  If the control result is not within the acceptable range, refer to the Accu-Chek Inform System Operator’s Manual before proceeding with patient testing.

- If LO is displayed, the blood glucose result may be lower than the reading range of the meter.  If this contradicts the patient’s condition, perform a quality control check with glucose control solution and a new test strip.  If the control result is within the acceptable range, review proper testing procedure and repeat the blood glucose test with a new test strip.  If LO still appears on the patient test, notify the Provider for specific follow up testing.  If the control result is not within the acceptable range, refer to the Accu-Chek Inform System Operator’s Manual before proceeding with patient testing.

- If a Strip Defect error message appears on the display, the test strip may be damaged or the test was not performed correctly.  The test strip should be inserted into the meter prior to applying blood to the test strip.  If this display appears before blood is placed on the strip, remove the test strip and reinsert.  If the error display remains, repeat the test with a new strip.

- If the meter displays “Error 88-Bad Dose”, there is an incorrect amount of blood on the strip.  A second drop of blood may be applied to the test strip within 15 seconds of the first drop.  If more than 15 seconds have passed, the test result may be erroneous and you should discard the test strip and repeat the test.

- If the temperature-warning message appears on the display, the temperature is above or below the operating range of the test strips.  Move to a testing area that is between 57° and 104° F (14° and 40° C) and wait five minutes before repeating the test.  Do not artificially heat or cool the meter.

XVI. 	LIMITATIONS OF PROCEDURE

- Test strips give dependable test results if the following limitations are understood:

- Use only Accu-Chek Comfort Curve test strips (Cat. No. 2030365) or Accu-Chek Advantage H test strips (Cat. No. 787) for testing capillary, venous, neonatal (including cord), and arterial whole blood samples.

- Blood glucose determinations using venous and arterial blood specimens should be performed within 30 minutes of specimen collection to avoid glycolysis. Avoid air bubbles if dosing with pipettes. Air bubbles may cause erroneous results.

- For best results with arterial and venous blood, the following anticoagulants/preservatives are recommended: heparin and EDTA.

- Serum separator tubes and red-topped tubes are acceptable if blood is used immediately before the clotting process begins.

- Iodoacetate or fluoride/oxalate should not be used as a preservative.

- Caution should be taken to clear arterial lines before blood is drawn and dosed on the test strip.

- Caution is advised in the interpretation of neonate glucose values below 50 mg/dL. Follow the recommendations for follow-up care that have been set by your institution for critical glucose values in neonates.

- Do not use during xylose absorption testing.

- No effect was found at 20% to 65% hematocrit and glucose concentrations up to 200 mg/dL.

- At glucose concentrations above 200 mg/dL, low hematocrits (below 20%) may cause elevated results and high hematocrits (above 55%) may cause reduced results versus a whole blood reference.

- System measurement range is 10-600 mg/dL.

- The Accu-Chek Inform System has been tested at altitudes ranging from sea level to 10,150 feet.

- The following compounds, when determined to be in excess of their limitations, may produce elevated glucose results:

- Compound	Limitation

- Galactose	>10 mg/dL

- Maltose	>16 mg/dL

- Bilirubin (unconjugated)	>20 mg/dL

- Lipemic Samples	>5000 mg/dL

- Acetaminophen	>8 mg/dL

- Uric Acid:

- Hypoglycemic range	>10 mg/dL

- Euglycemic range	>12 mg/dL

- Hypergylcemic range	>16 mg/dL

- In situations of decreased peripheral blood flow, fingerstick blood testing may not be appropriate as it may not reflect the true physiological state. Examples would include but are not limited to: severe dehydration caused by diabetic ketoacidosis or the hyperglycemic hyperosmolar nonketotic state, hypotension, shock, or peripheral vascular disease.

XVII. 	OPERATOR CERTIFICATION/RECERTIFICATION

A certified Accu-Chek Inform System Instructor may be a Diabetes Clinical Specialist, QA Nurse, Roche Diagnostics Account Manager, or another designated individual who is trained in the use of the Accu-Chek Inform System.

Each operator will be appropriately in-serviced to perform blood glucose testing on the Accu-Chek Inform System.

Each operator must successfully complete a Knowledge Test and Skills Checklist to be certified to perform blood glucose testing on the Accu-Chek Inform System.

A certification roster and in-service training roster will be maintained for each nursing unit.

Each operator will be evaluated for competency, initially and annually thereafter.

- The operator will learn the policies and procedures for:

- Performing blood glucose testing on the Accu-Chek Inform System in this facility.

- Proper calibration of the Accu-Chek Inform System, patient testing, and documentation of results.

- Proper specimen collection and infection control in this facility.

- Proper test performance with the Accu-Chek Inform System

- Proper documentation of patient blood glucose results and follow-up of abnormal test results.

- Proper quality control testing.

- Proper preventive maintenance.

- Learning the meaning of display codes and proper troubleshooting procedures.

The operator will successfully complete an Accu-Chek Inform Knowledge Test to be certified for blood glucose testing with the Accu-Chek Inform System.

XVIII. REFERENCES

- Mor, Juan-R. and Guarnaccia, Rocco: Assay of Glucose Using an Electrochemical Enzymatic Sensor, Analytical Biochemistry, 79:319-328 (1977).

- D’Costa, E.J., Higgins, I.J., et al., Quinoprotein Glucose Dehydrogenase and its Application in an Amperometric Glucose Sensor, Biosensors, 2:71-87 (1986).

- Hauge, J.G., Glucose Dehydrogenase of Bacterium Anitratum: an Enzyme with a Novel Prosthetic Group, Journal of Biological Chemistry, 239:3630-3639 (1964).

- Tietz, N.W., Textbook of Clinical Chemistry, p. 2190 (1994).

- American Diabetes Association Position Statement, Diabetes Care, Vol. 19 (suppl. 1), p. S4 (1996).

- Atkin, S.H.; Dasmahapatra, A.; Jaker, M.A.; Chorost, M.I.; Reddy, S., Fingerstick Glucose Determination in Shock, Annals of Internal Medicine, 114:1020-1024 (1991).

- Sandler, M.; Low-Beer, T., Misleading Capillary Glucose Measurements, Practical Diabetes, 7:210 (1990).

- Wickham, N.W.R.; Achar, K.N.; Cove, D.H., Unreliability of Capillary Blood Glucose in Peripheral Vascular Disease, Practical Diabetes, 3:100 (1986).

- AccuChek Inform Operator Manual

SIGNATURE APPROVAL:

_____________________________		_______________________

Chief Executive Officer            	Date		Clinical Director                    	Date

_____________________________		______________________

Laboratory Director		    	Date	          Laboratory Supervisor           	Date

| Date | Name and Title | Signature |
| --- | --- | --- |