# SCOPE

# This policy applies to the clinical laboratories of Harrisburg Medical Center.

# PURPOSE

# The purpose of this policy is to provide guidelines for urinalysis using the CLINITEK Status Connect System with Siemens Clinitest Human Chorionic Gonadotropin (hCG) Pregnancy Test.

# PRINCIPLE

The Clinitest hCG Pregnancy Test is a chromatographic immunoassay (CIA) for the rapid determination of hCG in urine. The membrane is precoated with anti-beta hCG capture antibody on the test line region (T) and goat anti-mouse IgG on the control line region (C). During testing, the urine specimen is allowed to react with colloidal gold particles coated with anti-beta hCG monoclonal antibody. The mixture then chromatographically moves along the membrane by capillary action. For a positive or borderline result, a pink-colored line with a specific antibody-hCG-antibody-colloidal gold particle complex will form on the membrane in the test line region. A pink-colored line at the reference region (R), the area between the control line region and the test line region, has been adjusted to a level approximating 25 mlU/mL hCG. Absence of a pink-colored line in the test line region indicates a negative result. The appearance of a colored line in the control region and the reference region serves as verification that sufficient volume has been added and that proper flow has occurred.

# Clinical Application and Usefulness

The Clinitest hCG Pregnancy Test is for in vitro diagnostic use as a qualitative method in the rapid detection of human chorionic gonadotropin (hCG) in urine specimens. The test is utilized with the CLINITEK Status Connect System and is intended for near-patient (point-of-care) and centralized laboratory locations.

# Specimen Collection and Handling

Specimen Collection

|  |  | BIOHAZARD All products or objects that come in contact with human or animal body fluids should be handled, before and after cleaning, as if capable of transmitting infectious diseases. Wear facial protection, gloves, and protective clothing. |
| --- | --- | --- |

Urine is the recommended sample type for this assay.

- This assay requires 200 μL of sample for a single determination.

- Collect urine into a clean, dry container.

- Specimens collected at any time of day may be used.

Specimen Storage and Stability

- Refrigerate specimens at 2° - 8°C for up to 72 hours, if the testing is not performed immediately.

- If samples are refrigerated, bring them to room temperature before testing.

- Do not use samples that have been stored for longer than 72 hours.

Specimen Rejection Criteria

Do not accept the following:

- Specimens that have remained at room temperature for longer than two hours

- Specimens with urine preservatives

- Specimens that arrive in homemade containers (glass jars, pill bottles, etc.)

- Leaking specimen containers. Do not use urine that leaks into the bag. Only use urine that is in the specimen container.

- Specimens that do not have complete patient identification: Name and date of birth.

If an unacceptable specimen is received, note the reason for rejection on the test requisition and request a new, acceptable specimen from the patient. Acceptable collection container and collection instructions should be provided to the patient.

WARNING: Some urine specimens may have been collected during a critical procedure or by means of an invasive procedure; therefore, it is important to never dispose of an unacceptable specimen until the caregiver has been notified.

Specimen Referral Criteria

Physician request.

# Reagents

Storage and Stability

- Store Clinitest hCG Cassettes either refrigerated or at room temperature, 2 to 30°C.

- If refrigerated, bring the wrapped cassettes to room temperature before opening the protective pouch to avoid moisture condensation on the membrane

- Cassettes are stable for the duration of the shelf life when packaged in the protective pouch.

CAUTION:

- Do not use the Clinitest hCG Cassettes beyond the expiration date.

Reagent Special Preparation

No special preparation for Clinitest hCG Cassettes is required.

# Instrument Operation and System Description

The CLINITEK Status connector is intended for use with the CLINITEK Status+ Analyzer.  Together, the two units comprise the CLINITEK Status Connect System.  The connector also supports importing certain information using an optional bar code scanner.  It is for in vitro diagnostic use in:

- The detection of human Chorionic Gonadotropin (hCG) in urine samples, when Clinitest  hCG Cassette Tests are used.

The optical system consists of six light emitting diodes, a light guide, a mirror, a lens and a detector. Light from the LEDs travels along the light guide and is reflected off the calibration bar, strip or cassette onto the mirror. It is then directed through an aperture on the lens, from where it is focused onto the detector. The light intensity detected is converted into electrical impulses, which are processed by the instrument’s microprocessor and converted into clinically meaningful results.

# System Start-up and Maintenance

The test table insert and the test table should be kept clean if the analyzer is to operate properly.

WARNING: Do not autoclave the test table or test table insert.

WARNING: Care should be taken not to scratch the white calibration bar. If it is scratched or scuffed, obtain a new test table. Solvents of any kind must not be used to clean the bar.

Refer to your CLINITEK Status®  Analyzer Operator’s Manual for detailed cleaning and maintenance instructions.

System Start-up

The system is turned on by pressing the on/off button located at the front of the instrument. The analyzer automatically runs a system diagnostic check during which it performs a series of electronic, signal and memory checks, as well as ensures there is sufficient battery voltage to operate the instrument (if powered by batteries).  The Select Ready screen displays after system initialization.

Calibration

The CLINITEK Status Connect System performs a “self-test” and calibration each time it is turned on. In addition the analyzer performs an automatic calibration each time a test is run. The white calibration bar (on the test table) provides NIST traceable calibration.

# Quality Control (QC)

Each test includes two procedural controls, which indicate that sufficient sample was added for capillary flow to occur and the correct procedural technique was used. If the instrument does not detect the Reference (R) and Control (C) regions within two minutes after starting the test, an error is reported and the test must be repeated.  It is  recommended that quality control specimens be tested with each new lot of reagents, new shipment of reagents, and monthly for reagents that have been stored for more than 30 days.

QC Materials

Use a recommended quality control material.  Water should not be used as a negative control.

MAS Level 1 and Level 2 controls to monitor performance.

- Solutions are liquid and ready for use.

- Mix gently before using.

- Quality control values are programmed into each analyzer.

- Analyzer will indicate Pass or Fail at the end of the quality control printout.

- Indicate date opened and expiration date on each open quality control bottle.

- MAS controls expire 6 weeks after opened if at room temperature and 3 months after opened if refrigerated at 2 – 8 degrees °C

QC Frequency

Test known negative and positive controls whenever

- A new kit is first opened.

- Whenever you have questionable test results.

- With every new operator

- Every month.

Troubleshooting Out-of-Range QC Values

A QC run is acceptable when all values fall within the expected ranges.

If the hCG QC results do not fall within the defined ranges then the run is rejected, and you must take the following corrective action:

- Review these instructions to ensure that the assay was performed according to the procedures recommended by Siemens Healthcare Diagnostics

- Verify that the cassettes and control materials are not expired

- If necessary, rerun the quality control samples or contact Siemens Healthcare Diagnostics for more assistance

Enabling the Connector to Communicate with the Analyzer

- At the Select Ready screen, select Instrument Set up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select Connectivity.

- Choose Select.  The Connectivity screen 1 of 2 displays.

- To enable communication between the instrument and the connector, select Enabled.  To prevent communication, select Disabled.

- Select Previous.

- Select Done twice to return to the Select Ready screen.

# Managing the Operators List

To manage the Operators List, you can enable or disable authorized operator options, and add operator IDs.

Setting Operators IDs

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select Authorized Operator.

- Choose Select.  The Authorized Operator screen displays.

- To permit access only by authorized operators, select Enabled.  To allow all operators access to the system, select Disabled.

- For Enabled, see Adding Operator IDs below to add at least one operator.  If Disabled is selected, select Done 3 times to return to the Select Ready screen.

Adding Operator IDs

- At the Authorized Operator’s Screen, select Add Operator.

- Enter the new Operator ID.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Authorized Operator screen displays the Operator ID and which functions the operator can perform.

- To edit this Operator ID, select Edit.

- To edit which functions this Operator ID can access, select Edit.  The Authorized Operator-Operator access screen displays 1 of 2 displays.

- To allow this operator to run patient tests, select Enabled.  To prevent patient tests, select Disabled.

- To allow this operator to run QC tests, select Enabled.  To prevent QC tests, select Disabled.

- Select Next.  The Authorized Operator-Operator access screen 2 of 2 displays.

- To allow this operator to recall results, select Enabled.  To prevent recall results, select Disabled.

- To allow this operator to set up the instrument, select Enabled.  To prevent instrument setup, select Disabled.

- Select Done twice.  The Authorized Operator-Operators list screen displays.

- Select Exit.

- Select Done 3 times to return to the Select Ready screen.

# Scanning Patient Information

To require patient information entry only by scanning:

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Operator and Patient Information.

- Choose Select.  The Input Settings screen displays.

- Select Custom Set Up.

- Select Next.  The Custom Settings-Operator screen 1 of 5 displays.

- Select Next.  The Custom Settings-Patient Information screen 2 of 5 displays.

- Select Bar Code Reader Settings.  The Patient Information screen displays.

- To permit entry of patient information by either bar code scanner or manually, select Disabled.

- Select Done.  The Customs Settings Patient Information screen 2 of 5 displays.

- Select Next 4 times.

- Select Done twice to return to the Select Ready screen.

# Setting Up QC Cassette Tests

Accessing QC Cassette Set Up

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select QC Settings.

- Choose Select.  The QC Settings screen displays.

- For QC cassette test, select Set Up.  The QC Setting-Cassette test screen 1 of 3 displays.

# Setting Up QC Cassette Tests

- To allow QC prompts, select Enabled.

To prevent QC prompts, select Disabled.

- To require QC tests when they are due and not permit patient tests, select Required.

To permit patient tests even if a QC is due, select Reminder.

- Select Next.  The QC Settings-Cassette test screen 2 of 3 displays.

- To have the instrument determine QC pass of fail, select Instrument.

To have the operator determine QC pass of fail, select Operator.

- To permit patient tests if the QC test fails, select Yes.

To prevent patient tests if the QC test fails, select No.

- Select Next.  The QC Settings-Cassette test screen 3 of 3 displays.

NOTE:  If you selected to have the operator determine pass of fail in Step 4 above, Defining QC Cassette Test Controls is not necessary.

# Defining QC Cassette Test Controls

- To set the number of levels to complete a QC test select 1 or 2.

- For each control set in Step 1, select the corresponding Control Level, 1 or 2.

The Control Level 1-Cassette test screen displays.

- Select Enter name of control.  The Name of Control screen displays.

- Enter the name of this control using the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Control Level 1-Cassette test screen displays.

- To set the control level as positive, select Positive.

To set the control level as negative, select Negative.

- Select Done.  The QC Settings-Cassette test screen 3 of 3 displays.

- If necessary, repeat defining QC cassette test controls for the next control level.

- When all QC levels are defined, select Done.  The QC Settings screen displays.

# Setting QC Cassette Schedule

- For QC cassette test, select Interval.  The QC Settings-Set QC schedule for cassette test screen displays.

- To schedule tests by hours, select the Hours radio button and use the arrow keys to indicate the number of hours between QC tests.  Go to Step 7.  To schedule QC tests by days, select the Days radio button and use the arrow keys to indicate the number of days between QC tests.

- If you selected Days, select Set QC time.  The QC Settings-Set QC time cassette test screen displays.

- Select the 1, 2, or 3 radio button for the number of QC tests per day.

- Use the arrow keys to indicate time of day for each test.

- Select Set.  The QC Settings-Set QC schedule for cassette tests screen displays.

- Select Done.  The QC Settings screen displays.

- Select Done 3 times to return to the Select Ready screen.

# Running a QC Cassette Test

- At the Select Ready screen, select QC Test Cassette test due.  The QC Test screen displays.

- Select QC Cassette Test Required.

- If the instrument is set to determine pass/fail, the Control Lot screen displays.  Go to step 8.

- If the operator is set to determine the pass/fail, the Enter Control Name screen displays.

- Enter the name of the control.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Enter Control Level screen displays.

- Enter the control level.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Control Lot screen displays.

- Enter the control lot.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Control Expiration displays.

- Use the arrow keys to indicate the control lot expiration date.

- Select Enter.  The Cassette Lot screen displays.

- Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- The Cassette Expiration screen displays.

- Use the arrow keys to indicate the cassette lot expiration date.

- Select Enter. The Prepare Test screen displays.

- Select Start. The Results  screen displays.

- To print the results, select Print.

- Select Done.  The QC Test-Select PASS or FAIL screen displays.

- If the operator is set to pass or fail QC, To pass the test, select QC PASS.  To fail the test, select QC FAIL.

- Select DONE.  The QC Test-Results Summary screen displays.

- To repeat a failed QC test, select Repeat failed QC test.

- Select Done to return to the Select Ready screen.

Routine Analysis

The CLINITEK Status Connect System can be operated in two modes, Quick Test or Full Test. By default the analyzer will run in Quick Test mode.   Refer to your CLINITEK Status Analyzer Operators Manual Appendix D for a complete listing of other default settings.

Quick Test Mode

In the Quick Test mode testing begins as soon as Cassette Test is selected on the main menu screen.

WARNING: If refrigerated, bring the Clinitest hCG Cassette and patient sample to room temperature 20 - 30°C (68 - 86°F) prior to testing.

- At the main Select screen, touch Cassette Test. The Test Type screen will appear.

- Touch Clinitest hCG cassette. A Prepare Test screen will appear displaying the following two steps:

- Make sure the test table insert is facing up and in position for a cassette test.

- Remove the test cassette from the foil package and place the cassette on the test table.

- Touch START. Another Prepare Test screen will appear displaying the next two steps:

WARNING: Once you touch the START button you have eight-seconds to draw the urine sample into the pipette and add the urine sample into the well on the cassette.

  - Hold pipette at slight angle

  - Squeeze the upper bulb and draw enough sample into the pipette to fill the stem completely, with an overdrawn amount going into the reservoir (lower bulb)

  - Discharge the sample in the pipette stem into the sample well of the test cassette by squeezing the upper bulb in one squeeze. The excess fluid will remain in the reservoir.

WARNING: Do not push or pull the test table.

  - At the end of the eight-second countdown, the test table and cassette will automatically be pulled into the instrument.

- The analyzer will perform an automatic calibration and finish analyzing the sample.

WARNING: Do not move or bump the table while the instrument is calibrating.

- When analysis is complete, the Results screen will be displayed.

- Remove the used cassette and dispose of it according to your standard laboratory procedures.

- Touch Done to complete the test and return to the main Select screen.

Full Test Mode

To access the Full Test mode you must customize your settings.

- At the main Select screen, touch Instrument Set Up. The Choose Settings screen will appear.

- Use the down arrow to highlight Operator and Patient Information. Touch Select. The Input Settings screen will appear.

- Touch the circle in the Full Test box and touch Next.  An Input Settings confirmation screen will appear.

- Touch Done and you will return to the main Select screen.

For complete information on Instrument Set Up and customizing your test sequence, see your CLINITEK Status+ Analyzer Operator’s Manual.

In the Full Test mode you are prompted to enter an Operator ID, Patient Name and/or Patient ID prior to running a test.

- At the main Select screen, touch Cassette Test. The Operator ID screen will appear.

- If you were the last operator to enter an ID on the analyzer, touch Last Operator. The Patient Information screen will appear. Go to step 6.

- If you are a new operator, touch Enter New Operator ID. The Enter Operator ID screen will appear.

- Use the keypad to enter your ID using a maximum of 13 characters.

- Touch Enter. The Patient Information screen will appear.

- If the test sample is from a previous patient, touch Recall Patient. A list of up to 200 patients will appear on the screen.

- Use the up and down arrows to scroll through the patient names. Once the patient you are looking for has been highlighted, touch Select. The Test Type screen will appear. Go to step 13.

- If the test sample is from a new patient, touch Enter New Patient. The Enter Patient Name screen will appear.

- Use the keypad to enter the patient’s name using a maximum of 20 characters.

- Touch Enter. The Patient ID screen will appear.

- Use the keypad to enter the patient’s ID using a maximum of 13 characters.

- Touch Enter. The Test Type screen will appear.

WARNING: If refrigerated, bring the test cassette and patient sample to room temperature 20 to 30°C prior to testing.

- Touch Clinitest hCG Cassette. A Prepare Test screen will appear displaying the following two steps:

- Make sure the test table insert is facing up and in position for a cassette test.

- Remove the test cassette from the foil package and place the cassette on the test table.

- Touch START. Another Prepare Test screen will appear displaying the next two steps:

WARNING: Once you touch the START button you have eight seconds to draw the urine sample into the pipette and add the urine sample into the well on the cassette.

  - Hold pipette at slight angle

  - Squeeze the upper bulb and draw enough sample into the pipette to fill the stem completely, with an overdrawn amount going into the reservoir (lower bulb)

  - Discharge the sample in the pipette stem into the sample well of the test cassette by squeezing the upper bulb in one squeeze. The excess fluid will remain in the reservoir.

WARNING: Do not push or pull the test table.

  - At the end of the eight-second countdown, the test table and cassette will automatically be pulled into the instrument.

- The analyzer will perform an automatic calibration and finish analyzing the sample.

WARNING: Do not move or bump the table while the instrument is calibrating.

- When analysis is complete, the Results screen will be displayed.

- Remove the used cassette and dispose of it according to your standard laboratory procedures.

- Touch Done to complete the test and return to the main Select screen.

STAT Samples

STAT samples are run like all other samples. Follow directions as outlined above in either Quick Test Mode or Full Test Mode.

# Reporting Results

As with all in vitro diagnostic assays, each laboratory should determine its own reference range(s) for the diagnostic evaluation of patient results.

Reference Interval for Healthy Men and Healthy Non-pregnant Women

Healthy pregnant and non-pregnant women do not have detectable hCG levels using the Clinitest hCG Pregnancy Test.

Reference Interval for Pregnant Females

- 100 mIU/mL on the day of the first missed menstrual period

- Peak levels of hCG occur at 8 – 10 weeks after the last menstrual period

- Lower levels of hCG occur during the remainder of the pregnancy

- A rapid decrease and usually a return to normal in hCG levels occurs within days of delivery

Critical Values

None

Reporting Protocol for Critical Values

Not applicable

Units for Reporting Results

The system reports hCG results in mIU/mL. Units are user-defined in the system software.

Reportable Range

The Clinitest hCG Pregnancy Test detects urinary hCG concentrations greater than 25 mIU/mL (calibrated to the World Health Organization 3rd International Reference Preparation).

Acceptable Results

Patient test results are acceptable and may be reported when:

- The control line appears

- No errors messages are present on the analyzer.

- There is sufficient volume of urine to run the test.

- The sample is correctly identified at all stages of testing with the patient’s name and date of birth.

Corrective Action

Patient test results must be repeated and corrective action taken when:

- Control line fails to appear.

- Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure.

- Review the procedure and repeat the test with a new test device.

- If the problem persists, discontinue using the test kit immediately and contact the Laboratory Administrative Director or designee.

- Technical Support at (800) 637-3717.

- Fill out a Corrective Action Form.

# Procedure Notes

Urine Specimens

Specimens collected at any time of day may be used.

Positive Results

The instrument automatically determined that the Test (T) region intensity is equal to or more intense than a 25 mIU/mL urine sample and confirm that the Control (C) and Reference (R) regions met minimum intensity specifications.

Borderline Results

The result is indeterminate, repeat in 48 – 72 hours.

Negative Results

The instrument will automatically determine the Test (T) region is less intense than the 25 mIU/mL hCG concentration level that the device can detect, and confirms the Reference (R) and Control (C) regions meet minimum intensity specifications.  Negative test results in patients suspected to be pregnant should be retested with a sample obtained 48 to 72 hours later, or by performing a quantitative assay.

Invalid Results

The instrument will automatically determine if a procedural error or test reagent deterioration has occurred by confirming the Reference (R) and Control (C) regions meet minimum intensity requirements. If not, the user will be advised to repeat the test and to contact the Siemens Healthcare Technical Solutions Center if the problem persists.

Disposal

Dispose of hazardous or biologically contaminated materials according to the practices of your institution. Discard all materials in a safe and acceptable manner, and in compliance with all federal, state, and local requirements.

# Method Limitations

The test is not intended to detect conditions other than pregnancy. A number of conditions other than pregnancy, including trophoblastic disease and certain nontrophoblastic neoplasms, can cause elevated levels of hCG.

As is true with any diagnostic test, clinical diagnosis should not be based solely on a single test result.  Clinical diagnosis should incorporate all clinical and laboratory data.  Because of lag between conception and the appearance of hCG in urine (see Summary and Explanation of the Test), to exclude pregnancy with the highest degree of certainty, it is traditional to repeat the test on a fresh sample obtained 2–3 days after obtaining a “negative” result on the initial sample.

Patients on antibody therapies may obtain invalid results due to the presence of interfering antibodies in the medications.

The presence of heterophile antibodies or non-specific protein binding may cause false-positive results in sensitive immunoassays. If a qualitative interpretation is inconsistent with the clinical evidence, results should be confirmed by an alternative hCG method.

The specificity of the Clinitest hCG Pregnancy Test was determined from cross-reactivity studies with known amounts of human Luteinizing Hormone (hLH), human Follicle Stimulating Hormone (hFSH) and human Thyroid Stimulating Hormone (hTSH). All tests yielded negative results when used with 300 mIU/mL hLH, 1000 mIU/mL hFSH and 1000 µIU/mL hTSH.

For additional information on performance characteristics including cross-reactivity, see the product information in the Clinitest hCG product insert.

# Equipment and Supplies

- Clinitest hCG Cassette with disposable pipette

- CLINITEK Status Plus Connect System

- Specimen collection container

# References

- Siemens Healthcare Diagnostics Clinitest hCG Test Package Insert, 06878007, 2007-08

- Siemens Healthcare Diagnostics CLINITEK Status Analyzer Operator’s Manual, 132387, Rev. T, 2003-07

- Siemens Healthcare Diagnostics CLINITEK STATUS  Connect System Operator’s Manual, 135055, Rev A, 2009-12

- National Committee for Clinical Laboratory Standards (NCCLS). Clinical Laboratory Procedure Manuals, Third Edition (GP2-A3), 1996

- Siemens Healthcare Diagnostics CLINITEK Status + Analyzer Operators Manual, 135057, Rev. A, 2009-11.

*NCCLS is now known as: Clinical and Laboratory Standards Institute (CLSI).

Trademark Information

CLINITEK, CLINITEK Status and Clinitest hCG and all associated  marks are trademarks of Siemens Healthcare Diagnostics Inc.  All other trademarks and brands are the property of their respective owners.

Technical Assistance

Siemens Healthcare Technical Solutions Center: 1-877-229-3711

Siemens Healthcare Diagnostics

511 Benedict Avenue

Tarrytown, NY  10591-5097

© 2011 Siemens Healthcare Diagnostics