- PRINCIPLE

The Sysmex XS-1000i™ is a quantitative automated hematology analyzer for     in-vitro diagnostic use for determining 21 hematological parameters. Examination of the numeric and/or morphologic findings of the complete blood count are useful in diagnosis of such disease states as anemias, leukemias, allergic reactions, viral, bacterial, and parasitic infections. The Sysmex XS-1000i analyzer directly measures the WBC, RBC, HGB, HCT, PLT, NEUT#, LYMPH#, MONO#, EO#, and BASO #. The remaining parameters are calculated or derived:  MCV, MCH, MCHC, RDW-CV, RDW-SD, MPV, and differential percentages.

The Sysmex XS-1000i counts and sizes red blood cells (RBC) and platelets (PLT) using electronic resistance detection enhanced by hydrodynamic focusing. Hematocrit (HCT) is measured as the ratio of the total RBC volume to whole blood using cumulative pulse height detection. Hemoglobin (HGB) is converted to SLS-hemoglobin, and read photometrically.

WBC count and differential are evaluated using flow cytometry with a semiconductor laser utilizing scattered light and fluorescence to determine the differences in cell size, complexity and RNA/DNA content. The WBC differential channel classifies neutrophils (NEUT), lymphocytes (LYMPH), monocytes (MONO), eosinophils (EO), and basophils (BASO) by cellular complexity and nucleic acid content. The differential cell placement is then enhanced utilizing Adaptive Cluster Analysis.

- SPECIMEN

- Required specimen

- Whole blood anticoagulated with EDTA is preferred.

- Sodium Citrate may be used when EDTA platelet clumping or platelet satellitism is noted on the EDTA specimen. Use Sodium Citrate results for platelet counts and WBC counts. Multiply instrument PLT and WBC result by 1.11 to correct for anticoagulant dilution.

- II.	SPECIMEN (continued)

- Specimen volumes required:

1.	Optimal draw is a tube drawn to capacity. The collection             tube should be filled to a minimum of one-half full for                  acceptable results. EXCEPTION:  a 2.5 mL EDTA tube filled       less than one-           half full is unacceptable.

2.	A minimum of 500 μL whole blood is required for manual           mode analysis.

3.	A minimum of 1.0 mL whole blood is required for auto mode      analysis (XS-1000i with sampler).

4.	An EDTA Microtainer® (Becton Dickinson), designed for                  collection of 500 μL filled above the 250 μL line, is adequate           for testing in the open mode. A minimum of 90 μL is required           for aspiration using sample tube adapter.

- Unacceptable specimens, including those listed below, must be redrawn:

- Clotted samples or those containing clots, fibrin strands, or platelet clumps. All specimens will be checked visually by the operator for obvious clots prior to sampling by the analyzer.

- Grossly hemolyzed samples.

- Samples drawn above an IV.

- Characteristics that may affect test results:  lipemia, icterus, cold agglutinins, warm agglutinins, hemolysis, electrolyte imbalances, megakaryocytes and WBC fragments.

E.	Stored Specimen Stability

- Stored at 4C, EDTA blood samples with normal results may be analyzed up to 48 hours without significant loss of stability.

- Sample stability at room temperature is 24 hours. Samples stored at room temperature may exhibit an increase in MCV after 24 hours; this may be minimized by refrigeration.

- Allow refrigerated samples to come to room temperature and mix well before analysis.

F.	Do not place samples on a mechanical rocker. Constant rocking may cause PLT clumping and alter white cell membranes resulting in false interpretive messages.

WARNING:

All patient specimens should be considered potentially infectious and must be handled with precautions used for human blood, as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29 CFR part 1910.1030.  Follow specimen handling as outlined by laboratory safety policy.

Recommended:

Wear gloves, a lab coat and safety glasses if there is a risk of splash.

- SUPPLIES & REAGENTS

- Supplies

    - De-ionized water

    - Lint-free plastic lined lab wipes

    - Test tubes

    - CLOROX ULTRA™ bleach (Use when CELLCLEAN™ is                      indicated)

    - Sysmex reagents

    - Commercial controls:

a.	e-CHECK™, 3 levels (L,N,H) - 4 x 4.5 ml vials,  # 199-4004-	1

b.	e-CHECK (XS)™, 2 levels (N,H) - 5 x 1.5 ml vials,  # 199-5002-0

c.	e-CHECK (XS)™, 1 level (L) – 5 x 1.5 ml vials, cat. # 199-5001-0

    - Sysmex SCS-1000 whole blood calibrator

- Sysmex Reagents

- Four Sysmex reagents and bleach are used on the Sysmex XS-1000i.

- All reagents are stored at room temperature and are to be used within the manufacturer's expiration date on each container.

- Record date received and date opened on container. Record the lot number, expiration date and opened date on XS-1000i reagent log.

- All reagents are azide free, and intended for in vitro diagnostic use only; do not ingest.

Reagent	          Abbreviation	Open Expiration

Sysmex CELLPACK™	          EPK		60 days

Sysmex STROMATOLYSER-4DL™	          FFD		60 days

Sysmex STROMATOLYSER-4DS™	          FFS		90 days

Sysmex SULFOLYSER™ (500 mL container)     SLS		60 days

- DILUENTS

- Sysmex CELLPACK (EPK) is a whole blood diluent for use in the determination of hemoglobin and impedance counting and sizing of blood cells. Sysmex CELLPACK also forms a laminar sheath flow around the diluted sample for hydrodynamic focusing of the RBC and PLT.

- Sysmex CELLPACK Active Ingredients

Sodium Chloride			6.38 g/L

Boric Acid				1.00 g/L

Sodium Tetraborate			0.20 g/L

EDTA-2K				0.20 g/L

- Sysmex CELLPACK Storage

- Store at a controlled temperature of 5-30o C.

- If frozen, thaw, mix thoroughly, and allow bubbles to disperse before use.

- Sysmex CELLPACK is clear and colorless. If there are signs of contamination, instability or color change, do not use.

- Sysmex CELLPACK Stability

- Unopened, 18 months after the date of production marked on the box.

- Opened, Sysmex CELLPACK is stable for 60 days.

- Sysmex CELLPACK Hazard Risk

The OSHA Hazard Communication Standard of 29CFR part 1910.1200 requires MSDS documentation of ingredients which have been determined to be health hazards, comprise 1% or greater of the composition, are physical hazards, are capable of release to exceed permissible exposure limit/threshold limit values or have been identified as carcinogens. Sysmex CELLPACK does not have ingredients with those characteristics.

- LYSE REAGENTS

- Sysmex STROMATOLYSER-4DL (FFD) is the lysing reagent and        diluent for the enumeration of NEUT, LYMPH, MONO, EO, and            BASO after eliminating RBC stroma.

Sysmex STROMATOLYSER-4DL Active Ingredients

Nonionic surfactant			0.18%

Organic quaternary Ammonium salts      0.08%

Sysmex STROMATOLYSER-4DL Storage

- Store at a controlled temperature of 2-35o C.

- If frozen do not use.

- Sysmex STROMATOLYSER-4DL is a clear, odorless liquid.  If        there are signs of contamination, instability or color change, do       not use.

Sysmex STROMATOLYSER-4DL Stability

- Unopened, 12 months from date of production marked on box.

- Opened, Sysmex STROMATOLYSER-4DL is stable for 60 days.

Sysmex STROMATOLYSER-4DL Hazard Risk

The OSHA Hazard Communication Standard of 29CFR part 1910.1200 requires MSDS documentation of ingredients which have been determined to be health hazards, comprise 1% or greater of the composition, are physical hazards, are capable of release to exceed permissible exposure limit/threshold limit values or have been identified as carcinogens. Sysmex STROMATOLYSER-4DL does not have ingredients with those characteristics.

CAUTION:

Do not ingest. Avoid skin and eyes contact. Flush with plenty of water immediately. Consult with a physician in case of ingestion and/or eyes contact.

- Sysmex SULFOLYSER (SLS) is the RBC lysing reagent that                     releases the hemoglobin to be measured by SLS hemoglobin                    method.

Sysmex SULFOLYSER Active Ingredients

Sodium Lauryl Sulfate			1.7 g/L

Sysmex SULFOLYSER Storage

- Store at controlled temperature of 2-30o C.

- If frozen, may form a white cloudy precipitate. Thaw and warm       reagent in a 30o C water bath to dissolve ingredients completely      and mix thoroughly before use.

- Sysmex SULFOLYSER is a clear, odorless liquid. If there are          signs of contamination, instability or color change, do not use.

Sysmex SULFOLYSER Stability

- Unopened, 12 months from the date of production marked on         the box.

- Opened, Sysmex SULFOLYSER (500 mL container) is stable         for 60 days.

- LYSE REAGENTS (continued)

2.	Sysmex SULFOLYSER (SLS)

Sysmex SULFOLYSER Health Risk

The OSHA Hazard Communication Standard of 29CFR part 1910.1200 requires MSDS documentation of ingredients which have been determined to be health hazards, comprise 1% or greater of the composition, are physical hazards, are capable of release to exceed permissible exposure limit/threshold limit values or have been identified as carcinogens. Sysmex SULFOLYSER does not have ingredients with those characteristics.

CAUTION:

Do not ingest. Avoid skin and eyes contact. Flush with plenty of water immediately. Consult with a physician in case of ingestion and/or eyes contact.

- STAINING REAGENTS

- Sysmex STROMATOLYSER-4DS (FFS) is used to stain leukocytes in       diluted, lysed blood samples for the determination of the 5-part                  differential including NEUT, LYMPH, MONO, EO and BASO.

Sysmex STROMATOLYSER-4DS Active Ingredients:

Ethylene Glycol			96.9%

Methanol				3.0%

Polymethine dye			0.002%

Sysmex STROMATOLYSER-4DS Storage

- Store at controlled temperature of 2-35 Co.

- If frozen, do not use.

- Sysmex STROMATOLYSER -4DS is a blue viscous liquid with a          faint odor. If there are signs of contamination, instability or color           change, do not use.

Sysmex STROMATOLYSER-4DS Stability

- Unopened, 12 months from the date of production marked on pouch.

- Opened Sysmex STROMATOLYSER-4DS is stable for 90 days.

Sysmex STROMATOLYSER-4DS Health Risk

Health hazard reflects unusual exposures, via ingestion. This reagent is relatively safe when handled while wearing gloves and when enclosed in the analyzer hydraulic system.  Skin contact may cause discoloration. Methanol component is a defatting agent, which may dry and crack the skin.  Prolonged and repeated skin contact may cause irritation.

CAUTION:

Do not ingest. Avoid skin and eyes contact. In the case of skin contact, wash immediately with plenty of soap and water. In case of contact with the eyes, rinse immediately with water or normal saline. Consult with a physician in case of ingestion and/or eyes contact.

Recommended:

Wear gloves, a lab coat and safety glasses for protection.

- CLOROX® ULTRA

- CLOROX ULTRA bleach is recommended for use in cleaning of the Sysmex XS-1000i analyzer whenever CELLCLEAN is indicated.

CLOROX ULTRA Ingredients

Sodium Hypochlorite			6.0%

CLOROX ULTRA Storage

Stable under normal use and storage conditions

WARNING:

CLOROX ULTRA:  Avoid acidification or contact with ammonia containing products, which can generate hazardous chlorine gas.

CLOROX ULTRA Health Risk

Respiratory irritant if mist or vapor is inhaled, nausea and vomiting if ingested. May irritate skin. Contact with eyes can cause severe, but temporary injury to the eye.

WARNING:

CLOROX ULTRA contains a strong oxidizing agent. Causes substantial but temporary eye injury. May irritate skin. May cause nausea and vomiting if ingested. Exposure to vapor or mist may irritate nose, throat and lungs.

Recommended:

Wear gloves, a lab coat and safety glasses for protection.

- COMMERCIAL CONTROL MATERIALS

Sysmex e-CHECK and Sysmex e-CHECK (XS) are whole blood commercial controls for use with the Sysmex XS-1000i hematology analyzer.

Sysmex e-CHECK and e-CHECK (XS) Ingredients (formulation)

Sysmex e-CHECK and e-CHECK (XS) consist of human red and white blood cells with a platelet component suspended in fluid medium. Each vial of Sysmex e-CHECK contains 4.5 mL of control material. Each vial of Sysmex e-CHECK (XS) contains 1.5 mL of control material.

Sysmex e -CHECK and e -CHECK (XS) Storage

Store vials as packaged, at 2-8o C.

Do not freeze or expose to excessive heat.

Sysmex e-CHECK and e-CHECK (XS) Stability

Unopened and properly stored, Sysmex e-CHECK and e-CHECK (XS) are stable until the expiration date stated on the vial.

- Open vial stability is 7 days for Sysmex e-CHECK and 14 days for Sysmex     e-CHECK (XS) when promptly refrigerated after each use.

- Record the date on each vial upon opening.

- Heat or freezing can damage Sysmex e-CHECK and e-CHECK (XS) without gross visible changes. Moderate hemolysis can be normal. Deterioration is suspected when the mean of the control results is not within the assay expected ranges after appropriate troubleshooting.

- If deterioration is suspected, call the Sysmex Technical Assistance Center at 1-888-879-7639 (1-888-8SYSMEX).

WARNING:	Potentially Infectious Material

The human blood used in Sysmex e-CHECK / e-CHECK (XS) is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens.  Sysmex e-CHECK / e-CHECK (XS) should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29 CFR, 1910.1030.

- Sysmex SCS-1000 Hematology Calibrator

Sysmex SCS-1000 Calibrator is a secondary whole blood calibrator for use with the Sysmex XS-1000i hematology analyzer. Assay values for primary parameters are traceable to reference methods.

Sysmex SCS-1000 Ingredients (formulation)

Sysmex SCS-1000 consists of human red and white blood cells with a platelet component suspended in fluid medium. Each vial contains 2.0 mL of calibrator material.

Sysmex SCS-1000 Storage

Store vials in the upright position, at 2-8 o C. Do not freeze or expose to excessive heat.

- Sysmex SCS-1000 Hematology Calibrator (continued)

Sysmex SCS-1000 Stability

- Unopened and properly stored, Sysmex SCS-1000 is stable until the expiration date stated on the vial.

- Open vial stability is 4 hours.

- Storage outside of 2-8o C can damage Sysmex SCS-1000 causing deterioration that risks inaccurate calibration. If deterioration is suspected, call the Sysmex Technical Assistance Center at 1-888-879-7639 (1-888-8SYSMEX).

- Use of the product at environmental temperatures that exceed 86o F (30 o C) can reduce calibration accuracy.

WARNING:	Potentially Infectious Material

The human blood used in Sysmex SCS-1000 is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens.  Sysmex SCS-1000 should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29 CFR, 1910.1030.

Reagent Replacement

Recommended:

Wear clean gloves, a lab coat and safety glasses for protection, and to prevent contamination of the reagent.

- When a reagent container is empty, the alarm sounds and the Help Dialog box displays on the IPU. “Replace Container” message with the name and abbreviation for the reagent to be replaced is displayed.

- Click [RESET ALARM] or press [F1] to silence the alarm. The Action area of the dialog box displays instructions for reagent replacement.

- Click [Execute] to open Reagent Replacement screen.

- Once open, the Reagent Replacement screen will display “Replace” next to the reagent to be replaced.

- Scan the long (lower) barcode label on the container using the handheld bar code reader. Updated information is displayed for lot number, expiration date and volume.

- Reagent information may also be entered manually.

- Enter the lot number and expiration date found on the reagent label.

- Enter the number of days for “Opened Expiration”. Open dating for all reagents is 60 days.

- Open the new reagent container. Using clean technique to avoid contamination, remove the line from the empty container and insert it directly into the new container. Tighten the cap on the new container.

- Click [Execute] to initiate priming.

- Reagent will be aspirated to satisfy the sensor.

- The Reagent Log will be updated once priming is complete.

- To view the Reagent Log, click [Menu] or press [F4], click on “Controller” icon, and “Reagent Log” icon.

- Double click the line for the reagent that was replaced to open the Comment box. A comment (e.g. Tech initials) may be added.

- IV.	CALIBRATION

Initial calibration is performed during installation. Calibration compensates for any bias inherent to the pneumatic, hydraulic, and electrical system that may affect the accuracy of results. Calibrators traceable to reference methods are used in the calibration of the instrument. WBC differential parameters are calibrated in the factory prior to shipment, and verified by the field service representative upon installation.

The laboratory must verify calibration every six months or on an "as needed" basis to ensure accuracy of the system. Calibration is also required if one or more of the following occur:

- Critical parts are replaced such as manometers, apertures or detector circuit boards.

- Controls show an unusual trend or are outside of acceptable limits and cannot be corrected by maintenance or troubleshooting.

- When advised by Sysmex Field Service Representative.

Calibration verification may include review and documentation of acceptable performance on all three levels of commercial control, and Xm QC data, proficiency testing results and patient control testing results. The operator may calibrate the following parameters: WBC, RBC, HGB, HCT and PLT.

Before calibration, ensure that the Sysmex XS-1000i is both clean and precise.

- IV.	CALIBRATION (continued)

- Precision Check

- Perform routine daily and weekly maintenance on the instrument, and perform a background count to ensure counts are within acceptable limits.

- Verify that there is sufficient volume of all reagents. Precision and Calibration procedures will be aborted if the Sysmex XS-1000i runs out of reagent.

- Obtain a sample of fresh normal whole blood. Do not use commercial controls or calibrators for precision. The blood donor specimen should:

- be free from medication, and interfering substances such as lipemia, icterus, platelet clumps, hemolysis, etc.

- have morphologically and numerically normal CBC.

- be drawn in EDTA anticoagulant tube using proper collection technique.

- be a minimum of 2 mL of sample.

- On the IPU, click [Menu] or press [F4].

- Click on the “Controller” icon in the Menu window.

- Double click on the “Precision Check” icon on the Controller menu.

- Attach appropriate tube adapter to the sample position.

- Place the well-mixed sample tube in the tube position.

- Analyze the sample 11 times in the Manual Mode, mixing gently before each analysis. After each analysis, the results display on the IPU screen, and the cursor moves to the next line.

- The mean, SD, and CV% are calculated on the last 10 analyses.         The CV% values are displayed in red if they exceed the Precision        Limit %. If the CV% exceeds the precision Limits, click [CANCEL].

- If all results are within Precision Limits, click [OK] on the Precision       Check screen to save the data and close the Precision Check               window.

- To print a hard copy of the Precision results, click on “Precision            Check List” on the “Controller” screen, then click [File] on the               Menu Bar and click [Print].

- IV.	CALIBRATION (continued)

- Calibration Check

- Prepare the Sysmex SCS-1000 calibrator according to the product       insert.

WARNING:	Potentially Infectious Material

The human blood used in Sysmex SCS-1000 is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens. Sysmex SCS-1000 should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29 CFR, 1910.1030.

- Click [Menu] or press [F4].

- Click the “Controller” icon on the Main menu window.

- Double-click the “Calibrator Calibration” icon on the Controller              menu.

- Click [ASSAY TARGET], then click in the parameter fields and use       the numeric keys to enter the target values for each parameter             from the Sysmex SCS-1000 assay sheet. Click in the next field            after each entry. Click [OK] when all values have been entered.

- Attach the appropriate sample tube adapter.

- Remove the cap from the calibrator vial and perform 6 consecutive       analyses in the sampling mode.  Do not mix in between 6                     analyses. Wait until the 6th analysis is complete, and then review          results.

- After 6 analyses, the last 5 analyses are used to calculate the              Range Values and Delta percents. The parameter values are               highlighted in red if they are out of range.

- Calibration cannot be performed if the Range Value exceeds the Max Range Value, or if the Delta % exceeds the Service limit. Take corrective action. See Troubleshooting section, in the Sysmex XS-Series Instructions for Use manual.

- If the Range Values and Delta Percents are within acceptable limits and require calibration, click [ACCEPT]. The parameters are displayed with a check in the check box beside the new Compensation Rates. Parameters that are not within acceptable limits cannot be selected for calibration.

- Click [OK] to execute the calibration or [Cancel] if calibration is not desired or necessary. The display returns to the Controller menu.

- Calibration Check (continued)

- Using another fresh vial of the calibrator, verify the                      calibration by repeating the Calibration Check procedure.           The analyzed values should all be within acceptable limits.         Do not execute calibration; exit the Calibration function.

- Print the Calibration History.

- Click [Menu] or press [F4].

- Click the “Controller” icon.

- Click the “Calibration History” icon.

- Click on the calibration date in the Calibration History list.

- To print a hard copy of the Calibration results, click on Calibration Check List on the “Controller” screen, then click [File] on the Menu Bar and click [Print].

- Following calibration, analyze commercial controls. Adjust          target values if necessary.

- V.	QUALITY CONTROL

A.	At least two levels of control material are analyzed and within range daily or when a test is requested.  In addition, these controls should be run with each new calibration, with each new reagent, and after specific maintenance or troubleshooting procedures.

B.	The following controls should be prepared and used in accordance with the package inserts.

a.	Sysmex e-CHECK (4.5 ml vials) and e-CHECK (XS) (1.5 ml        vials) instructions for use.

1.	Remove Sysmex e-CHECK / e-CHECK (XS) vials from               refrigerator and allow them to come to room temperature           (18-25o C), for approximately 15 minutes.

2.	Mix vials by gentle end to end inversion until the cell button        in the bottom of the vial is completely suspended.

WARNING:	Potentially Infectious Material

The human blood used in Sysmex e-CHECK / e-CHECK (XS) is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques. However, no current tests can assure the absence of these pathogens. Sysmex e-CHECK / e-CHECK (XS) should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29 CFR, 1910.1030.

C.	Entering Lot Information for a New Lot of Controls

- Click [QC Files] or press [F5].

- Click on file number (1-20).

- Click [Input] or press [F9]. The QC dialog box displays.

- Click on [▼] beside “Material” to select the control level.

- Click in the field for “LOT NO.”. Using the keyboard, enter the lot number from the vial or assay sheet.

- Click on [▼] beside the “Exp. Day” to display the calendar. Click on the expiration date in the calendar to display it in the field.

- Click [OK] to update the lot information.

- Repeat steps 2-7 to enter lot information for the other levels of control.

- QUALITY CONTROL (continued)

D.	QC Analysis

- QC analysis in the Manual mode. (XS-1000i)

- Click [Manual] or press [F2]. The Manual Sample No. dialog box opens.

- Enter the lot number using one of the following methods:

- Use the handheld barcode reader to scan the bar code from the vial.

- Use the keyboard to type the lot number.

Note:	“QC” must be upper case followed by a hyphen.

- Click the [QC] button on the Manual Sample No. dialog box. Click on the QC level to be analyzed. Click on [OK].

- Press the Open/Close switch to open the sample position.

- Attach the appropriate tube adapter for the vial to be analyzed.

- Place the well-mixed vial in the adapter and press the “Start” switch on the Main Unit.

- The QC results display when analysis is complete. Results which are outside of acceptable limits display with red background. Click [ACCEPT] to accept the results, or [CANCEL] to reject.

- Follow steps a-f to analyze other levels of controls.

- QC analysis in the Manual mode. (XS-1000i with Sampler)

  - Click [Manual] or press [F2]. The Manual Sample No. dialog box opens.

- Enter the lot number by one of the following methods:

- Use the handheld barcode reader to scan the barcode from the vial.

- Use the keyboard to type the lot number.

Note:	“QC” must be upper case followed by a hyphen.

    - Click the [QC] button on the Manual Sample No. dialog box. Click on the QC level to be analyzed. Click on [OK].

- Attach the appropriate tube adapter for the vial to be analyzed.

- Place the well-mixed vial in the adapter and press the “Start” switch on the right side inside the sampler.

- The QC results display when analysis is complete.  Results which are outside of acceptable limits display with red background. Click [ACCEPT] to accept the results, or [CANCEL] to reject.

- Follow steps a-e to analyze other levels of controls.

- QC analysis in the Sampler mode with Barcodes. (XS-1000i with sampler)

- Place a Sysmex rack in the rack position of the Sampler with the notch on the rack to the right.

- Place the well-mixed control vials in positions 8, 9, 10 of the Sysmex rack.

- Attach the appropriate sample tube adapter.

- Close the Sampler cover.

- Click [Sampler] or press [F3].

- The Sampler Sample No. dialog box displays. Click on the starting position for the rack and tube position in which the vials have been placed.

- Press the sampler Start switch on the left side of the Main Unit.

A dialog box displays when analysis is complete

- V.	QUALITY CONTROL (continued)

E.	Review QC

      - Radar Charts

  - On the IPU, click [QC Files] or press [F5].

  - Click on the file to be reviewed.

  - The Radar Chart displays beside the selected file.

  - Results of the most recent analysis display in blue. Date and time of analysis are displayed in the Analysis Date column.

  - Results outside of acceptable limits are displayed with a red “X” and the parameter name is displayed with a red background.

      - L-J Charts

      - On the IPU, click [QC Files] or press [F5].

      - Double click on the file to be reviewed. The Levy-Jennings chart will be displayed.

      - Results outside of acceptable limits are displayed with a red “X” on the L-J chart. The parameter name and the result value will be displayed with a red background.

      - Scroll through the screens to view all parameters by using the scroll bar on the right of the screen or press the down arrow.

Verify that all parameters fall within established limits or within the package assay range. Contact the Sysmex Technical Assistance Center to investigate any suspected control product failure.

- V.	QUALITY CONTROL (continued)

F.	Auto-Set Targets

Parallel test new controls by analyzing each level of control a minimum of twice a day for 5 days prior to expiration of the previous lot. After a minimum of 10 data points are accumulated, auto-set the targets.

- On the IPU, click [QC Files] or press [F5].

- Double-click on the appropriate file number (1-20) to open the L-J chart.

- Set the range of data for target calculation by clicking on the green cursor and dragging to the left to include all points to be used in calculation, or press [Ctrl] and [A] to select all data.

- Click [Input] or press [F9].

- Click on RBC in the Target/Limit window and click and drag to include all parameters.

- Click [Auto Setting].

- Click the “Target” check box to select.  Click [OK].

- The targets will be auto-set to the mean of the range of points selected in the QC file and will display in the Target/Limit window.

- Click [OK] on the “Input Lot Information” screen.

- Repeat steps 2-9 for each level of control.

G.	Corrective Action for Out of Range QC Results

Corrective Actions:

1.	Check the lot number of the control and the reference range and rerun the control.

2	If the new value is within control limits, proceed with patient specimens.  Retain both control values in the Quality Control File.

3	If the new value is also out-of-control, open new controls and rerun using fresh control material. As long as two levels are within range, patient testing may begin. The third level will be monitored and after two consecutive days  if the value remains out of control, contact Technical Support.

H.	Recording and Storage of QC Data

Results of each quality control run are stored by the analyzer. Hard copies are printed monthly and kept in SYSMEX QC binder located in the department.

1.	Printing QC Data

- On the IPU, click [QC Charts] or press [F5].

- Double-click the appropriate file number (1-20) to open the L-J chart.

- Set the range of data to be printed by clicking on the green cursor and dragging to the left to include all points, or press [Ctrl] and [A] to select all data.

- Click [Out] or press [F12].

- Click on either “Report (GP)” for a graphic printout of the L-J Chart of data, statistics, targets and limits or click on “Report (LP)” for line print of data, statistics, targets and limits.

- Repeat steps b-e to print each level of QC data.

2.	Saving QC Data on Floppy Disk

- On the IPU, click [QC Files] or press [F5].

- Click the appropriate file number (1-20).

- Click on “Record” from the menu bar and select “Backup”.

- Insert 3-1/2“ floppy disk when prompted or USB flash-drive into USB port.

- The Save dialog box appears. Enter a file name for the QC file. All file information and data will be saved under this file name.

- Click [SAVE]. Click [CANCEL] if you do not wish to save the file.

3.	Saving QC Data on a removable drive (USB flash)

- Select [QC Files] from Main Menu or [F5].

- Highlight appropriate file (1 – 20) to save to Insight format.

- Select Insight icon or [F12].

- Insert USB Flash Drive on lower right side of IPU into USB port.

- D:drive window opens.

- Verify the lot number and level in the “File Name” box, click [SAVE].

- Repeat steps b – g for remaining levels.

I.	Establishing Historical LIMIT %’s for Commercial Controls

For each parameter of each level of control, an acceptable range around the mean must be established. This range, called the ‘LIMIT %’ is based on historical performance of the commercial control material when the instrument is in good working condition.

Historical LIMIT %s are established using three different lots of Sysmex       e-CHECK / e-CHECK (XS) controls (over a 6 month period for the 84-day dated lot). Interim Limit %s suggested by Sysmex are used prior to establishing the analyzer-specific limits during the evaluation period. Once three lots of QC data are collected, the CV% for each parameter is averaged. To establish a 3CV% limit, multiply the average CV's x 3. These historical limits are manually entered for the LIMIT % in each file for each level of control and are used for all subsequent lots of controls. These limits should provide acceptable error detection with a low probability of false rejection, and need not be re-established.

1.	Entering QC Limit %

- On the IPU, click [QC Files] or press [F5].

- Click on the appropriate file (1-20).

- Click [Input] or press [F9].

- Each parameter for the level selected will be displayed in the “Manual Setting” window on the right of the “Input Lot Information” screen.

- Manually enter the LIMIT % for the first parameter in the “Limit Range %” field and press [ENTER] on the keyboard 2 times to move to the limit for the next parameter. Repeat for all parameters of the QC level.

- Click [OK] to save and return to QC file list.

- Repeat steps b-f for other levels.

J.	Insight™ Quality Assurance Program (QAP)

The Insight account number is: ___23374__________.

The Sysmex XS-1000i serial # is _____69973_______.

Lab Tech or Supervisor is responsible for saving the data to a floppy disk and uploading the data to Sysmex Insight on the Sysmex Web Site. Data must be submitted no later than 3 days after the due date on the Sysmex e-CHECK / e-CHECK (XS) assay sheet.

- Each lot has 2 data submission dates, approximately every        30 days for the 56-day dated product. The submission dates      are printed on the QC Assay sheet.

2.	Saving Sysmex Insight data to floppy disk.

- On the IPU, click [QC Files] or press [F5].

- Click on the appropriate file (1-20).

- Click [Insight] or press [F12].

- Insert 3-1/2” floppy disk into disk drive.

- Verify lot number and level in the “File Name” box. Click [Save].

- Repeat steps b-e for remaining levels.

- Submit the data by uploading to the Sysmex Insight program on the Sysmex Web Site:  www.sysmex.com. Additional information may be obtained by referring to the Insight Participant Overview Guide Report Guide provided by the Sysmex Data Center upon enrollment in the Insight program. Contact the Sysmex Data Center for questions regarding data submission by telephone or    e-mail.

Sysmex Insight phone:	847-996-4563

Sysmex Insight Email:	insight@sysmex.com

Note:	Keep the diskette containing the Sysmex Insight files until the Insight report is received.

- VI.	OPERATING PROCEDURE

A.	Start-Up Procedure

- Check pneumatic trap for fluid and drain if necessary.

- Check reagent boxes for sufficient run volume.

- Check printer paper supply.

- Power Up Sequence

- Press power switch on Information Processing Unit (IPU). IPU log on must be done before powering up the Main Unit.

- Sysmex XS-1000i program log-on box displays. Log on the IPU with your User Name (___XS___) and press [ENTER]. No password is required. Substitute your user name and password if they have been assigned.

- Press the power switch on right side of the Main Unit.

- The instrument automatically performs self check on the:

Microprocessor 		Mechanical parts

Temperatures 			Background counts

| Press the power switch on the printer. - Sysmex XS-1000i Acceptable Background Counts |  |
| --- | --- |
| Parameters | Acceptable Limit |
| RBC | 0.02 x 106/ μL |
| HGB | 0.1 g/dL |
| PLT | 10 x 103/ μL |
| WBC - C | 0.1 x 103/ μL |
| WBC - D | 0.1 x 103/ μL |

- Analyze Quality Control.

B.	Patient Sample Processing

1.	Manual Mode – (20 μL aspirated sample volume) minimum        of 50 μL in tube or 90 μL in a micro-sample container

- On the IPU, click [Manual] or press [F2].

- Enter the specimen number (alpha or numeric characters) using the keyboard or using the handheld bar code reader.

- Click on CBC or CBC+Diff if this information is not being provided by the Host Computer.

- Click [OK].

- Attach appropriate sample tube adapter.

- Mix the patient sample 10 times by end-to-end inversion.

- Place sample in sample tube adapter. It is not necessary to remove the cap except when using non-pierceable micro-sample containers.

WARNING:

Potential biohazard exposure when handling open patient specimens. 	Follow standard precautions outlined by laboratory safety guidelines.

Recommended:

Wear gloves, a lab coat and safety glasses. Use plastic lined gauze when opening specimen tubes.

- Press Start switch. (Located above the sample tube position on the Main Unit of the XS-1000i without Sampler; inside the sampler cover on the XS-1000i with Sampler.)

- When Ready LED is lit green, repeat steps a-i for each additional sample.

2.	Sampler Mode with Bar Codes – XS-1000i with Sampler

(20 μL aspirated sample volume). A minimum of 1.0 cc of          blood is required in the tube for the sampler mode.

- Place a Sysmex rack in a rack position of the Sampler with the notch on the rack to the right.

- Place up to 2 racks at one time (up to 20 samples).

- Place bar coded specimens in the rack. Ensure that labels are smooth with no loose edges.

- Attach the appropriate sample tube adapter.

- Close the Sampler cover.

- On the IPU, click [Sampler] or press [F3]. The Sample number dialog box displays.

- Click on the starting position for the rack and tube position in which the tubes have been placed. Press [OK].

- Press sampler Start switch on the left side of the Main Unit.

- The Sysmex XS-1000i automatically mixes the sample 10 times, aspirates, and analyzes the sample according to the bar code discrete order if bidirectional interface is used. Results print as they are completed if auto-output is selected.

- A dialog box displays when analysis is complete.

- VI.	OPERATING PROCEDURE (continued)

  - Shut Down – Perform every 24 hours

Cleans the detector and dilution lines

- Click [Menu] or press [F4].

- Double click on the “Shutdown” icon.

- After 2 minutes, a dialog box on the IPU displays “Please power           off the analyzer”.

Note:	To continue analysis, click [RESTART] from XS-1000i Shutdown screen. After auto-rinse and background check is completed,            XS-1000i is “Ready”.

- Power off the Sysmex XS Main Unit.

- To power off the IPU, click [File] from the menu bar, then click [Exit].

- Dialog box displays ”Do you really want to Log off?” Click [OK].

- Click on Start button at the bottom of Windows desktop.

- Click [Shut Down].

Note:	The Restart key displays on this dialog box. If desired, click [RESTART] to begin IPU start up process.

- The system displays: “Please wait while the system writes unsaved       data to the disk”.

- Record on Maintenance Log.

- VII.	MAINTENANCE

This section includes written procedures for performing monthly and 30,000 cycle maintenance. Refer to Sysmex XS-Series Instructions for Use manual, Chapter 9 for detailed, illustrated procedures.

WARNING for All Maintenance:

CLOROX ULTRA contains a strong oxidizing agent. Causes substantial but temporary eye injury. May irritate skin. May cause nausea and vomiting if ingested. Exposure to vapor or mist may irritate nose, throat and lungs. If contact with eyes, flush with copious amounts of water. Also, potential biohazard exposure when performing maintenance on the Sysmex XS-1000i.

Recommended:

Wear gloves, lab coat and safety glasses for protection.

Important Note For All Maintenance

CLOROX ULTRA is a 6% (by volume) Sodium Hypochlorite solution. The Sysmex XS-Series Instructions for Use manual recommends using a 5% Sodium Hypochlorite solution as a stock solution for maintenance procedures. To make a liter of 5% stock solution from CLOROX ULTRA, use the formula below.

One Liter of 5% from 6% Sodium Hypochlorite:

(Conc. 1) x (Vol. 1) = (Conc. 2) x (Vol. 2)

so (6%) x (Vol 1) = (5.00%) x ( 1 liter, 1000 mL)

V1 = 5.00/6 x 1000 mL

V1 = 833 mL bleach and 167 mL of distilled water will make one liter of 5% Sodium Hypochlorite solution.  Store stock 5% bleach in a dark place to prevent solution degradation from exposure to light.

A.	Monthly:

      - Rinse Sequence – to be performed monthly or every 1200 cycles. Cleans the optical detector block

(Sysmex XS-Series Instructions for Use manual)

- Click [Menu] or press [F4].

- Click “Controller” icon on the menu screen.

- Click “Maintenance” icon. The maintenance screen displays.

- Click on “Monthly Rinse”. The Monthly Rinse dialog box displays.

- Attach the appropriate sample tube adapter.

- Place a tube of 5% Sodium hypochlorite solution (CELLCLEAN) in the tube adapter.

- Press the Start switch. (Located above the sample tube position on the Main Unit of the XS-1000i without Sampler; inside the sampler cover on the XS-1000i with Sampler) to initiate the cleaning.

- Record on Maintenance Log.

- The Power Off dialog box displays when the process is complete.

- Press [Restart] to resume operation or power off the instrument.

- B.	As-needed Maintenance

1.	Waste Chamber Cleaning

(Sysmex XS-Series Instruction for Use manual)

- Click [Menu] or press [F4].

- Click “Controller” icon on the menu screen.

- Click “Maintenance” icon. The maintenance screen displays.

- Click on “Rinse Waste”. The Rinse Waste dialog box displays.

- Attach the appropriate sample tube adapter.

- Place a tube of 5% Sodium Hypochlorite solution (CELLCLEAN) in the tube adapter.

- Press the Start switch. (Located above the sample tube position on the Main Unit of the XS-1000i without Sampler; inside the sampler cover on the XS-1000i with Sampler) to initiate the cleaning.

- The dialog box closes when the sequence is complete.

- Record on Maintenance Log.

2.	Perform Rinse Flow Cell Cleaning

Perform if Flow Cell in optical detector is suspected to be           dirty.

(Sysmex XS-Series Instructions for Use manual)

- Click [Menu] or press F4.

- Click “Controller” icon on the menu screen.

- Click “Maintenance” icon. The maintenance screen displays.

- Click on “Rinse Flow Cell”. The Rinse Flow Cell dialog box displays.

- Attach the appropriate tube adapter.

- Place a tube of 5% Sodium Hypochlorite solution (CELLCLEAN) in the tube adapter.

- Press the Start switch (Located above the sample tube position on the Main Unit of the XS-1000i without Sampler; inside the sampler cover on the XS-1000i with Sampler) to initiate the cleaning.

h.  The dialog box closes when the sequence is complete

i.    Record on Maintenance Log.

3.	Perform Air Bubble Removal for Flow Cell

Perform if air bubbles in Flow Cell create abnormal aggregate pattern on scattergram.

(Sysmex XS-Series Instructions for Use manual)

- Click [Menu] or press [F4].

- Click “Controller” icon on the Menu screen.

- Click “Maintenance” icon. The maintenance screen displays.

- Click “Remove Air Bubbles”. The sequence begins.

- The dialog box closes when the sequence is complete.

- Record on Maintenance Log.

    - Remove RBC Clogs with Remove Clog Sequence

Perform when “RBC Clog Error,” “RBC Bubble Error”, or                  “RBC/PLT 	Sampling Error” message displays.

(Sysmex XS-Series Instructions for Use manual)

- When error message is displayed, the HELP dialog box displays the error. Click on the error.

- Click [OK] to begin Clog Removal Sequence.

Alternate Method for Performing Clog Removal

(when no error is displayed)

c.	Click [Menu] or press [F4].

- Click “Controller” icon on the Menu screen.

- Click “Maintenance” icon. The Maintenance screen displays.

- Click on “Remove Clogs” icon and the sequence begins.

- VIII.	CALCULATIONS

- If correcting the HGB or HCT due to interfering substances recalculate and correct the affected indices:

MCHC	=	HGB/HCT x 100

MCH		=	HGB/RBC x 10

MCV		= 	HCT/RBC x 10

- If a sodium citrate tube is used for EDTA induced platelet clumping, multiply the citrate platelet count by 1.11 to correct for anticoagulant dilution.

- REPORTING RESULTS

- A.	Adult Reference

| Parameter | Reference Range | Parameter | Reference Range |
| --- | --- | --- | --- |
| WBC: | 4.8-10.8 K/cmm | Neut % | 42.2-75.2 |
| RBC | male: 4.7-6.1M/cmm | Lymph % | 20.5-51.1 |
|  | female:4.2-5.4 | Mono % | 1.7-9.3 |
| HGB | male:14-18 g/dL | Eo % | 0-10 |
|  | female:12-16 | Baso % | 0-0.8 |
| HCT | male:42-52 % | Neut # | 1.4-6.5 |
|  | female:37-47 % | Lymph # | 1.2-6.5 |
| MCV | male:80-94 fL | Mono # | 0.1-0.6 |
|  | female: 81-99 fL | Eo # | 0.0-0.7 |
| MCH: | 27-31 pg | Baso # | 0.0-0.2 |
| MCHC: | 33-37 g/dL |  |  |
| PLT: | 130-400 K/cmm |  |  |
| RDW-CV: | 11.5-14.5 % |  |  |
| RDW-SD: |  |  |  |
| MPV: | 7.4-10.4 fL |  |  |

- B.	Pediatric Reference Range:

| Parameter | Reference Range | Parameter | Reference Range |
| --- | --- | --- | --- |
| WBC: | 0-<2: 6.0-17.5 2-12: 5.0-15.0 >12: Adult Range | Neut % | 42.2-75.2 |
| RBC | 4.7-6.1 | Lymph % | 20.5-51.1 |
| HGB | 0-3mos: 13.7-20 3 mo-<3yrs:9.5-14.5 3yrs-12: 10.5-14 >12: Adult Range | Mono % | 1.7-9.3 |
| HCT | 0-3mos: 42-60 3 mo-<3yrs: 31-41 3yrs-12: 34-44 >12: Adult Range | Eo % | 0-10 |
| MCV | 80-94 | Baso % | 0-0.8 |
| MCH: | 27-31 | Neut # | 1.4-6.5 |
| MCHC: | 33-37 | Lymph # | 1.2-3.4 |
| PLT: | 130-400 |  |  |
| RDW-CV: | 11.5-14.5 | Mono # | 0.1-0.6 |
| RDW-SD: |  | Eo # | 0.0-0.7 |
| MPV: | 7.4-10.4 | Baso # | 0.0-0.2 |

- C.	Acceptable Reporting Format:

- The Sysmex calculates results and transmits these electronically to the RPMS Lab Package. These results are then verified by the technologist and become part of the patient’s electronic health record. The patient report displays units and normal ranges for each analyte.

- A copy of all samples run will be printed each day and retained in the laboratory for 2 years.

- D.	Reporting Critical Results to Physicians:

- The requesting physician or physician's nurse should be notified immediately when test results exceed the following limits:

| TEST | LOW | HIGH |
| --- | --- | --- |
| Hemoglobin | < 7.5 g/dL | > 18.9 g/dL |
| Hematocrit | < 27 % | > 54 % |
| WBC | < 2,500/mm | > 20,000/mm |
| RBC | - | > 6.9 million/mm |
| Platelets | < 50,000/mm | > 600,000/mm |

- PANIC VALUES are hand delivered to the requesting provider, the providers nurse or the clinical director; a notation is made on the lab report.

- Abnormal Values or flagged values should be repeated and verified that the results duplicate before they are turned out to the provider.

- E.	Slide Review: Slide reviews and manual differentials are referred to Regional Medical Lab. A manual differential will be added to the CBC through the RPMS Lab package and sent to the reference lab. The protocol for ordering a Manual differential is as follows:

- IMMATURE GRAN? (>100)

- SUSPECT BLAST MESSAGE (>100)

- ANY UNREASONABLE DIFF OR SCATTER PLOT

- WBC > 25.0

- WBC < 2.0

- NEUTROPHILA > 85%

- LYMPHOCYTES > 65%

- MONOCYTES   > 25%

- WBC < 3.5 WITH > 75% NEUTROPHILS

- SPECIFIC VERBAL REQUEST FOR MANUAL DIFF FROM PROVIDER

- ANY SIGNIFICANT RBC MORPHOLOGY, PLT FLAGS, OR Q- FLAGS (LEFT TO THE DISCRETION OF TECH)

- X.	PROCEDURE NOTES

- Nucleated RBC's. Perform manual differential and correct Sysmex        XS-1000i WBC count for the NRBC’s per 100 WBC’s counted.

WBC  =  XS WBC count x 100 / 100 + NRBC

- Megakaryocytes. When megakaryocytes are present, perform a WBC and PLT estimate.

- Capillary Analysis:

- Use when insufficient patient sample is available for aspiration      (<500 μL) or a sample has a parameter above the linearity limits of the analyzer.

- Use if there are marked changes in plasma constituents, (i.e. very low sodium or very high glucose), prepare a 1:7 dilution and allow to equilibrate before analysis.

- Analysis of the specimen on Sysmex XS-1000i is recommended before removing the cap to make a smear.

- Do not place samples on a mechanical rocker. Excessive mixing may induce platelet clumping and alter white cell membranes resulting in false interpretive messages.

- CLOROX ULTRA, a filtered bleach, is recommended for use in cleaning. If CLOROX ULTRA is not available generic bleach may be used, but must be 5% Sodium Hypochorite concentration and be free of particles that may cause background contamination when used on the analyzer. CLOROX ULTRA must be diluted from 6% by volume (straight from the bottle) to 5% before preparing further dilutions recommended for maintenance.

G.	For troubleshooting specifics refer to the Sysmex XS-Series Instructions for Use manual.

- XI.	LIMITATIONS OF PROCEDURE

A.	Sysmex XS-1000i MANUFACTURER STATED LINEARITY:

| Parameter | Range | Units |
| --- | --- | --- |
| WBC | 0-400.0 | x103/μL |
| RBC | 0-8.00 | x106/μL |
| HGB | 0-25.0 | g/dL |
| HCT | 0-60.0 | % |
| PLT | 0-5000 | x103/μL |

1.  Parameters that exceed these limits are flagged with @ beside the result. The sample must be diluted, rerun and multiplied by the dilution factor, or dilute 1:7 and use the capillary mode.

2.  Note the use of dilution for linearity on the patient report.

B.	Known Interfering Substances

- Specimens must be free of clots and fibrin strands.

- Marked changes in plasma constituents (e.g., low sodium,         extremely elevated glucose) may cause cells to swell or             shrink.  The blood to anticoagulant ratio is important.

- Red cell fragments, microcytic RBC's or white cell                       cytoplasmic fragments may interfere with automated platelet       counts.

- Cold agglutinins produce spurious macrocytosis, elevated           MCH's,  MCHC's, falsely decreased RBC counts and HCT's.       Rare Warm agglutinins produce the same spurious results         as a cold agglutinin.

- Extremely elevated WBC's may cause turbidity and increase       the hemoglobin.

- Severely hemolyzed samples (in vitro) falsely decrease RBC       and hematocrit.  Recollect hemolyzed specimens.

- Giant platelets and clumped platelets may falsely elevate the       WBC count and falsely decrease the platelet count.

- Platelet clumping and/or "platelet satellitism" can occur in           specimens collected in EDTA. This may falsely elevate the        WBC and falsely decrease the platelet count. Recollect the        specimen in Sodium Citrate anticoagulant and multiply by          1.11 dilution factor.

- Lipemia falsely elevates the HGB and MCHC. Perform a             plasma replacement, plasma blank procedure or make 1:7        dilution with CELLPACK.

- Severely Icteric samples may falsely elevate the HGB value       and  related indices. Make a 1:7 dilution with Sysmex                 CELLPACK.

- Rocking specimen excessively may affect the WBC                    differential.

12. 	Megakaryocytes may falsely increase WBC counts on                                         automated hematology analyzers.

- XII.	REFERENCES

- Sysmex XS-1000i Instructions for Use, Sysmex Corporation, Kobe, Japan, February 2006.

- Sysmex XS-1000i Users Guide, Sysmex Corporation, Kobe, Japan, February 2006.

- Sysmex NE-Series User's Guide, Sysmex Corporation (USA), Inc., Clinical Applications Division, Los Alamitos, CA, 1991 pg. 39.

- Koepke, John.  Practical Laboratory Hematology.  Churchill Livingstone Inc. 1991.   p. 24-25, 36-39.

- NCCLS. Clinical Laboratory Technical Procedure Manuals-Forth Edition; Approved Guideline. (GP2-A4, 2002).

- Sysmex America Inc., Mundelein, IL. Instructions for use of Sysmex SCS-1000™ with Sysmex XE-Series Calibration Program, Rev. 11, 21-Apr-00.

- Sysmex America, Inc., Mundelein, IL. Sysmex e-CHECK Hematology Control insert sheet for Sysmex XE-Series and XS-1000i Analyzers, Rev. 1, 30-Aug-02.

- Sysmex America, Inc., Mundelein, IL. Sysmex e-CHECK Insight™ Overview Guide, Appendix A-Xm Quality Control, Version 1.0a, 14-September-00.

- Sysmex America, Inc., Mundelein, IL. Sysmex Insight™ Overview Guide for Sysmex    e-CHECK, Document No: WI-2493A Rev. 06, May 23, 2006.

- Garrity, P., Walters, J., Concepts in New Age Hematology, A Hematology Monograph, Baxter Healthcare Corporation, Scientific Products Division, Hematology Support Services. August 1990

- Cornbleet J. Spurious results from automated hematology cell counters.  Laboratory Medicine. 1983:8:509-514.

- CLOROX ULTRA Professional Products Company, Oakland, CA. CLOROX ULTRA product label, 1998.

- Sysmex Reagents of America, Inc. MSDS sheets and reagent product inserts.

- College of American Pathologists (CAP) Hematology and Coagulation Checklist, October 2001.

- Brigden, Malcom L. Cell Counter-Related Abnormalities, Laboratory Medicine, May 1999, Vol. 30, #5, p.325-334.

- Gould, N., Connell, B., Dyer, K., Richmond, T., Performance Evaluation of the Sysmex XE-2100 Automated Hematology Analyzer, Sysmex Journal International, 1999 Vol. 9, No. 2, pp. 120-128.

Sysmex XS-1000i, Sysmex CELLPACK, Sysmex STROMATOLYSER-4DL, Sysmex STROMATOLYSER-4DS, Sysmex SULFOLYSER, Sysmex e-CHECK, Sysmex          e-CHECK (XS), Sysmex Insight are trademarks of the Sysmex Corporation.

All other trademarks are the property of their respective owners.