| Title: | BinaxNOW Malaria POC |
| --- | --- |
| Department: | Hematology Lab |
| Originator: | L Freeman, CLS Specialist |

| Background: | As a part of the Ebola testing protocol it is recommended that screening for malarial organisms be undertaken as the patient is statistically more likely to have a malarial infection than the Ebola virus at this institution. A point of care malarial test has become available eliminating the need to perform thick and thin smears eliminating the need for decontamination of smears for microscopic examination, and greatly decreasing turn-around times since there is no smear preparation or staining undertaken. Once validated we plan to implement the test in the Hematology Department for all malarial testing. The BinaxNOW is an immunochromatographic assay for the qualitative detection of Plasmodium antigens circulating in human EDTA blood samples or fingerstick samples. The test targets the histone rich protein II (HRPII) antigen specific to P. falciparum and a pan-malarial antigen common to all four types of Plasmodium organisms seen in humans. |
| --- | --- |

| Scope: | This process validation plan includes - Method evaluation for the BinaxNOW malarial test cards |
| --- | --- |

| BinaxNOW Malaria Operational Qualification | At a minimum include the following: - Following vendor information and recommendations, write procedures and forms for instrument - Operation - Set up quality control schedules - Add any additional information regarding issues that occurred during operational qualification |
| --- | --- |

| Performance Specification BinaxNOW Malaria | Performance of the BinaxNOW Malaria will be evaluated as summarized below: Performance Specification; Testing Performed; Expected Results. Accuracy; 1. Run external controls. 2. Run a minimum of 20 patient samples 15 negative for malaria and 5 negative samples spiked with positive control material of various levels and one RMAL CAP survey 3. For implementation in Hematology department 10 samples will be examined microscopically concurrent to BinaxNOW testing; 1. Control band will be positive on both positive and negative control and T1 and T2 bands will also be positive on positive control only. 2. All known negative samples will test positive on the C line only while all positive samples will test positive on the C line and at least 1 of the T lines 3. All 10 samples will correlate between the BinaxNow and smear examination.. Precision; 1. Run positive control and negative control 10 X each over at least 10 days; Each control will show positivity at C, T1 and T2 for the positive control and positivity at C only for negative control. Reference range; Reference Range is negative for this test.. Clinical Sensitivity/ Specificity; Review of Manufacturer’s literature. Analytical Specificity; Review of manufacturer’s literature Regarding interfering substances Run at least one icteric sample to see if visual impairment of control lines occurs; No interferences with hemolysis, icterus and lipemia expected. |
| --- | --- |

| Performance Specification | Performance Specification; Testing Performed; Expected Results. Carryover; There is no carryover for this test as each test is contained on its own card and uses disposable pipette tips.. |
| --- | --- |

| Report Check | For all test parameters - Include examples of LIS worksheets that are called up - LIS reports - EMR reports Verify that normal ranges and report formats are correct. |
| --- | --- |

| Training Plan | Prepare training and initial competency plans and forms for each of the following: - Test Operation Procedure - Competency records for all staff |
| --- | --- |

| Validation Packet | Assemble the validation packet and submit for supervisor and director review. |
| --- | --- |
|  | The validation packet should include the following; - Approved Validation Plan - Final SOPs for implementation - Product inserts used during validation - MSDS forms for all reagents - BinaxNOW Operational Qualification Data Summary - Test Performance Qualification Data Summary - Initial training and competency documentation - SOP readership documentation . - Billing code summary - Evidence of enrollment in a proficiency testing program - Final sign-off by the director |

| Comments/ Requests For Additional Information | Comments or requests for additional information about the validation plan follow: |
| --- | --- |

| Validation Plan Approval | Name; Signature; Title; Date. Denis Dwyre, MD; Hematology Director. |
| --- | --- |