Siemens Xpand Competency Quiz

Name:________________________________________________Date:________________________

- What is the best order to proceed for daily start up on the Expand?

- Calibrated assays, cut waste cuvettes, pull reagent inventory list, and perform daily maintenance

- Pull  reagent inventory list, calibrate assays, cut waste cuvettes, perform daily maintenance

- Cut cuvette waste, perform daily maintenance, Pull reagent inventory list, calibrate assays

- Pull reagent inventory list, calibrate assays, perform daily maintenance, cut cuvette waste

- If daily system checks fail, you should-

- Ignore and continue

- Change both reagent and samples probes

- Clean windows

- Check lot # of Absorbance Flex to reagent ABS Flex and repeat with fresh solution.

- Above assay range on analyzer print out means-

- The analysis is above the reference range of the assay.

- The analysis is above the linear range of the instrument.

- The analysis is above the CAV range for the assay.

- All of the above

- Calibration for the ISE should be-

- Done once daily during start up

- Automated about every 2 hours

- Every 60-90 days

- Only done when there is an issue with QC.

- Quality controls should be run-

- Daily

- Weekly

- Every 8 hours

- Daily on assays being performed for patient testing, after calibrations, and after maintenance or major service to analyzer.

- A calibration failure could be due to-

- Wrong calibrator used

- Fresh calibrator needed

- Calibrator loaded incorrectly in sample segment

- All of the above.

- When calling service you should have __________available.

- Your  name and social security number

- Instrument serial and model numbers

- Your name and phone number

- B and C

- Calibration of a reagent is performed when…

- There is a new lot of reagent

- When QC is consistently out of range

- When shifts or tends in QC are identified

- All of the above

- A  yellow Needs Check icon will appear when

- The system is shutting down

- the systems wants you to add more reagent

- the system is checking itself for any needs to process load list

- All of the above

- A sample tube icon will appear when….

- The system is waiting for samples to be loaded

- The system does not have any HM Vessels left

- The cuvette waste needs to be emptied

- The system does not have enough sample volume to process request

Reviewed by:______________________________Date:________________Scrore:______________