CHEM.ARCH.7.0	ACCELERATOR APS LINE OPERATING PROCEDURE

PRINCIPLE

The ACCELERATOR APS workcell is a modular system designed to automate sample handling and processing in the clinical laboratory. The system allows consolidation of multiple clinical chemistry and immunoassay analytical instruments into a unified workstation.

PURPOSE

The ACCELERATOR APS (Automated Processing Systems) workcell is a modular system designed to automate pre-analytical processing, sample handling, and processing in the clinical laboratory.

DOCUMENT OWNER

Supervisor, Chemistry and Special Chemistry

INTRODUCTION

The system consolidates multiple analytical instruments into a unified workstation by employing a common sample processing capability. The workcell software provides for workload management, sample order management, and instrument operational status monitoring. This is accomplished through communication connections between the workcell, analyzers, and LIS (laboratory information systems) or middleware.

The 7 major components of the APS workcell include the Input/Output module, the Track module, one or two Centrifuge modules, the Decapper module, the Resealer module, the Storage and Retrieval Module, and the Analyzer Interface Module.

    - Input/Output Module:

The Input/Output module provides a single point of sample tube input and output, barcode tube identification and tracking, and operator interface.

The Input/Output module holds up to 15 48-position routine racks in numbered lanes, which can be designated for specific purposes, such as input lane, output lane, sorting lane and Priority output lane. In addition, there are two lanes for 12-position Priority input racks and three lanes for 12-position Priority output racks. Lanes can also be designated for specific tubes, such as capped or uncapped, pre-spun or unspun, etc.

The Input/Output module is equipped with a two-color LED indicator for each rack lane. The following bullets describe the different LED indicator conditions.

• Red indicator: rack is locked into a lane and needs to be processed.

• Green indicator: rack is locked and processing is completed.

• Off indicator: rack is unlocked or no rack is loaded.

The Touch Screen monitor and mouse allows users to interface with the workcell. Users can select onscreen options by selecting function and navigation buttons, menu items, and text areas. In addition, the Input/Output module stores the computer and UPS for the workcell.

    - Resealer Module:

The Resealer module allows automatic sealing of plastic sample tubes after analytical processing. At the Resealer module, tube grippers hold the sample tube while it is sealed with aluminum foil from the Resealer coil. Individual foil seals are cut for each sample tube and heat-sealed to the plastic sample tube.

    - Storage and Retrieval Module:

The Storage and Retrieval module stores sample tubes from the workcell in a temperature-controlled and protected environment. The Storage and Retrieval module consists of the storage worktable and automatic refrigerator. The automatic refrigerator consists of 20 floors, with each floor holding 16 48-position routine racks, for a total of 15,360 sample tubes. After a pre-defined, configurable time interval, the sample tubes are discarded into the storage waste.

    - Analyzer Interface Module:

The analyzer Interface module provides the path required to move sample tubes to the analyzer. If a sample tube has test requests for an analyzer, the tube carrier is moved to the analyzer for sample aspiration and processing. After sample aspiration, the sample tube is returned to the main lane of the workcell. The following are different types of analyzer interface modules designed for the following instrument types:

    - ARCHITECT c Systems

    - ARCHITECT i 2000

    - ARCHITECT i 2000SR (see Section -- for further Information)

    - Non-Abbott instruments section (see specific instrument section for further information)

    - Track Module:

The Track module provides the path required to move sample tubes from the Input/Output module to the Centrifuge, Decapper, Resealer and Storage and Retrieval modules, and the analytical instruments.

The Track Module consists of the following 5 parts for samples to be processed:

      - Routine Racks:

Routine racks are four-column, 48-position racks used for routine samples and for Priority output samples, e.g., samples with automation errors or those requested by the operator. The routine racks have a handle and rack ID.

      - Priority Racks:

Priority racks are one-column, 12-position racks used for STAT samples in the two Priority input lanes and for Priority output samples in the three Priority output lanes. The priority racks have a handle.

      - Storage Racks:

Storage racks are four-column, 48-position racks used to store sealed sample tubes in the Storage and Retrieval module. This rack does not have a handle.

      - Tube Carriers:

Tube carriers are used to transport tubes on the Track module. The tube carriers have a unique Radio-Frequency ID (RF ID) tag used to identify and locate sample tube positions along the Track module.

      - Reselaer Coil:

Resealer coil is the aluminum foil used by the Resealer module to seal plastic sample tubes.

SCOPE

The scope of this procedure encompasses laboratory personnel who are directly involved in the receipt and processing of blood specimens. This procedure should be used as a knowledge resource.

SAFETY

  - Bloodborne Pathogen Risk: HIGH

The exposure to a Bloodborne Pathogen is relatively high considering the state of some specimens which may be tainted with blood due to the conditions of collection prior to laboratory receiving.

  - Chemical Risk: LOW

There is no known Chemical Risk associated with the Accelerator APS Line workcell other than pathogenic risk of the specimens it processes

  - PPE required: LAB COATS, GLOVES, GOGGLES, FACE SHIELDS, AND COUNTER-MOUNTED SHIELDS.

Follow all PPE precautions when working in a laboratory setting as stated in the OSHA regulation found in the Safety Laboratory Manual.

- Refer to Task Assessment Log, Chemical Inventory Log and Chemical Registry for further information.

SPECIMEN REQUIREMENTS

- Sample Tube Specifications:

Sample tubes must meet the following specifications.

See Table 4.1.1 and Figures 4.1 thru 4.3 for detail.

Table 4.1.1: Sample tube specifications

- Acceptable Sample Tubes:

The ACCELERATOR APS workcell will only process plastic sample tubes with a removable closure on the open end.

VERY IMPORTANT: Never use glass tubes if the Resealer module is installed and in the on line status. Glass tubes cannot be sealed properly. If there are glass tubes that require processing, verify the Resealer module is off-line.

IMPORTANT: Follow the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed, assay-specific information about specimen collection, preparation, and storage.

Failure to follow tube and assay manufacturer's recommendations for sample tube type and handling could result in erroneous and misleading results.

- Use only sample tubes with screwed and pressure plastic closures.

- The tube surface must be smooth and free of any ridges or collars that could interfere with removal of the tube cap or closure.

- Use only plastic sample tubes when the Resealer module is installed and in the on-line status. The type of plastic tubes that can be sealed using the Resealer module are limited to:

  - PET (Polyethylene terephthalate)

  - PS (Polystyrene)

  - PP (Polypropylene)

Examples of acceptable tubes for processing on the ACCELERATOR APS are shown below. See figures 4.4.

NOTE: Refer to previous section for acceptable specimen tubes for the ARCHITECT i 2000SR.

- Unacceptable Sample Tubes:

The following are examples of unacceptable sample tube types.

- Description of Unacceptable Sample Tubes:

- Pediatric tubes: These include Red and Green tops along with STATfuge conical tubes used in separation of lipemic specimens.

- Cone bottoms tubes: These include those used as aliquots for urine drug screen specimens.

- Sample Cup: These were used to accommodate short specimens for sampling on the Architect prior to the APS Line.

- Open False Bottom: No longer used in the laboratory.

- Rubber Stopper Closure: UA tubes and 10mL red tops are unacceptable

- Ridged Tops: Do not carry in the lab.

- Flat Bottom Tube: Similar to send out tubes

- Barcode Label Requirements:

Depending on the type of barcode label, requirements include information about guidelines, format, sample ID length, label length, and placement.

    - Sample Barcode Symbology:

The barcode readers support the following barcode symbologies:

- Code 128

- Code 39 with and without check digit

- Interleaved 2 of 5 with and without check digit

- USS Codabar

NOTE: When the check digit is enabled, it is enabled for all symbologies that allow for check digit configuration (Code 39 and Interleave 2 of 5).

    - Sample ID Identification Length:

The following factors affect the number of characters that can fit on a sample barcode label.

- Sample barcode label length

- Sample barcode label placement

- Symbology used

- Density

- Ratio

- Quiet zone

For the ARCHITECT System, the maximum number of characters allowed on the sample barcode label is 20 characters. This value may be less for some symbologies, due to overall label length, narrow bar width, and ratio.

    - Sample Barcode Length:

Generally, a 50 mm label fits a 75 mm sample tube, and a 75 mm label fits a 100 mm sample tube.

To ensure the entire label is visible to the barcode reader when the tube is loaded into the tube carrier, do not place the barcode label lower than approximately 20 mm from the bottom of the sample tube. Place barcode labels on tubes vertically as straight as possible.

The presence of an external plastic cap on a sample tube will decrease the available space for applying the barcode label. The size of an external plastic cap may range from 15 to 20 mm.

The following are maximum sample barcode label lengths acceptable for 75 mm and 100 mm sample tubes, with and without external caps.

    - Sample Barcode Label Placement:

The barcode label should be affixed to the sample tube in a “ladder” orientation, where the printed bars are horizontal. The correctly labeled tube in Figure 4.12 illustrates an example of correct label placement.

Position the barcode label on the sample tube so that:

- Label is oriented along the length of the tube. Do not wrap the label horizontally around the tube.

- Label does not overlap the top or bottom of the tube.

- Label is vertically as straight as possible.

- Label does not touch the tube closure.

- Label should be free of excess moisture.

- Label should be placed at the top of the sample tube to ensure the barcode label does not extend beyond the top of the tube carrier.

NOTE: The operator is responsible for placing the correct barcode label on the sample tube to ensure proper sample identification.

IMPORTANT: When a new barcode label is added to a sample tube, ensure the original barcode label is covered entirely by the new barcode label.

    - Sample ID Allowable Characters:

The allowable character set for sample barcode IDs consists of all printable ASCII characters, except the characters listed below.

- The following printable ASCII characters cannot be used in a sample ID.

- & (ASCII 26 hex - decimal 038)

- \  (ASCII 5C hex - decimal 092)

- ^ (ASCII 5E hex - decimal 094)

- | (ASCII 7C hex - decimal 124)

- ~ (ASCII 7E hex - decimal 126)

- The following printable ASCII characters cannot be used as the first digit of a sample ID.

- (ASCII 2A hex - decimal 042)

- ? (ASCII 3F hex - decimal 063)

- Examples of valid sample IDs include:

- 001040

- 001040A

- 001040-B

- 001040/C1

- 001040?D

- 001040*

- ABCDEFG/10H

- Examples of invalid sample IDs include:

- 001040\A

- 001040^B

- ?001040

- *001040

NOTE: Sample tubes with an invalid sample ID are automatically sorted to a priority output rack without processing.

NOTE: Use of allowable non-alphanumeric character may cause the SID to be improperly displayed in some screens. This is a display issue only, the samples will be processed as expected.

- The effected ACSII codes are:

- ASCII 20 hex - decimal 32 to ASCII 2E hex - decimal 46

- Characters <space> to /

- ASCII 3A hex - decimal 58 to ASCII 40 hex - decimal 64

- Characters : to &

- ASCII 5B hex - decimal 91 to ASCII 60 hex - decimal 96

- Characters [ to `

- ASCII 7B hex - decimal 123 to ASCII 7E hex - decimal 126

- Characters { to ~

NOTE: The APS will not process samples with duplicate sample IDs until the original sample has been removed from the system. For workcells with a Storage and Retrieval Module, do not re-use sample IDs until the Storage-Rack Disposal Timeout has expired. Refer to Storage and Retrieval Module – Configuration.

SUPPLIES AND EQUIPMENT

The following List Numbers are unique identifiers that are used when ordering products. These numbers are provided for guidance only and are subject to change. Contact your local Abbott representative for the most current List Numbers.

  - ACCELERATOR APS Accessory and Consumable List Numbers:

Accessories and consumables are used during routine processing and component replacement. Contact your Abbott representative to order the following supplies.

| Item | Quantity | List Number |
| --- | --- | --- |
| ACCELERATOR APS I/O Module Routine Rack | 1 | 07L10 |
| ACCELERATOR APS Priority/Exception Rack | 1 | 07L11 |
| ACCELERATOR APS Resealing Aluminum Coil | 1 | 07L12 |
| ACCELERATOR APS Tube Carriers | 10 | 07L15 |
| ACCELERATOR APS Computer Air Filter | 1 | 07L22 |
| ACCELERATOR APS Hex Tool Set | 1 | 07L24 |
| ACCELERATOR APS Tube Storage Rack | 1 | 07L37 |
| ACCELERATOR APS Balance Rack Silicone Insert | 1 | 07L43 |
| ACCELERATOR APS i 2000SR Tube Carriers | 4 | 07L47 |

  - Manual List Numbers

Support documentation is available in printed format. Contact your Abbott representative to order the following manual.

| Item | Quantity | List Number |
| --- | --- | --- |
| ACCELERATOR APS Operations Manual | 1 | 07L19 |

MAINTENANCE

Proper service and maintenance of the ACCELERATOR APS workcell is one of the most important aspects of a complete quality assurance program. A thorough service and maintenance program will:

- Minimize down time

- Maintain records for inspection and accreditation

- Maintain optimal workcell operation to provide optimal performance

- Refer to Section 9 of the APS Operational Manual for more information.

Performance of a maintenance procedure should be documented in the Maintenance Log under the appropriate maintenance section. Refer to ACCELERATOR APS Maintenance Log,

  - Maintenance Schedule

| FREQUENCY | PROCEDURE |
| --- | --- |
| bASED on workload | - Workload based shutdown |
| weekly | - Clean touch screen - Clean input/output module - Empty desealer waste - Shut down - Empty storage |
| monthly | - Clean resealer module - Clean computer air filter – storage retrieval module - Clean input/output module robot gripper - Clean desealer module - Clean storage and retrieval module robot gripper - Clean computer air filter – input/output module - Clean UPS – input/output - Clean UPS – storage retrieval module |
| as needed | - Check storage and retrieval module temperature - Clean racks - Replace resealer coil - Record storage retrieval module rack errors |

  - General Routine Care and Cleaning

Some maintenance procedures are performed routinely.

WARNING: Potential Biohazard. Only trained operators should perform the following maintenance procedures. Follow your laboratory standard operating procedures and guidelines for biohazardous materials handling.

IMPORTANT: Do not allow bleach or soapy water to spill through the workcell cover, or set any bleach solution or soapy water containers on the workcell surface. This could lead to contamination of samples, possibly causing erroneous patient results.

  - Bleach Preparation Solution for Routine Cleaning

Decontamination procedures may require use of diluted sodium hypochlorite. Use the following formula to calculate the volume of water required to mix with manufacturer-supplied sodium hypochlorite.

X = (B – A)/A

Where:

| A= | % of sodium hypochlorite solution desired |
| --- | --- |
| B= | % of sodium hypochlorite (active or available chlorine) in manufacturer-supplied solution |
| C= | Number of parts of water required to mix with one part of manufacturer-supplied sodium hypochlorite (active or available chlorine) solution |

NOTE: Stability of the Bleach solution is 30 days after preparation.

  - Logging In/Out For Routine Maintenance

Log in with the appropriate user access level as necessary to perform the following maintenance procedures.

- To log in:

- Press the System navigation button.

- Press the Configuration option button.

- Press the Login function button. Enter the User ID using the pop-up screen then press the Confirm button.

- To log out after completing a procedure:

- Press the System navigation button.

- Press the Configuration option button.

- Press the Login function button. Enter 000 on the pop-up screen then press the Confirm button.

The Supervisor login password is needed for many of the maintenance procedures. Refer to the Login section for additional information on user logins.

  - MAINTENANCE Based on Workload

Perform the following maintenance procedure whenever interface communication slows down.

    - Workload Based Shutdown/Powercycle

    - Weekly Maintenance

Perform the following procedures weekly:

    - Clean the Touch Screen

    - Clean the Input/Output Module

    - Clean and Lubricate the Centrifuge(s) and Inspect Centrifuge(s) Buckets for Cracks

NOTE: For other Centrifuge maintenance procedures, refer to the manufacturer’s literature.

    - MONTHLY MAINTENANCE

Perform the following maintenance procedures monthly:

    - Clean the Centrifuge Module(s) Robot Gripper.

    - Clean the Decapper Module

    - Clean the Input/Output Module Robot Gripper

    - Clean the Resealer Module

    - Clean the Storage and Retrieval Module Robot Gripper

    - Clean the Computer Air Filter(s)

    - Clean the UPS

    - AS NEEDED MAINTENANCE

Perform the following maintenance procedures As Needed:

    - Check the Storage and Retrieval Module Temperature

    - Clean the Racks

    - Replace the Resealer Coil

    - Replace the Silicone Insert of the Balance Rack

PROCEDURE

The ACCELERATOR APS is flexible and accommodates many laboratory environments and work flows. Before attempting to operate the workcell, it is important to be familiar with the hardware components of your workcell and the fundamental principles of the software user interface. Refer to the PRINCIPLE section on page 1 for additional information.

- Workcell Shutdown and Startup:

The operator may need to shut down or start up the workcell and its components to:

- Perform routine maintenance

- Perform troubleshooting diagnostic procedures

- Perform unexpected or emergency shutdown

- Power Shutdowns

There are three procedures to remove power from the ACCELERATOR APS workcell:

      - Emergency Power Shutdown

      - Slow or non-responsive software

      - Controlled Power Shutdown

NOTE: Incompleted request for sample tubes to be retreived from the Storage module at the time of system shutdown, will be delivered to the I/O module after powering on the workcell and selecting Run.

      - Shutdown Prodecures

| Shutdown | Description |
| --- | --- |
| Emergency | Removes power from all areas of the workcell. Shutdown is accomplished by pressing the UPS OFF button(s) and removing the power cords from their receptacles. No data backup is performed before power loss, and the files could become corrupted. |
| Software | If a shutdown cannot be requested via the software, the computer(s) must be shut down. This may occur when software is slow to respond or non-responsive. |
| Controlled | Shutdown is requested via the software. Before the instrument shuts down, a data backup is performed. The UPS OFF button(s) is(are) pressed after the workcell shuts down. Controlled shutdown is used when power must be removed for routine maintenance. NOTE: To completely remove the power, unplug the workcell and storage and retrieval module power cord and the centrifuge(s) power cord from their receptacles. The workcell and storage retrieval module share the same power cord. |

- Important Power Safety Information

The power shutdown procedures must be performed in the prescribed order for the following reasons:

- Turning OFF the computer power switch activates a signal that prevents the UPS from providing power to workcell components when the workcell power cord is disconnected from the receptacle.

- If the computer power switch is ON, the UPS power button must be pressed to turn the UPS OFF, even if the workcell power cord was removed from the receptacle.

WARNING: Electrical shock hazard. If the UPS OFF button is not pressed, full power is provided to the instrument for approximately 10 minutes, even if the workcell power cord is unplugged. To completely remove power from the workcell, the UPS must be turned OFF and the main power cords unplugged.

- Emergency Shutdown

Performing an Emergency Shutdown

IMPORTANT: Sample integrity could be compromised if samples are left on the system or the Storage and Retrieval module when power is shut down. Refer to your laboratory guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

NOTE: This procedure does not power down the analyzers. Refer to the specific analyzer operations manual for additional information.

NOTE: Use this procedure if the ACCELERATOR APS workcell must be shut down in an emergency situation when a controlled shutdown is not possible.

- Unplug the main power cord from its power receptacle.

WARNING: Electrical shock hazard. To remove power completely, the workcell must be unplugged from its power receptacle even if the UPS is switched off and no power is being supplied to the workcell.

- Power down the Input/Output module.

- Open the Input/Output module door, with the handle.

- Press the UPS OFF button (#2, below) for approximately three seconds. The self-test “ ” indicator turns off when shutdown is complete.

WARNING: Electrical shock hazard. If the UPS OFF button is not pressed, full power is provided to the instrument for approximately 10 minutes, even if the instrument power cord is unplugged. To completely remove power from the workcell, the UPS must be turned OFF and the main power cords unplugged.

- Power down the Storage and Retrieval module (if available).

- To access the Storage and Retrieval module UPS/CPU, pull the panel forward and remove it.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

The panel is identified by following label.

- Press the Storage and Retrieval module UPS OFF button for approximately three seconds. The self-test “ ” indicator LED turns off when shutdown is complete.

NOTE: The workcell and Storage and Retrieval module share the same power cord.

- Unplug the power cord of the Storage and Retrieval module shelves and refrigerator from its receptacle or any laboratory-supplied UPS/generator.

NOTE: When the Storage and Retrieval module UPS OFF button is selected, power is disconnected to the sample robot, rack robot, shutter, and computer. Power for the Storage and Retrieval module storage shelves and refrigerator must be disconnected by unplugging the shelves and refrigerator power cord.

- Power down each ARCHITECT i 2000SR Interface module (i 2000SR IM), if available.

- Our power switch is type 2, turn to the “APS/OFF” position.

- Restoring Power after an Emergency Shutdown

IMPORTANT: Sample integrity could be compromised if samples are left on the workcell or storage table when power is shut down. Refer to your laboratory guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

NOTE: This procedure does not restore power to the analyzers. Refer to the specific analyzer operations manual for additional information.

- Power on each ARCHITECT® i 2000SR IM i 2000SR, if available.

f the power switch is type 2, make sure i 2000SR IM power switch is in the “APS/OFF” position. The ARCHITECT i 2000SR IM will be automatically turned on when the APS UPS is switched on.

- Power on the Input/Output module.

- Plug the workcell power cable cord into its receptacle.

NOTE: The workcell and Storage and Retrieval module share the same power cord.

- Open the Input/Output module door, with the handle.

- Press the UPS ON button. When the self-test “ ” indicator LED turns green, proceed to the next step.

- If the green Power indicator is on, go to next step. Otherwise, push the computer power switch up and hold approximately three seconds, then release.

- Close the Input/Output module front-left door.

- Power on the Storage and Retrieval module (if available).

- Empty the Storage and Retrieval module waste container.

- Remove the Storage and Retrieval module UPS/CPU access panel (if necessary) to access the UPS.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- Press the UPS ON button. When the self-test “ ” indicator LED turns green, proceed to the next step.

- If the green Power indicator is on, go to next step. Otherwise, push the computer power switch up and hold approximately three seconds, then release.

- Replace the Storage and Retrieval module UPS/CPU access panel.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- Plug the power cord of the Storage and Retrieval module shelves and refrigerator into its receptacle.

- After the ACCELERATOR APS software is launched, the workcell powers up and the Snapshot screen displays. The workcell is in Pause mode. Press the Run navigation button to unpause the workcell.

NOTE: After restoring power, all modules initialize when the Run navigation button is pressed. Wait for initialization to complete before processing sample tubes.

- When the message “Make sure that the Workcell is free of tubes” displays, check for the presence of sample tubes.

- Any sample tubes present on the workcell are read at the Input/Output module barcode reader and processed if included in the Expected or Incomplete worklist.

WARNING: Potential Biohazard. Biohazardous material could be present on workcell surfaces. Take appropriate precautions when performing this procedure. Follow laboratory standard operating procedures and guidelines when performing this procedure.

- Confirm the prompt by pressing the “ ” button.

- Remove all filled racks from the Input/Output module.

IMPORTANT: Review the middleware/LIS host to determine which samples have pending tests. Ensure sample tubes requiring reruns or additional testing are unloaded from the Storage and Retrieval module. If a sample tube has not been retrieved, a manual delivery of the sample tube may be required.

- Performing a Software Shutdown

Use this procedure if the ACCELERATOR APS workcell must be shut down when a controlled shutdown is not possible due to the software not being responsive to the Shutdown command.

IMPORTANT: Sample integrity could be compromised if samples are left on the system or the Storage and Retrieval module when power is shut down. Refer to your laboratory guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

NOTE: This procedure does not power down the analyzers. Refer to the specific analyzer operations manual for additional information.

- Open the Input/Output module (IOM) door with the handle.

- Push the IOM computer’s main power switch (#1) up until the green power light goes out, and then release the switch.

- Press the IOM UPS OFF button (#2, below) for approximately three seconds. The self-test “ ” indicator (#1) turns off when UPS power is completely off.

WARNING: Electrical shock hazard. If the UPS OFF button is not pressed, full power is provided to the instrument for approximately 10 minutes, even if the instrument power cord is unplugged. To completely remove power from the workcell, the UPS must be turned OFF and the main power cords unplugged.

- To access the Storage and Retrieval module (SRM) UPS/CPU, pull the panel forward and remove it. The panel is identified by following label.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- Push the SRM computer’s main power switch (#1) up until the green power light goes out, and then release the switch.

- Press the SRM UPS OFF button for approximately three seconds. The self-test   “ ” indicator turns off when UPS power is completely off.

NOTE: When the Storage and Retrieval module UPS OFF button is selected, power is disconnected to the sample robot, rack robot, shutter, and computer.

- Restoring Power after a Software Shutdown

IMPORTANT: Sample integrity could be compromised if samples are left on the workcell or storage table when power is shut down. Refer to your laboratory guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

NOTE: This procedure does not restore power to the analyzers. Refer to the specific analyzer operations manual for additional information.

- Open the Input/Output module (IOM) with the handle.

- Press the IOM UPS ON button (see #2 below). When the self-test “ ” indicator LED turns green (#1), proceed to the next step.

- If the green Power indicator is on, go to next step. Otherwise, push the computer power switch up and hold approximately three seconds, then release.

- Close the Input/Output module door.

- Power on the Storage and Retrieval module (SRM).

- Empty the Storage and Retrieval module waste container.

- Press the SRM UPS ON button (#2). When the self-test “ ” indicator LED turns green (#1), proceed to the next step.

- Push the SRM computer power switch (#1) up and hold approximately three seconds, then release. The green power light will illuminate.

- Replace the Storage and Retrieval module panel.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- After the ACCELERATOR APS software is launched, the workcell powers up and the Snapshot screen displays. The workcell is in Pause mode. Press the Run navigation button to unpause the workcell.

NOTE: After restoring power, all modules initialize when the Run navigation button is pressed. The IOM will also read the RFID of all tube carriers on the workcell. Wait for the steps to complete before processing sample tubes.

- When the message “Make sure that the Workcell is free of tubes” displays, check for the presence of sample tubes.

  - Any sample tubes present on the workcell are read at the Input/Output module barcode reader and processed if included in the Expected or Incomplete worklist.

WARNING: Potential Biohazard. Biohazardous material could be present on workcell surfaces. Take appropriate precautions when performing this procedure. Follow laboratory standard operating procedures and guidelines when performing this procedure.

  - If sample tubes are present in the Centrifuge module, perform the following steps.

  - Turn the locking knob of the Centrifuge counterclockwise to unlock.

  - Use the hex tool (3 mm) to remove the two screws holding the safety shield to the Centrifuge. Set aside the screws.

  - Lift the Centrifuge safety shield and open the Centrifuge lid.

CAUTION: Use caution when opening the lid to ensure it is fully open. For Centrifuge modules with the locking mechanism installed, ensure the safety shield is opened completely to engage the locking mechanism.

NOTE: If the sliding door of the Centrifuge is open, it is not possible to turn the locking knob. In this case, remove sample tubes through the sliding door.

  - Open the Centrifuge lid to access the rotor buckets and remove all sample tubes loaded in rotor buckets.

  - Close the Centrifuge lid.

  - Remove all sample tubes from bucket inserts in the Centrifuge loading and unloading areas.

  - Close the Centrifuge safety shield and tighten the screws. When closing the Centrifuge module safety shield with a locking mechanism, pull out the lock and lower the safety shield.

WARNING: Pinch Hazard. Use caution when pulling the lock for the locking mechanism. Fingers can be pinched.

  - Turn the locking knob of the Centrifuge clockwise to lock.

  - If two Centrifuge modules are installed, repeat the above steps for the second Centrifuge module.

- Confirm the prompt by pressing the “” button.

- Remove all filled racks from the Input/Output module.

IMPORTANT: Review the middleware/LIS host to determine which samples have pending tests. Ensure sample tubes requiring reruns or additional testing is unloaded from the Storage and Retrieval module. If a sample tube has not been retrieved, a manual delivery of the sample tube may be required.

- Performing a Controlled Power Shutdown

NOTE: This procedure does not shut down the analyzers. Refer to the specific analyzer operations manual for additional information.

WARNING: Potential Biohazard. Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

IMPORTANT: Sample integrity could be compromised if samples are left on the system or the Storage and Retrieval module when power is shut down. Refer to your laboratory guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

- Remove all samples from the workcell before performing this procedure. If necessary, perform the Purging Samples procedure.

  - Press the System navigation button.

  - Press the Configuration option button.

  - Press the Purge function button at the bottom of the screen.

  - Press “ ” to confirm the Purge procedure.

  - Allow the purge to complete. Verify all samples were purged by checking the on-line samples on the Worklist screen.

- Remove all racks from the Input/Output module.

- Press the Pause navigation button to pause the workcell. Ensure the Run/Pause indicator symbol displays.

- Ensure all Centrifuges are on-line.

  - Press the System navigation button.

  - Press the Configuration option button.

  - Press the System function button at the bottom of the screen.

  - Highlight the Centrifuge module(s).

  - If off-line, press the Selection button to change to on-line.

- Shut down the workcell.

  - Press the System navigation button.

  - Press the Configuration option button.

  - Press the Shutdown function button at the bottom of the screen.

  - Confirm the message “Are you sure you want to shut down?” on the pop-up screen by pressing the “ ” button.

- Power off the Input/Output module.

  - Open the Input/Output module door, with the handle.

  - If necessary, push the computer power switch up, then release. The green power indicator will turn off. Some computers will automatically power off, depending on the software.

WARNING: Electrical shock hazard. The remaining steps must be performed to remove power from all areas of the workcell.

  - Press the UPS OFF button for approximately three seconds. The self-test   “” indicator turns off when shutdown is complete.

  - Close the Input/Output module front-left door.

- Power off the Storage and Retrieval module (if available).

  - To access the Storage and Retrieval module UPS/CPU, pull the panel forward and remove it.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

  - If necessary, push the Storage and Retrieval module computer power switch up, then release. The green power indicator will turn off. Some computers will automatically power off, depending on the software.

  - Press the Storage and Retrieval module UPS OFF button for approximately three seconds. The self-test “ ” indicator turns off when shutdown is complete.

  - Replace the Storage and Retrieval module UPS/CPU access panel.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- Power down each ARCHITECT i 2000SR Interface module, if available.

  - If the power switch is type 1, turn to the i 2000SR IM power switch from the “APS” position to the “0” position.

  - If the power switch is type 2, make sure the i 2000SR IM power switch is in the “APS/OFF” position. The ARCHITECT i 2000SR IM will be automatically turned off when the APS UPS is switched off.

  - If it is necessary to completely remove power, unplug the workcell and Storage and Retrieval module power cord, and the Centrifuge(s) power cord from their receptacles. The workcell and the Storage and Retrieval module share the same power cord.

- Restoring Power after a Controlled Shutdown

NOTE: This procedure does not restore power to the analyzers. Refer to the specific analyzer operations manual for additional information.

    - If necessary, plug the workcell and Storage and Retrieval module power cord into its receptacle.

NOTE: The workcell and Storage and Retrieval module share the same power cord.

    - Power on each ARCHITECT i 2000SR IM i 2000SR, if available.

      - If the power switch is type 1, turn to the i 2000SR IM power switch from the “0” position to the “APS” position.

      - If the power switch is type 2, make sure the i 2000SR IM power switch is in the “APS/OFF” position. The ARCHITECT i 2000SR IM will be automatically turned on when the APS UPS is switched on.

    - Power on the Input/Output module.

  - Open the Input/Output module front-left door.

  - Press the UPS ON button. When the self-test “” indicator turns green, proceed to the next step.

  - If the green Power indicator is on, go to next step. Otherwise, push the computer power switch up and hold approximately three seconds, then release.

  - Close the Input/Output module front-left door.

    - Power on the Storage and Retrieval module (if available).

- Empty the Storage and Retrieval module waste container.

- Remove the Storage and Retrieval module UPS/CPU access panel (if necessary) to access the computer and UPS.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

- Press the Storage and Retrieval module UPS ON button. When the self-test “” indicator turns green, proceed to the next step.

- Push the Storage and Retrieval module computer main power switch up and hold approximately three seconds, then release.

- Replace the Storage and Retrieval module UPS/CPU access panel.

WARNING: Pinch Hazard. Use caution when removing or replacing the cover panel. Fingers can be pinched.

    - Plug the Centrifuge(s) power cords into the receptacles.

    - Power on the Centrifuge module(s).

NOTE: Toggle the Centrifuge power switch to the right to turn on the Centrifuge module. Repeat for the second Centrifuge module, if necessary.

    - After the ACCELERATOR APS software is launched, the workcell powers up and the Snapshot screen displays. The workcell is in Pause mode. Press the Run navigation button to unpause the workcell.

NOTE: When power is restored, all modules initialize when the Run navigation button is pressed. The Centrifuge module weighs all balance tubes and bucket inserts during initialization. Wait for initialization to complete and the Resealer module to reach operating temperature before processing sample tubes.

    - When the message “Make sure that the workcell is free of tubes” displays, check for the presence of sample tubes.

NOTE: This pop-up appears only if sample tubes were left inside the Centrifuge module or in racks locked in lanes during Shutdown.

- If a Purging Samples procedure was not completed before shutdown, then any sample tubes present on the workcell are read at the Input/Output module barcode reader and processed if included in the Expected or Incomplete worklist.

- If sample tubes present in the Centrifuge module were not purged, perform the following steps.

    - Turn the locking knob of the Centrifuge counterclockwise to unlock.

    - Use the hex tool (3 mm) to remove the two screws holding the safety shield to the Centrifuge. Set aside the screws.

    - Lift the Centrifuge safety shield and open the Centrifuge lid.

CAUTION: Use caution when opening the lid to ensure it is fully open. For Centrifuge modules with the locking mechanism installed, ensure the safety shield is opened completely to engage the locking mechanism.

NOTE: If the sliding door of the Centrifuge is open, it is not possible to turn the locking knob. In this case, remove sample tubes through the sliding door.

    - Remove all sample tubes loaded in rotor buckets.

    - Close the Centrifuge lid.

    - Remove all sample tubes from bucket inserts in the Centrifuge loading and unloading areas.

    - Close the Centrifuge safety shield and tighten the screws. When closing the Centrifuge module safety shield with a locking mechanism, pull out the lock and lower the safety shield.

WARNING: Pinch Hazard. Use caution when pulling the lock for the locking mechanism. Fingers can be pinched.

    - Turn the locking knob of the Centrifuge clockwise to lock.

    - If two Centrifuge modules are installed, repeat the above steps for the second Centrifuge module.

    - Confirm the prompt by pressing the “ ” button.

    - Remove all filled racks from the Input/Output module.

IMPORTANT: Review the middleware/LIS host to determine which samples have pending tests. Sample tubes for which orders were sent for reruns or additional testing may not be unloaded from the Storage and Retrieval module. A manual delivery of the sample tube may be required. Refer to Delivery of a Sample to a Priority Output Rack.

- Removing Tubes from Workcell if unable to Power-up

Sample tubes should be removed manually from the workcell if all of the following conditions occur:

- Workcell power went off unexpectedly

- There are sample tubes on the workcell

- Power to the workcell is going to be off for an extended period of time such that the integrity of the sample will be compromised.

CAUTION: Do not remove tubes manually unless ALL three of the above conditions are met. Positive identification of samples could be compromised if tubes are manually removed from the workcell when the power is on.

If sample tubes need to be removed manually, use the following procedure.

- Manual removal of sample tubes from the workcell

    - Ensure power is off to the APS workcell

    - Put all Abbott analyzers in a Ready or Stopped status, or equivalent status for non-Abbott analyzers.

    - Open the safety shields or covers for the following modules:

      - I/O module

      - Centrifuge module(s)

      - Decapper module

      - Resealer module

      - Storage and Retrieval module

      - All analyzers

    - Manually remove all accessible sample tubes from tube carriers.

NOTE: Place all samples in a sample rack. They may need to be removed from the On-line worklist later in this procedure

NOTE: All sample tubes located in lanes 1 and 2 of the i 2000SR IM will have to be removed by the local Abbott representative. See Figure E.5 i 2000SR IM Lanes, page E-6

    - Close all the safety shields and replace covers.

    - If there are sample tubes in the Centrifuge Module perform the following steps:

NOTE: If the sliding door of the Centrifuge is open, it is not possible to turn the locking knob. In this case, remove sample tubes through the sliding door.

- Turn the locking knob of the Centrifuge counterclockwise to unlock.

- Use the hex tool (3 mm) to remove the two screws holding the safety shield to the Centrifuge. Set aside the screws.

- Lift the Centrifuge safety shield and open the Centrifuge lid.

- Remove all sample tubes loaded in rotor buckets. Place them in the rack noted in Step 4.

- Close the Centrifuge lid.

- Remove all sample tubes from bucket inserts in the Centrifuge. Place them in the rack noted in Step 4.

- Place the buckets and the balancing tubes in their initial position. See Polar Centrifuge module components, page 1-8 and Cartesian Centrifuge module components, page 1-8.

- Close the Centrifuge safety shield and tighten the screws.

- Turn the locking knob of the Centrifuge clockwise to lock.

- If two Centrifuge modules are installed, repeat the above steps for the second Centrifuge module.

    - If there are sample tubes that are still on the workcell that were not accessible before, the track covers must be removed to have access to those sample tubes:

- Locate the sample tubes and determine which track cover needs to be removed.

- Remove the clear plastic tape that covers the gap between the ends of the track covers.

- Use a screwdriver to remove the two screws from the track covers. Set the screws aside.

- Lift straight up on the track cover to remove it and set the cover aside.

- Remove the sample tubes. Place them in the rack noted in Step 4.

- Replace the track cover. Ensure that the cover is seated in the grooves on the track.

- Replace the track cover screws to secure the cover to the track.

- Replace the clear plastic tape that covers the gap between the ends of the track covers.

- Repeat steps from a through h for all sample tubes remaining on the track.

    - At the next start up all sample tubes that were manually removed may need to be deleted from the On-Line worklist. These tubes should all be in the rack referenced in step 4, step 6d, step 6f and step 7e. Perform procedure Remove a Sample Tube from the On-line Worklist, page 10-108

- Sample Management

Sample management consists of the activities associated with preparing and loading samples, initiating processing, and unloading samples.

- Sample Requirements

It is important to follow the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed, assay-specific information about specimen collection, preparation, and storage.

NOTE: The ACCELERATOR APS does not assess integrity of a specimen for testing after centrifugation and before running on an analyzer. It is the laboratory’s responsibility to determine an acceptable process for specimen preparation and assessment of integrity. Hemolysis, icterus, and lipemia can be semi-quantitatively measured on some chemistry analyzers. Samples requiring operator inspection can be routed to a Sorting lane. This can be done via the host LIS or middleware by ordering an off-track inspection test. Contact your local Abbott representative for additional details.

- Sample Tube Requirements

A sample tube should be plastic with a removable closure on the open end. The tube surface must be smooth and free of any ridges or collars that could interfere with the removal of the tube cap or closure.  Sample tubes must be filled and handled according to laboratory guidelines and procedures. Follow all usual precautions for collecting blood by venipuncture to avoid specimen hemolysis. Refer to Acceptable Sample Tubes for additional tube type information for the ACCELERATOR APS.

IMPORTANT: Use only sample tubes in Acceptable Sample Tubes, section 4-2, pg 12. Sample tubes must meet these specifications.

IMPORTANT: Do not use Glass tubes. Glass tubes cannot be sealed properly. Refer to System (On-line/Off-line), page 2-7 for additional information.

IMPORTANT: Do not use rubber stopped (conventional closure) sample tubes in Dynamic or Input spun lanes. The rubber stoppers must be removed manually due to the rubber causing damage to the Decapper module. Refer to Acceptable Sample Tubes, page 4-16.

IMPORTANT: Follow the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed, assay-specific information about specimen collection, preparation, and storage. Failure to follow tube and assay manufacturer's recommendations for sample tube handling and type could result in erroneous and misleading results.

IMPORTANT: If adequate sample volume is not used, reliability of assay results cannot be guaranteed. Ensure there is sufficient sample volume to run the ordered tests. Due to the workcell’s enclosed, automatic processing of sample tubes, sample volume cannot be inspected visually after centrifugation and before sampling. It is the laboratory’s responsibility to determine an acceptable process to assess sample volume. It may be necessary to centrifuge some samples off-track due to inadequate sample volume. Refer to the specific analyzer operations manual and reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for information regarding sample volume.

- Special Handling

- Always review the analyzer and specific reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) before selecting test methods to be processed on the ACCELERATOR APS workcell. Some methods may have specific instructions that are not appropriate for workcell processing (e.g., methods requiring preprocessing, pretreatment, or special handling).

- Sample carryover requirements for some clinical chemistryanalyzers may be less stringent than for immunoassay analyzers. The possibility of sample carryover should be considered when samples are processed on multiple analyzers. The ACCELERATOR APS workcell routes sample tubes to maximize throughput and sample carryover cannot be controlled by configuring the sample routing order on the APS. It is the laboratory’s responsibility to verify sample carryover is acceptable among all analyzers.

- Sample Processing

Each laboratory should determine their acceptable sample processing times for priority samples. If this time is not being met by loading priority samples in the Priority Input lanes of the Input/Output Module, centrifuge the sample tubes off-track and manually remove caps. Load the sample tubes manually on the analyzer (for example, the instrument sample carousel or the interface module). Using this procedure eliminates any increase in sample processing time.

NOTE: When a Routine sample is loaded into Priority Input lanes the priority of the sample is upgraded to STAT and the workcell sends an upgrade status notification to the Host, if this function has been enabled by the local Abbott representative.

NOTE: When handling uncapped sample tubes, avoid splashing sample outside of the sample tubes.

WARNING: Potential Biohazard. Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

To load samples for priority processing, perform the following steps.

- Determine the lane assignment of the lanes. Refer to Table 5.6 Loading racks, to determine what type of tubes can be placed in racks based on the following:

- Lane assignment for the lane to be used.

- On-line/off-line status of the Centrifuge and Decapper.

IMPORTANT: Verify the rack is loaded correctly with sample tubes, according to instructions described in Loading Sample onto Racks, page 5-47.

- Verify there is not a rack in the lane.

- Grasp the one-column, 12-position Priority rack by the handle and position the rack so the handle is facing you.

- Load the rack into either the Priority I-1 or I-2 lane, performing the following steps.

- Place the rack flat on the worktable.

- Align the edge of the rack with the guides on the worktable.

- Slide the rack into the lane.

- Push firmly until the rack clicks into position and the indicator light turns red.

IMPORTANT: Before processing sample tubes, always verify the rack is properly locked into the lane and detected by the workcell. The LED indicator should be red

WARNING: Pinch Hazard. Never reach under the LED bar to load or unload tubes. You could be injured if your hand enters the Input/Output robot area. Remove the racks from the Input/Output module before loading and unloading sample tubes. Always use the rack handles to insert or remove a rack.

IMPORTANT: Do not use routine or Priority racks if the silicone tube supports are broken or missing. The tubes should be upright and centered when added to the rack. If a rack is damaged, set it aside and contact your local Abbott representative.

IMPORTANT: Always load empty Priority racks into Priority output lanes (Lanes O-1, O-2, O-3) before sample processing. If an error occurs, the sample tube will be returned to the Priority output rack.

It is possible to load an empty routine rack into a routine lane configured as Priority output. If an error occurs, the sample tube will be returned to this rack in case of unavailability of racks in standard Priority output Lanes ((0-1, 0-2, 0-3).

NOTE: When a Routine sample is loaded into Priority Input lanes the priority of the sample is upgraded to STAT and the workcell sends an upgrade status notification to the Host, if this parameter has been enabled by the FSE.

NOTE: When loading samples manually on the c System analyzers, it is recommended to set the analyzers to Off-line on the APS screen prior to opening the cover on the sample carousel. If workcell sample tubes are in the analyzer queue, use Transitioning To Off-line to allow the samples to complete aspiration.

- Loading Racks and Patient Samples

WARNING: Electrical shock hazard. Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

- When loading samples on the ACCELERATOR APS workcell, ensure the following conditions are met.

- Inspect sample tubes and verify they are not cracked or leaking before insertion in the sample racks. Damaged or leaking sample tubes could result in workcell contamination and/or sample carryover.

- Sample tubes must be barcode labeled.

- Ensure barcode label alignment is correct and meets recommended specifications. Refer to Acceptable Sample Tubes, section 4-2 starting on pg 12 for additional information.

- Ensure the label does not extend below the top of the tube carrier. Refer to Sample Barcode Label Length, page 14.

- Do not load duplicate sample IDs on the APS workcell.

- Verify the sample tube size meets the specifications listed in Table 4.9, Sample tube specifications, section 4-2 page 12.

- Ensure there is sufficient sample volume to run the ordered tests.

NOTE: Due to the workcell’s enclosed, automated processing of sample tubes, sample volume cannot be inspected visually after centrifugation and before sampling. It is the laboratory’s responsibility to determine an acceptable process to assess sample volume. It may be necessary to centrifuge some samples off-track due to inadequate sample volume. Refer to the specific analyzer operations manual and reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for information regarding sample volume.

- Follow the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed, assay-specific information about specimen collection, preparation, and storage.

- Ensure complete clot formation in serum specimens has taken place. Some specimens, especially those from patients receiving anticoagulants or thrombolytic therapy, may exhibit increased clotting times. If the specimen is centrifuged before a complete clot forms, the presence of fibrin may cause erroneous results.

- Refer to Sample Requirements, page 4-13 for additional information.

IMPORTANT: Do not run sample tubes if these conditions are not met. Patient results could be affected.

- Loading samples onto racks

There are three different types of lanes on the Input/Output module.

- Routine

- Priority Input

- Priority Output

The table below explains their function:

The racks use silicone rubber guides on two levels to hold the tubes upright and centered so they can be gripped securely by the Input/Output robot.

To load a tube, insert it straight into a rack position, ensuring it is seated firmly against the bottom of the rack and not tilted. Place the first tube in Column A, Row 1.

Continue to fill that column by placing each subsequent sample in the next row of the same column. Continue with the next column, filling each column before inserting a tube in the next. This method of filling racks ensures the most efficient robot access. If the robot encounters an empty position, it does not check for tubes in subsequent positions in that column. This could result in unrecognized sample tubes existing in an output or sorting rack.

WARNING: Failure to load tubes as described could result in sample tube damage, loss of sample, and sample contamination due to placement of tubes into occupied positions.

WARNING: Pinch Hazard. Never reach under the LED bar to load or unload tubes. You could be injured if your hand enters the Input/Output robot area. Remove the racks from the Input/Output module before loading and unloading sample tubes. Always use the rack handles to insert or remove a rack.

Samples can be loaded onto routine and Priority racks at or away from the Input/Output module worktable.

The operator can determine the order of sample processing by loading racks in specific designated positions. Sample tubes that are loaded in the Priority racks in positions I-1 and I-2 are processed before sample tubes loaded in the routine racks in Lanes 1 through 15. Routine racks are processed in the order in which they are loaded into the workcell. For example: if an Input rack is loaded in Lane 10, then another Input rack is loaded in Lane 5, the rack in Lane 10 is processed first followed by the rack in Lane 5.

NOTE: Lanes configured as Input/Dynamic/Spun/Uncapped have a higher priority than a Sealed lane. Sealed lanes have the lowest priority, however, once the I/O robot has started picking up tubes from a rack in a Sealed lane, it will continue to load all the tubes from that rack until it is empty.

NOTE: Priority Output lanes (O-1, O-2, O-3) are used before the routine lane(s) that has been configured as PO.

The workcell accepts capped, uncapped, spun, unspun, and sealed sample tubes, provided they are loaded correctly on the workcell. The routine and Priority racks can be used to group sample tubes. For example, uncapped tubes must be loaded in a separate rack from capped tubes. Uncapped sample tubes must not be centrifuged.

Similarly, spun sample tubes should be placed in a separate rack from unspun sample tubes. Sealed sample tubes can be loaded on a separate rack for routing to the Storage and Retrieval module. Each type of sample tube must be loaded into separate racks and identified appropriately for the workcell. A single rack can hold different acceptable tube sizes.

NOTE: It is very important to place racks in the correct lanes. Always check lane assignments before placing the racks on the Input/Output module.

To check lane assignments, perform the following steps.

- Press the Racks navigation button in the upper part of the screen.

- Press the Lane Setup function button to determine the current lane configuration.

Refer to Checking Status of Racks and Lanes for additional information on lane status.

Refer to I/O Module Lane Assignment Setup for additional information on lane setup.

IMPORTANT: Uncapped tubes must be placed on a separate rack and in a lane specifically assigned for uncapped tubes. Serious damage could result if these sample tubes do not bypass the Centrifuge and, Decapper modules.

IMPORTANT: Placing capped whole blood (unspun) tubes in a lane configured for uncapped or spun tubes could lead to erroneous and misleading results.

IMPORTANT: Sealed tubes already processed that need to be sent directly to the Storage and Retrieval module must be placed in a lane specifically assigned for sealed tubes. Serious damage could result if a sealed tube is transported directly to the Decapper module or analyzers.

IMPORTANT: Before loading a sealed tube for retest, manually remove the foil from the top of the tube. Serious damage could result if a sealed tube is transported to the Decapper module or analyzers.

IMPORTANT: Never use glass tubes if the Resealer module is installed and in on-line status. Glass tubes cannot be sealed properly. Turn the Resealer module off-line when glass tube processing is needed.

- Lane Status

The LED bar is positioned over the Input/Output module worktop. Each LED indicator light illuminates as indicated below.

The following table describes all possible lane statuses depending on lane setup and LED indicators.

Review the Worklist screen for more information about sample tubes contained in the racks.

WARNING: Pinch Hazard. Never reach under the LED bar to load or unload tubes. You could be injured if your hand enters the Input/Output robot area. Remove the racks from the Input/Output module before loading and unloading sample tubes. Always use the rack handles to insert or remove a rack.

- Loading Racks into the Input/Output Module

Before processing, ensure the input racks are in lanes assigned for their appropriate tube types, and empty racks are loaded in both the Priority and routine output lanes.

- Loading Output Racks

Before processing, ensure output racks have been loaded in the Input/Output module. Use the following table to determine where to load output racks in the Input/Output module.

NOTE: If a rack unlatches unexpectedly, do not reinsert the rack into the lane. Check status of all samples located in the rack and take appropriate action. Perform the following procedures to load output racks into the Input/Output module.

WARNING: Potential Biohazard. Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

To load output racks in Priority output lanes, perform the following steps.

- Verify there is not a rack in the lane.

- Grasp an empty one-column, 12-position Priority rack by the handle and position the rack so the handle is facing you.

- Load the rack into the Priority O-1, O-2, or O-3 lane, performing the following steps.

- Place the rack flat on the worktable.

- Align the edge of the rack with the guides on the worktable.

- Slide the rack into the lane.

- Push firmly until the rack clicks into position and the indicator light turns red.

To load output and sorting racks in routine lanes:

- Verify there is not a rack in the lane.

- Grasp an empty four-column 48-position routine rack by the handle and position the rack so the handle is facing you.

- Load the rack into the appropriate routine lane (Lanes 1 through 15), performing the following steps.

- Place the rack flat on the worktable.

- Align the edge of the rack with the guides on the worktable.

- Slide the rack into the lane.

- Push firmly until the rack clicks into position and the indicator light turns red.

- Loading Input Racks for Priority Processing

Perform this procedure to load Priority input racks into the Input/Output module. Samples loaded in the Priority input lanes are processed before samples loaded in routine lanes.

- Loading Input Racks for Routine Processing

Perform this procedure to load routine input racks into the Input/Output module.

NOTE: When handling uncapped sample tubes, avoid splashing sample outside of the sample rack.

WARNING: Potential Biohazard. Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

To load samples for routine processing, perform the following steps.

- Determine the type of tubes which can be placed in the racks, based on the following:

        - Lane assignment for the lane to be used.

        - On-line/off-line status of the Centrifuge and Decapper.

        - Refer to Table 5.6 Loading racks.

IMPORTANT: Verify the rack is loaded correctly with sample tubes, according to instructions described in Loading Samples onto Racks.

- Verify there is not a rack in the lane.

- Grasp the four-column 48-position routine rack by the handle and position the rack so the handle is facing you.

- Load the rack into a routine lane (1-15), performing the following steps

        - Place the rack flat on the worktable.

        - Align the edge of the rack with the guides on the worktable.

        - Slide the rack into the lane.

        - Push firmly until the rack clicks into position and the indicator light turns red.

WARNING: Pinch Hazard. Never reach under the LED bar to load or unload tubes. You could be injured if your hand enters the Input/Output robot area. Remove the racks from the Input/Output module before loading and unloading sample tubes. Always use the rack handles to insert or remove a rack.

IMPORTANT: Do not use routine or Priority racks if the silicone tube supports are broken or missing. The tubes should be upright and centered when added to the rack. If a rack is damaged, set it aside and contact your local Abbott representative.

IMPORTANT: Always load empty Priority racks into Priority output lanes (Lanes O-1, O-2, O-3) before sample processing. If an error occurs, the sample tube will be loaded into the Priority output rack.

- Loading Racks into Appropriate Lanes

Use the following table to load racks into the appropriate lane. For information regarding changing a lane assignment, refer to I/O Module Lane Assignment Setup.

- Initiate or Resume Sample Processing

To initiate or resume processing samples from Pause mode on the ACCELERATOR APS workcell, perform the following steps.

      - If the Exceptions navigation button is yellow or red, check for errors or warnings.

- Press the Exceptions navigation button.

- View any errors or warnings that occurred in a workcell module or instrument.

- Take action as needed. Refer to Section 10, Troubleshooting for additional information.

NOTE: If a rerun is requested and the APS is in Paused mode, APS will not consider the request until the APS is put into Run. When the APS is then put into Run, if the sample is not found on track or in the Input/Output module, a Rerun Exception is generated.

      - Press the Run button to place the workcell in Run mode. The Symbol “” displays when the Run button is selected. When the workcell is in Run mode, the Input/Output module is ready to begin processing sample tubes.

NOTE: After 10 minutes of inactivity, the workcell Track and motor stop, but the workcell remains in Run mode. To resume running, load a rack into the Input/Output module.

- Pause Sample Processing

To pause sample processing from Run mode on the ACCELERATOR APS workcell, perform this procedure. Press the Pause button to place the workcell in Pause mode. This pauses all workcell components. The symbol “” displays when Pause is selected. The workcell continues communication with all modules, analyzers, and host LIS or middleware while in Pause mode.

- Pause

When the Pause navigation button is selected, the following actions take place before the module pauses.

- Checking Status of Racks and Lanes

To check the lane configuration and status of racks loaded on the workcell, perform the following steps.

- Press the Racks navigation button in the upper part of the screen.

- A set of function buttons displays at the bottom of the screen.

- To view information about samples contained in a rack, highlight the rack then press the Selection button. Sample information can also be viewed by pressing the Worklist navigation button. Refer to Checking Details of Samples, page 5-71.

- Racks – On-line Status

To view the status of racks loaded on the workcell, perform the following steps.

- Press the Racks navigation button.

- Press the On-line function button.

The following table describes on-line rack status, depending on lane setup.

- Checking the Status of Samples

Perform this procedure to view sample information. Requests sent from the LIS or middleware and acknowledged by the workcell are included in the Worklist screens.

- Press the Worklist navigation button.

- Select the On-line function button. A list of all samples currently on the workcell displays.

NOTE: For actual sample locations, refer to Figures 5.42 through 5.59.

- How to Check Sample Status on Worklists

Perform this procedure to view sample information from the different Worklist screens.

- Press the Worklist navigation button.

- Press a function button to display a list of specific samples.

NOTE: During normal operation, the selected list must be refreshed to view the current status of samples. Press the function button again to refresh the desired sample list.

NOTE: The screen displays the samples list by pages of 3,000 samples. Highlight Next Page at the top or at the bottom of the screen content area then press the Selection button to display the next 3,000 samples. To return to the previous page highlight Previous Page at the top or at the bottom of the screen content area then press the Selection button.

NOTE: If a worklist screen (Expected, Incomplete or Complete) contains more than 32,768 samples, the list of samples is not displayed. A pop-up displays, refer to Pop-up Message Table, page B-28 to troubleshoot the error condition.

- Highlight a sample to view specific information for that sample.

- Press the Selection button.

See sample screen on the next page…

The following set of function buttons displays in the Sample screen.

From the More... function button, the following specific sample information may display, depending on information available.

- Unloading Racks and Samples

To maintain continuous processing, samples and racks should be unloaded on a routine basis. In addition, it may be necessary to deliver a sample from the workcell or the Storage and Retrieval module to the Priority output lanes.

WARNING: Potential Biohazard. When handling uncapped sample tubes, avoid splashing sample outside of the sample tubes.

Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

To maintain continuous processing, racks should be unloaded on a routine basis.

If the LED indicator is off, proceed to step 5. If it is red or green, begin procedure at step 1.

NOTE: If the LED indicator is red and the sample robot is processing tubes on that lane, a pop-up message displays. Wait for the sample robot to complete processing then the rack can be unlocked.

NOTE: The Priority output rack holds samples with automation errors and those samples requested by the operator to be delivered. Before removal of this rack, an operator may review and/or print the Priority Output screen for a list of sample tubes, their rack location, and the associated error code.

- Press the Pause navigation button to pause the workcell. Ensure the Run/Pause indicator symbol “” displays.

- Press the Racks navigation button.

- Press the On-line function button.

- Highlight the lane of the rack then press the Unlock function button, Refer to Checking Status of Racks and Lanes, page 5-66 for additional information.

- Grasp the handle and slide the rack out of its lane.

- Ensure the rack is completely clear of the guides, and both ends are supported when it is removed from the workcell.

WARNING: Pinch Hazard. Never reach under the LED bar to load or unload tubes. You could be injured if your hand enters the Input/Output robot area. Remove the rack from the Input/Output module before loading and unloading sample tubes. Always use the rack handles to insert or remove a rack.

- Unloading Samples

Perform this procedure to remove a sample tube from the Input/Output module.

WARNING: Potential Biohazard. When handling uncapped sample tubes, avoid splashing sample outside of the sample tubes.

Sample tubes are potentially biohazardous. Follow your laboratory standard operating procedures and guidelines when handling and disposing of tubes.

- Before unloading samples, determine if any sample tubes require additional sample processing by performing one of the following:

- Review the Exceptions screen from the ACCELERATOR APS screen. Refer to the Troubleshooting Section page 10-6 of the Procedure Manual for additional information.

- Review the middleware/LIS host to determine if any samples have pending tests for this analyzer.

- The operator must use laboratory operating procedures to ensure all samples have been processed, and all tests have been completed.

- Store sample tubes according to laboratory guidelines.

- Delivery of a Sample to a Priority Output Rack

Use this procedure to return a sample from the workcell or the Storage and Retrieval module to the Input/Output module. The sample tube is delivered to the Priority output rack with no further processing performed.

NOTE: If Delivery is requested for a sample tube that is in the Centrifuge module, but has not yet started spinning, all samples will be removed from the centrifuge module without being spun. The requested tube is placed in a priority output rack; the remaining tubes are unloaded to the track and rerouted to the centrifuge module.

- Press the Worklist navigation button.

- If the sample is currently on the workcell, press the On-line function button. A list of samples currently on the workcell displays.

- If the sample tube has been stored in the Storage and Retrieval module, perform the following steps.

      - Press the Search function button to locate the sample.

      - Enter the sample ID in the pop-up keyboard then press the Confirm button.

      - Highlight the sample to be accessed from the Worklist screen.

      - Press the Selection button.

      - Select the Deliver function button.

      - Confirm the pop-up “Are you sure?” by pressing the “ ” button to accept.

      - The sample tube is delivered to the Priority output lane.

NOTE: It is possible to configure a routine lane as Priority Output. Priority Output lanes (O-1, O-2, O-3) are used before the routine lane(s) that has been configured as PO.

- Purging Samples

Samples may need to be purged from the workcell if the workcell must be paused or powered down for an extended period of time, e.g., when maintenance will be performed. Purged samples are returned to Priority Output lanes.

NOTE: During the Purging Samples procedure, samples in the Centrifuge are returned to the Priority output lane when centrifugation is complete.

IMPORTANT: Sample integrity could be compromised if samples are left on the workcell when power is shut down. Refer to your laboratory specific guidelines or the reagent manufacturer’s assay-specific documentation (e.g., a package insert or reagent application sheet) for detailed specimen collection, preparation, and storage information.

NOTE: This Purge procedure does not purge tube carriers from the ARCHITECT i 2000SR Interface module (i 2000SR IM) or the ARCHITECT i 2000SR Entry gate. Perform Purging Samples from i 2000SR IM, page E-148 for each ARCHITECT i 2000SR connected to the APS workcell.

- Perform the following steps to display the LOCATION column of the Worklist screen.

- Press the Worklist navigation button.

- Press the On-line function button.

- If sample tubes display in the LOCATION column, it is necessary to perform the Purging Samples procedure to remove these sample tubes from the workcell. If no sample tubes remain on the workcell, no further action is needed.

- Remove racks from Lanes 1 through 15. Refer to Unloading Racks from the Input/Output Module, page 5-91.

- Press the Racks navigation button.

- Press the On-line function button.

- Highlight the lane of the rack then press the Unlock function button, if necessary.

- Verify all installed Centrifuge modules are on-line.

- Press the System navigation button.

- Press the Configuration option button.

- Press the System function button at the bottom of the screen.

- Highlight the necessary module to turn on-line.

IMPORTANT: If the Centrifuges are off-line, samples cannot be purged from the Centrifuge area. Refer to System (On-line/Off-line), page 2-7.

Select “” to accept or “” to cancel.

- Allow any sample tubes loaded in tube carriers at the Input/Output module to be scanned by the barcode reader before proceeding.

- Perform the following steps to start the Purge procedure.

- Press the System navigation button.

- Press the Configuration option button.

- Press the Purge function button at the bottom of the screen.

- The pop-up message “Purge all samples from System” displays.

- Select the confirmation button by pressing “” to confirm the Purge procedure.

- Load empty Priority output racks and replace until all samples have been removed from the Track module.

NOTE: It is possible to configure a routine lane as Priority output. If a routine lane is configured as Priority output, it is used to unload purged samples if the standard Priority output racks are unavailable.

- The Purge procedure continues until all samples are unloaded from the workcell.

NOTE: Sample tubes in the Centrifuge must complete centrifugation before delivery to the Priority output lane.

- To verify the Purge procedure is running:

- Press the System navigation button.

- Press the Diagnostics option button.

- Press the Automation function button.

- Press the I/O Module function button.

- Locate the line Purge all samples from system to determine procedure status.

- Verify all samples were purged by checking for on-line samples on the Worklist screen.

- Press the Worklist navigation button.

- Press the On-line function button.

- There will not be any on-line samples when the Purge procedure has completed.

- When sample tubes are located only on the workcell Track, the Purge procedure must be ended manually. Refer to Checking Details of Samples, page 5-71.

- Press the Purge function button at the bottom of the Configuration screen.

- The pop-up message “Quit purging samples” displays.

- Select the confirmation button by pressing “” to confirm the Quit Purging procedure.