## Slide 1: Architect maintenance, qc and Calibration process update
Laboratory Meeting October 4, 2017

## Slide 2: Architect current process – day shift
Analyzer Maintenance
  Daily
  Weekly
  Monthly
  Quarterly
Calibration
  As Needed
Daily QC on CAIV 1 & 2
  Chem1 & 2
  Urine Chem 1 & 2
  CSF 1 & 2
  Ped 2
  BHOB 1, 2, 3
  Transplant 1, 2, 3

## Slide 3: Architect current process – Evening shift
Analyzer Maintenance
  None
Calibration on CAIV 1
  Lytes
  Mg
  Acetaminophen
  Ammonia
  ETOH
  As Needed
Daily QC on CAIV 1
  CHEM 1 & 2
    Lytes
    Mg
  LIG 1 & 2
  AAM 1 & 3
  CRD 1 & 2
  CRP 3
  BIOF 3
  Urine Chem 1 & 2
    Lytes
Daily QC on CAIV 2
  CHEM 1 & 2
    Lytes
  LIG 1 & 2
  CRD 1 & 2
  CRP 3
  BIOF 3
  Urine Chem 1 & 2
    Lytes

## Slide 4: Architect current process – Night shift
Analyzer Maintenance
  None
Calibration on CAIV 2
  Lytes
  Mg
  Ammonia
  ETOH
  As Needed
Daily QC on CAIV2
  CHEM 1 & 2
    Lytes
    Mg
  AAM1 & 3
  Urine Chem 1 & 2
    Lytes
Daily QC on CAIV1
  CHEM 1 & 2
    Lytes
  Urine Chem 1 & 2
    Lytes

## Slide 5: Architect NEW process – day shift
Analyzer Maintenance
  Daily
  Weekly
  Monthly
  Quarterly
Daily QC on CAIV 1 & 2
  Chem1 & 2
  Urine Chem 1 & 2
  CSF 1 & 2
  Ped 2
  BHOB 1, 2, 3
  Transplant 1, 2, 3
No Changes to the Current Process

## Slide 6: Architect NEW process – Evening shift
Analyzer Maintenance
  All Maintenance
Calibration on CAIV 1
  Lytes
  Mg
  Acetaminophen
  Ammonia
  ETOH
Daily QC on CAIV 1
  CHEM 1 & 2
    All Tests
  LIG 1 & 2
  AAM 1 & 3
  CRD 1 & 2
  CRP 3
  BIOF 3
  Urine Chem 1 & 2
    Lytes
Daily QC on CAIV 2
  CHEM 1 & 2
    All Tests
  LIG 1 & 2
  CRD 1 & 2
  CRP 3
  BIOF 3
  Urine Chem 1 & 2
    Lytes
CHEM 1 & 2 – all tests will be run
All Maintenance – when adequately staffed

## Slide 7: Architect NEW process – Night shift
Analyzer Maintenance
  None
  Calibration on CAIV 2
  Lytes
  Mg
  Ammonia
  ETOH
Daily QC on CAIV2
  CHEM 1 & 2
    Lytes
    Mg
  AAM1 & 3
  Urine Chem 1 & 2
    Lytes
Daily QC on CAIV1
  CHEM 1 & 2
    Lytes
  Urine Chem 1 & 2
    Lytes
CHEM 1 & 2 – all tests will be run

## Slide 8: Architect Current Calibration Process
Calibrations performed as needed and daily as required
Calibration reports printed automatically
Lot to Lot correlations performed for new lot of reagents loaded
QC performed

## Slide 9: Architect NEW Calibration Process
Calibrations performed as needed and daily as required
Calibration reports printed automatically
Calibration Curve Reviewed on the Architect – refer to procedure
Calibration Curve Printed from the Architect – refer to procedure
2 patient pre/post calibration correlation performed – 1 normal & 1 abnormal result
Correlations documented on the lot to lot correlation workbook and print
QC performed
Calibrations and Correlations will be given to the technical supervisor to be filed by instrument and test for easy retrieval during troubleshooting

## Slide 10: Reason for the architect process changes
During a recent PHOS calibration, an air bubble was aspirated as part of the blank which resulted in a bad calibration; however the instrument did not fail the calibration.
QC was acceptable and patient results were reported
A physician and a tech questioned patient results around the same time
Patients were repeated on the other Architect with significantly higher results
In all, over 60 patient tests results had to be corrected

## Slide 11: Reason for the architect process changes
Lipase was calibrated
New Lipase reagent (same lot number) was loaded
QC was run by day shift and was acceptable
3 patient had results of less than the lower reportable limit
QC was performed and failed
It was identified that the reagent handling instructions were not performed prior to loading the reagent

## Slide 12: Reason for the architect process changes
Chemistry controls were acceptable on night shift
Chemistry controls were not acceptable on day shift
The mixer was bent but the analyzer did not alert the tech to the issue
Patient correlations were performed to ensure accurate results

## Slide 13: Reason for the architect process changes
Due to the previously identified issues
The lack of notification by the Architects when there is a bad calibration
The lack of alerts by the Architects when the instrument malfunctions
To ensure accurate results for all patients, especially during the busy morning time
ALL SHIFTS WILL RUN CHEMISTRY 1 & 2 CONTROLS
MAINTENANCE WILL BE PERFORMED ON EVENING SHIFT ONCE STAFFING LEVELS ARE ADEQUATE