

ALL TESTS GO LIVE WITH CONNECTIVITY TO EPIC

SIMULTANEOUS WITH COVID IMPLEMENTATION

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | | | | |
| Covid training points in RED | **ID NOW™ COVID-19** | **ID NOW™ Strep A 2** | **ID NOW™ Flu A & B 2** | **ID NOW™ RSV** |
| **LAB ORDER**  **Must be placed in EPIC before analyzing test or result will not upload to EMR** | SPL555  AFFILIATE POC COVID-19 Symptomatic  SPL557  AFFILIATE POC COVID-19 Exposure | SPL525  AFFILIATE POC MOLECULAR GROUP A STREP SCREEN | SPL515  AFFILIATE POC INFLUENZA A&B (MOLECULAR) | SPL516  AFFILIATE POC RSV (MOLECULAR) |
| **GIVE PATIENT**  **FACT SHEET** | **Yes – required by FDA**  (provided in kit cartridge boxes) | No | No | No |
| **SWAB TYPE** | NP (1 nare)  (NP not provided in Covid cartridge boxes. Order and use RSV swabs from Abbott  for Covid testing) | Throat swab  (provided in kit cartridge boxes) | Nasal swab (1 nare)  (provided in kit cartridge boxes) | NP (1 nare)  (provided in kit cartridge boxes) |
| **USE OF TRANSPORT MEDIA**  **VIRAL OR OTHER** | **No** | No | No | No |
| **ONE PATIENT SWAB USED TO PERFORM MULTIPLE TESTS** | No – Single use only | No – Single use only | No – Single use only | No – Single use only |
| **SAMPLE COLLECTION TO**  **ANALYSIS TIME FRAME** | 1 hour  room temp  (Do not refrigerate) | Up to 72 hours  room temp & refrigerated | Up to:  2 hours room temp  24 hours refrigerated | Up to:  2 hours room temp  24 hours refrigerated |
| **ID NOW INSTRUMENT PATIENT TESTING PROCEDURE** | Identical | Identical  (except ID Now doesn’t initiate warming cycle) | Identical | Identical |
| **NEGATIVE RESULT**  **TEST TIME**  Can be several minutes longer for results  to post automatically to EPIC | 13 min | 6 min | 13 min | <13 min |
| **POSITIVE RESULT**  **TEST TIME** | 5-13 min | 2 - 6 min | 5-13 min | <13 min |

**NEW STEPS/REQUIREMENTS FOR ALL IDNOW TESTS AFTER INTERFACING TO EPIC:**

* Patient results do not need to be recorded as results will automatically upload to epic.
* Neg and pos qc lot numbers still need to be recorded on qc log.

**TESTING REMINDER FOR ALL IDNOW TESTS:**

* IF YOU GET AN INVALID, YOU CAN NOT RE-TEST USING THE ORIGINAL SWAB. Once the swab has been dipped in a reagent, it must be discarded. Please see the IDNow procedure for instructions on re-testing a specimen after getting an invalid test. It is best not to get them as this may cause issues with posting the result to epic.
* **CLEAN INSTRUMENT** with 10% bleach before each patient and qc test.
* **SCAN** your employee badge barcode when prompted, DO NOT USE INITIALS.
* **SCAN** patient beaker label (large label) when prompted.
* **CHANGE GLOVES AFTER SWAB MIXING – WILL NOT PROMPT.**