## Slide 1: Nitrazine pH Test for Rupture of Membranes

## Slide 2: Nitrazine pH test for ROM
Tests for pH of the vaginal pool in expectant mothers
Vaginal pH is normally acidic, pH <7.0
pH >7.0 indicates that the amniotic membranes have ruptured

## Slide 3: Nitrazine pH paper
qUAntify control material for Nitrazine pH

## Slide 4: Quality Control
Quality Control (QC) is performed daily to verify that all components are working properly
Two levels of QC are run
  Negative; pH <7.0 yellow
  Positive ; pH >7.0 blue
Record information on QC and patient log
  Nitrazine pH paper and QC material lot number, open and expiration dates (expires 31 days after opening)
  QC results
Do not proceed with patient testing if QC fails
  Corrective action for failed QC:
    Repeat test once – if QC fails, then
    Use new pH paper or new vial QC
    If QC fails again, report to POC Coordinator in Lab

## Slide 5
 | Lot Number | Open Date | Expiration Date | Expected Value
qUAntify Level 1 |  |  |  | pH <7.0 (yellow) = Negative
qUAntify Level 2 |  |  |  | pH >7.0 (blue) = Positive
Nitrazine Paper |  |  |  | ------------------------------------
 | Level 1 | Level 2 | Initials
1 |  |  | 
2 |  |  | 
3 |  |  | 
4 |  |  | 
5 |  |  | 
6 |  |  | 
7 |  |  | 
8 |  |  | 
9 |  |  | 
10 |  |  | 
11 |  |  | 
12 |  |  | 
13 |  |  | 
14 |  |  | 
15 |  |  | 
16 |  |  | 
17 |  |  | 
18 |  |  | 
19 |  |  | 
20 |  |  | 
21 |  |  | 
22 |  |  | 
23 |  |  | 
24 |  |  | 
25 |  |  | 
26 |  |  | 
27 |  |  | 
28 |  |  | 
29 |  |  | 
30 |  |  | 
31 |  |  | 
Date | Name | DOB | Result | Initials
MBCU Nitrazine Paper pH Quality Control and Patient Log
Month/Year____________________
Quality Control
Patient Testing

## Slide 6: Quality Control (continued)
QC material is:
  Kept at Room Temp in MBCU nurses station once opened
  Good opened, at room temperature for 31 days
  Stored unopened refrigerated at 2-8 C in the Laboratory Point of Care (POC) refrigerator for monitoring of temperature
  Good unopened, at 2-8 C until manufacturer’s expiration date

## Slide 7: Quality Control (continued)
Procedure
  Tear off two strips of nitrazine pH paper
  Place a drop of level one on one paper
  Place a drop of level two on the other
    Level one = Negative; pH <7.0 yellow
    Level two = Positive ; pH >7.0 blue
  Record results on log sheet
  Do not proceed with patient testing if QC fails
  Corrective action for failed QC:
    Repeat test once – if QC fails, then
    Use new pH paper or new vial QC
    If QC fails again, report to POC Coordinator in Lab

## Slide 8: Nitrazine pH Test Results
Yellow=negative
Blue=positive

## Slide 9: Patient Testing
Patient testing may not be performed unless QC has been performed on that day
Verify patient name and date of birth
Obtain small strip of nitrazine paper
Touch the nitrazine paper to the sample from the vaginal pool within 2-5 minutes at room temperature
Compare the color of the nitrazine paper to the color chart on the paper dispenser’s pH scale
  pH below 7.0 (yellow) is negative
  pH above 7.0 (blue) is positive
Document results on log sheet and in patient’s chart

## Slide 10: Patient Testing (continued)
TEST INTERFERENCE:
  Exposure to air may cause a change in pH
    Test within 2-5 minutes
  Bloody specimens may make it difficult to interpret the color reaction
  False positives may be due to contamination with
    Alkaline urine
    Bathwater

## Slide 11: Proficiency Testing
Proficiency Tests:
  Unknown samples that are tested 2-3 times a year
  Testing is performed on a rotational basis by staff who perform patient testing
  Results are submitted to CAP for a grade
  Failing grade of <80% consecutively may result in removal of testing from CMH

## Slide 12: Competency Testing
Training is performed initially on MBCU by the designated trainer
Competency is required at 6 months, on initial anniversary, and annually thereafter.
  For example: A new RN begins on 9/1/13
    Training – 9/1/13
    6 month Competency - 3/1/14
    Annual – 9/1/14, 9/1/15, etc.